## CoverSheet
Analytical Analysis Cover Sheet for Hepatic Clearance Assessment of PFAS Compounds
QAPP: | In Vitro Toxicokinetic (TK) Assessments of PFAS (D-CCED-0031343-1-1)
PI: | Barbara A. Wetmore
Analysis Set ID: | HepCl_ReRun
Analytes: | 6:3 FTCA, 6:2 FTS, 9H-PFNA, 8:2 FTS, 11H-PFUnDA, PFUnDA
Calibration Curve Prepared Date: | 2023-11-03
Preparer: | KR
Sample Type: | Calibration Curve (CC1 - CC17), Curve Precision Check (QC-A - QC-D)
Hepatic Clearance Assay Conducted Date: | 2023-11-03
Preparer: | KR, MP
Sample Type: | Hepatic Clearance Samples
Analytical Analysis Date: | 11/6-9/2023
Analytical Technique: | UPLC-MS/MS
Instrument: | Waters Xevo TQ-S micro (QEB0036)
Analyst: | KR
Relevant SOPs: | I-CCED-ETTB-4568; I-CCED-ETTB-SOP-3632
Notebook Entries: | Hepatocytic Clearance (Human) of Select Compounds
Total Samples Number (including blanks): | 255
Actual Sample Number: | 160
Method Detection Limit (MDL) Definition: | Minimum measured concentration of a substance that can be reported with 99% confidence that the measured concentration is distinguishable from method blank results
MDLXi = method detection limit determined for target analyte X from solution i = t(n-1,1-α=0.99) x Si
t(n-1,1-α=0.99) = 3.143, the student's t-value for a 99% confidence level and a standard deviation estimate with six degrees of freedom (n=7) |  |  |  | 40 CFR Part 136 | EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016)
Si = standard deviation of the seven replicate analyses for solution i |  |  |  | https://www.ecfr.gov/cgi-bin/text-idx?SID=a68a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
estimated Limit of Quantitation (eLOQ) Definition: | Lowest Calibration Curve point at which precision checks can be quantitated against curve ± 30%, provided other curve points have accuracy to ± 25%
QC Samples and Acceptance Criteria
QC Type | Frequency | Acceptance Criteria | Type | Corrective Action if QC Fails
Calibration Curve Linearity (Analytical) | 1 each run | 7-point curve min; r2≥0.98 | Linearity | Re-run samples and curve
Curve Precision Check (Analytical) | 2-3 samples per run | 75-125% of expected | Accuracy | Flag in comments
Blanks (Analytical) | 1 every 6-10 injections | <1/2 MDL of analytes | Specificity | Flag in comments, assess and resolve
Technical Replicate Assessment (Assay) | Every assay sample | Individual Replicates within ± 25% of Sample Mean | Precision | Flag in comments
Positive Control Reference Chemical(s) for Protein Binding (Assay) | 1 per assay batch | ± 20% of historical or published values | Binding Assay Performance | Flag in comments
Note: one of the three blanks (solvent, crash blank, crash mixed matrix) should be run every 6 injections
Analytes, Internal Standards Assessed
Analyte (Common Name) | DTXSID | Sample ID | Average Molecular Weight (g/mol) | Internal Standard | Peak Assignment from CoA | Mix | Concentration at Instrument (pg/uL ¦ µM)
4,4,5,5,6,6,7,7,8,8,9,9,9-Tridecafluorononanoic acid (6:3 FTCA) | DTXSID70379917 | 1213130502 | 392.116 | M2-6:2 FTS | H | 1 | 0.01
6:2 Fluorotelomer sulfonic acid (6:2 FTS) | DTXSID6067331 | 351896986 | 428.16 | M2-6:2 FTS | H | 2 | 0.01
9H-Perfluorononanoic acid (9H-PFNA) | DTXSID50226894 | 1208990480 | 446.087 | M9PFNA | J | 3 | 0.01
8:2 Fluorotelomer sulfonic acid (8:2 FTS) | DTXSID00192353 | 351896630 | 528.18 | M2-8:2 FTS | L | 2 | 0.01
11-H-Perfluoroundecanoic acid (11H-PFUnDA) | DTXSID5061954 | 1208990465 | 545.973 | M7PFUdA | Q | 3 | 0.01
Perfluoroundecanoic acid (PFUnDA) | DTXSID8047553 | 351896652 | 564.093 | M7PFUdA | Q | 1 | 0.01
Phenacetin | DTXSID1021116 |  | 179.219
Propranolol | DTXSID6023525 |  | 259.349
Internal Standard Information
Standard | Vendor | Catalog Number | Lot Number | Other Descriptors
MPFAC-24ES | Wellington | MPFAC-23ES | MPFAC24ES0123
Hepatocytes | Vendor | Catalog Number | Lot Number | Species
LiverPool 20 Donor Hepatocytes Mixed Donor | BioIVT | X008000 | WGR | Human
Sample ID Key
CC## | Calibration Curve
QC-X | Quality Check
Human Mix # Active T# X | Species Mix # SampleType Time# Replicate | [6 empty] | An Example Certificate of Analysis for Wellington MPFAC-24ES Mix (Mass-labelled PFAS Solution)
*MP = Matt Phillips
*KR = Krishna Ravindra
*BAW = Barbara Wetmore
Additional notes
Concentrations listed are at the instrument, unless otherwise specified

## Clint Data Summary
Hepatic Clearance Data Summary
Human | [21 empty] | No background clearance | HT
DTXSID | Common Name | Sample ID | Mix/SPE | Analyte | Best-Fit Values | [8 empty] | No Cells BkgdCl |  |  |  | HT BkgdCl |  |  |  | Clearance (µL/min/106 hepatocytes) | Bkgd-Adj Clearance
[5 empty] | Slope | r2 | Sy.x | F | DFn | DFd | P value | Half Life | Clearance | pvalue | Slope | Half Life | Clearance | pvalue | Slope | Half Life | Clearance
DTXSID70379917 | 6:3 FTCA | 1213130502 | 1 | 6:3 FTCA | -0.002523 | 0.7085 | 0.1414 | 38.89 | 1 | 16 | <0.0001 | 274.673 | 5.046 | 0.1409 | -0.0002502 | 2769.7842 | 0.5004 | 0.4364 | -0.00019 | 3647.3684 | 0.38 | 5.046
DTXSID6067331 | 6:2 FTS | 351896986 | 2 | 6:2 FTS | -0.0007007 | 0.2844 | 0.09708 | 6.359 | 1 | 16 | 0.0227 | 989.011 | 1.4014 | 0.7279 | -7.627e-05 | 9086.1413 | 0.15254 | 0.5505 | -0.0002074 | 3341.3693 | 0.4148 | 1.4014
DTXSID50226894 | 9H-PFNA | 1208990480 | 3 | 9H-PFNA | -0.0006392 | 0.07815 | 0.1918 | 1.356 | 1 | 16 | 0.2612 | 1084.1677 | 1.2784 | 0.086 | -0.001034 | 670.2128 | 2.068 | 0.9328 | -5.303e-05 | 13068.0747 | 0.10606 | 1.2784
DTXSID00192353 | 8:2 FTS | 351896630 | 2 | 8:2 FTS | -0.0007642 | 0.4051 | 0.08088 | 10.9 | 1 | 16 | 0.0045 | 906.8307 | 1.5284 | 0.1256 | -0.0003565 | 1943.899 | 0.713 | 0.834 | -5.617e-05 | 12337.5467 | 0.11234 | 1.5284
DTXSID5061954 | 11H-PFUnDA | 1208990465 | 3 | 11H-PFUnDA | -0.000286 | 0.05744 | 0.1012 | 0.9751 | 1 | 16 | 0.3381 | 2423.0769 | 0.572 | 0.5549 | -0.0004507 | 1537.6082 | 0.9014 | 0.7516 | 0.0001967 | -3523.1317 | -0.3934 | 0.572
DTXSID8047553 | PFUnDA | 351896652 | 1 | PFUnDA | -0.0001011 | 0.02235 | 0.05843 | 0.3658 | 1 | 16 | 0.5538 | 6854.5994 | 0.2022 | 0.477 | -0.0001391 | 4982.0273 | 0.2782 | 0.0087 | 0.000402 | -1723.8806 | -0.804 | 0.2022
[14 empty] | All slopes not significantly non-zero |  |  |  | Significantly non-zero slope, but positive

## 6 3 FTCA
Human Active Cells | [8 empty] | Media Only | [7 empty] | 6:3 FTCA
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.88 | 447.461 | 8778 | 3034.11 | 0.59 | 1.825 | 1.43 | -21.8
0 | Mix 1 Active T0 A | 47.47 | 6.6327 | 759.52 | 842.24 | 142.9982 | 16.98% |  | 0 | Mix 1 Media T0 A | 49.15 | 6.6675 | 786.4 | 782.1867 | 11.8555 | 1.516% | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.88 | 478.13 | 8539 | 2368.634 | 0.807 | 1.825 | 2.22 | 21.7
| Mix 1 Active T0 B | 62.96 | 6.9151 | 1007.36 | [5 empty] | Mix 1 Media T0 B | 48.05 | 6.6448 | 768.8 |  |  |  | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.87 | 418.264 | 7300 | 2974.036 | 0.563 | 1.825 | 1.33 | -27.3
| Mix 1 Active T0 C | 47.49 | 6.6331 | 759.84 | [5 empty] | Mix 1 Media T0 C | 49.46 | 6.6738 | 791.36 |  |  |  | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.88 | 523.429 | 9106 | 2930.851 | 0.714 | 1.825 | 1.88 | 3.1
15 | Mix 1 Active T15 A | 39.94 | 6.46 | 639.04 | 733.8667 | 121.5411 | 16.56% |  | 15 | Mix 1 Media T15 A | 53.95 | 6.7606 | 863.2 | 809.2267 | 54.6522 | 6.754% | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.88 | 452.235 | 7649 | 2982.801 | 0.606 | 1.825 | 1.49 | -18.5
| Mix 1 Active T15 B | 43.23 | 6.5391 | 691.68 | [5 empty] | Mix 1 Media T15 B | 47.12 | 6.6253 | 753.92 |  |  |  | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.88 | 472.381 | 7598 | 3064.16 | 0.617 | 1.825 | 1.52 | -16.5
| Mix 1 Active T15 C | 54.43 | 6.7695 | 870.88 | [5 empty] | Mix 1 Media T15 C | 50.66 | 6.6977 | 810.56 |  |  |  | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.88 | 432.128 | 7512 | 3262.376 | 0.53 | 1.825 | 1.21 | -33.9
30 | Mix 1 Active T30 A | 40.08 | 6.4635 | 641.28 | 656.6933 | 26.9743 | 4.108% |  | 30 | Mix 1 Media T30 A | 54.99 | 6.7797 | 879.84 | 813.6533 | 58.1851 | 7.151% | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.88 | 465.458 | 7889 | 3313.979 | 0.562 | 1.825 | 1.32 | -27.5
| Mix 1 Active T30 B | 42.99 | 6.5336 | 687.84 | [5 empty] | Mix 1 Media T30 B | 49.41 | 6.6727 | 790.56 |  |  |  | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.88 | 779.956 | 11885 | 2993.711 | 1.042 | 2.919 | 3.08 | 5.5
| Mix 1 Active T30 C | 40.06 | 6.463 | 640.96 | [5 empty] | Mix 1 Media T30 C | 48.16 | 6.6471 | 770.56 |  |  |  | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.88 | 712.941 | 14349 | 3350.341 | 0.851 | 2.919 | 2.38 | -18.4
60 | Mix 1 Active T60 A | 38.14 | 6.4139 | 610.24 | 573.76 | 35.4132 | 6.172% |  | 60 | Mix 1 Media T60 A | 46.73 | 6.617 | 747.68 | 789.44 | 36.1759 | 4.582% | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.88 | 1201.809 | 19563 | 3101.963 | 1.55 | 4.669 | 4.94 | 5.7
| Mix 1 Active T60 B | 33.72 | 6.2907 | 539.52 | [5 empty] | Mix 1 Media T60 B | 50.59 | 6.6963 | 809.44 |  |  |  | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.87 | 1077.497 | 18622 | 2541.916 | 1.696 | 4.669 | 5.47 | 17.1
| Mix 1 Active T60 C | 35.72 | 6.3483 | 571.52 | [5 empty] | Mix 1 Media T60 C | 50.7 | 6.6985 | 811.2 |  |  |  | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.88 | 1982.509 | 35805 | 3138.186 | 2.527 | 7.45 | 8.51 | 14.2
120 | Mix 1 Active T120 A | 40.62 | 6.4768 | 649.92 | 597.4933 | 45.7769 | 7.661% |  | 120 | Mix 1 Media T120 A | 54.43 | 6.7695 | 870.88 | 824.8533 | 57.2112 | 6.936% | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.87 | 2015.593 | 34908 | 3151.321 | 2.558 | 7.45 | 8.62 | 15.7
| Mix 1 Active T120 B | 35.34 | 6.3376 | 565.44 | [5 empty] | Mix 1 Media T120 B | 52.68 | 6.7368 | 842.88 |  |  |  | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.87 | 2973.767 | 47181 | 3163.276 | 3.76 | 11.875 | 13.02 | 9.6
| Mix 1 Active T120 C | 36.07 | 6.3581 | 577.12 | [5 empty] | Mix 1 Media T120 C | 47.55 | 6.6344 | 760.8 |  |  |  | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.88 | 3366.109 | 59308 | 3307.122 | 4.071 | 11.875 | 14.15 | 19.2
240 | Mix 1 Active T240 A | 26.76 | 6.0595 | 428.16 | 417.4933 | 85.7789 | 20.55% |  | 240 | Mix 1 Media T240 A | 45.68 | 6.5942 | 730.88 | 743.3067 | 34.5964 | 4.654% | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.88 | 5696.95 | 94449 | 3229.261 | 7.057 | 19.125 | 25.07 | 31.1
| Mix 1 Active T240 B | 20.43 | 5.7896 | 326.88 | [5 empty] | Mix 1 Media T240 B | 48.9 | 6.6624 | 782.4 |  |  |  | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.88 | 5013.81 | 76829 | 3103.845 | 6.461 | 19.125 | 22.89 | 19.7
| Mix 1 Active T240 C | 31.09 | 6.2095 | 497.44 | [5 empty] | Mix 1 Media T240 C | 44.79 | 6.5746 | 716.64 |  |  |  | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.87 | 6181.179 | 122226 | 2767.938 | 8.933 | 30.5 | 31.93 | 4.7
[17 empty] | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.88 | 8168.249 | 142368 | 3365.248 | 9.709 | 30.5 | 34.77 | 14
Human Inactive Cells | [16 empty] | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.88 | 11224.8 | 178932 | 3439.733 | 13.053 | 48.75 | 46.99 | -3.6
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.88 | 12265.34 | 222118 | 3673.549 | 13.355 | 48.75 | 48.1 | -1.3
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.88 | 22367.092 | 397198 | 3919.615 | 22.826 | 78.125 | 82.72 | 5.9
0 | Mix 1 Inactive T0 A | 46.64 | 6.615 | 746.24 | 764.2667 | 42.9386 | 5.618% | [9 empty] | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.88 | 21361.063 | 375481 | 4066.942 | 21.009 | 78.125 | 76.08 | -2.6
| Mix 1 Inactive T0 B | 45.83 | 6.5975 | 733.28 | [12 empty] | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.88 | 30114.336 | 533010 | 3888.294 | 30.979 | 125 | 112.53 | -10
| Mix 1 Inactive T0 C | 50.83 | 6.7011 | 813.28 | [12 empty] | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.88 | 32258.174 | 560369 | 3980.426 | 32.417 | 125 | 117.79 | -5.8
240 | Mix 1 Inactive T240 A | 41.94 | 6.5088 | 671.04 | 718.72 | 67.4297 | 9.382% | [9 empty] | 20231106_HepCl_ReRun_042 | QC B | QC | 3.87 | 710.963 | 12145 | 3127.347 | 0.909 | 2.506 | 2.59 | 3.5
| Mix 1 Inactive T240 B | 47.9 | 6.6417 | 766.4 | [12 empty] | 20231106_HepCl_ReRun_081 | QC B | QC | 3.87 | 752.553 | 12162 | 2887.97 | 1.042 | 2.506 | 3.08 | 22.9
| Mix 1 Inactive T240 C | 47.17 | 6.6263 | 754.72 | [12 empty] | 20231106_HepCl_ReRun_043 | QC C | QC | 3.88 | 1686.431 | 30672 | 3371.759 | 2.001 | 6.269 | 6.58 | 5
[17 empty] | 20231106_HepCl_ReRun_082 | QC C | QC | 3.88 | 2095.343 | 37365 | 3610.88 | 2.321 | 6.269 | 7.76 | 23.7
| Calibration Curve | [15 empty] | 20231106_HepCl_ReRun_044 | QC D | QC | 3.88 | 6550.212 | 126329 | 3225.131 | 8.124 | 25 | 28.97 | 15.9
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20231106_HepCl_ReRun_083 | QC D | QC | 3.88 | 7293.705 | 119126 | 3571.287 | 8.169 | 25 | 29.14 | 16.5
| 20231106_HepCl_ReRun_013 | CC06 | 1.825 | 1.43 | 1.55 | 0.33671 | 0.21723 | 1.0583 | -0.15068 | [7 empty] | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.89 | 15743.503 | 247765 | 4776.272 | 13.185 |  | 47.47
| 20231106_HepCl_ReRun_031 | CC06 | 1.825 | 2.22 | [12 empty] | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.88 | 13915.396 | 245491 | 3195.141 | 17.421 |  | 62.96
| 20231106_HepCl_ReRun_035 | CC06 | 1.825 | 1.33 | [12 empty] | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.88 | 14715.912 | 240678 | 4462.737 | 13.19 |  | 47.49
| 20231106_HepCl_ReRun_039 | CC06 | 1.825 | 1.88 | [12 empty] | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.88 | 12368.906 | 206371 | 4447.937 | 11.123 |  | 39.94
| 20231106_HepCl_ReRun_052 | CC06 | 1.825 | 1.49 | [12 empty] | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.88 | 12134.313 | 203497 | 4036.639 | 12.024 |  | 43.23
| 20231106_HepCl_ReRun_070 | CC06 | 1.825 | 1.52 | [12 empty] | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.88 | 11903.52 | 204453 | 3155.708 | 15.088 |  | 54.43
| 20231106_HepCl_ReRun_074 | CC06 | 1.825 | 1.21 | [12 empty] | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.88 | 11315.929 | 169106 | 4054.992 | 11.162 |  | 40.08
| 20231106_HepCl_ReRun_078 | CC06 | 1.825 | 1.32 | [12 empty] | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.87 | 12633.037 | 196143 | 4225.446 | 11.959 |  | 42.99
[17 empty] | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.88 | 11869.145 | 199049 | 4254.767 | 11.158 |  | 40.06
| Blanks should be less than… | 0.52913 |  | nM | [12 empty] | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.88 | 11134.584 | 184254 | 4188.782 | 10.633 |  | 38.14
| eLOQ | 1.825 |  | nM | [12 empty] | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.87 | 10594.069 | 161173 | 4496.683 | 9.424 |  | 33.72
[17 empty] | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.88 | 10354.753 | 164488 | 4153.806 | 9.971 |  | 35.72
| QC FLAGS | [15 empty] | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.87 | 9074.265 | 137502 | 3209.497 | 11.309 |  | 40.62
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.88 | 9561.769 | 153073 | 3877.044 | 9.865 |  | 35.34
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.88 | 9736.915 | 150985 | 3869.138 | 10.066 |  | 36.07
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.87 | 6488.23 | 94317 | 3451.207 | 7.52 |  | 26.76
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.88 | 4897.804 | 78547 | 3384.839 | 5.788 |  | 20.43
| Beloq eLoQ | [15 empty] | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.88 | 7423.349 | 109685 | 3411.54 | 8.704 |  | 31.09
[17 empty] | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.87 | 13458.945 | 179811 | 3946.084 | 13.643 |  | 49.15
[17 empty] | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.88 | 13394.592 | 210981 | 4015.599 | 13.343 |  | 48.05
[17 empty] | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.88 | 12774.548 | 184155 | 3722.044 | 13.729 |  | 49.46
[17 empty] | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.87 | 13906.821 | 198513 | 3719.442 | 14.956 |  | 53.95
[17 empty] | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.88 | 12754.647 | 175426 | 3898.412 | 13.087 |  | 47.12
[17 empty] | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.88 | 13496.068 | 184112 | 3840.235 | 14.058 |  | 50.66
[17 empty] | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.88 | 14466.164 | 202071 | 3796.952 | 15.24 |  | 54.99
[17 empty] | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.87 | 13676.497 | 198516 | 3988.584 | 13.716 |  | 49.41
[17 empty] | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.87 | 13672.899 | 183759 | 4089.779 | 13.373 |  | 48.16
[17 empty] | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.87 | 13819.075 | 185596 | 4257.981 | 12.982 |  | 46.73
[17 empty] | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.86 | 14415.857 | 197477 | 4107.876 | 14.037 |  | 50.59
[17 empty] | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.87 | 13589.694 | 179296 | 3864.161 | 14.067 |  | 50.7
[17 empty] | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.87 | 13366.214 | 177634 | 3543.577 | 15.088 |  | 54.43
[17 empty] | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.87 | 12986.323 | 189758 | 3555.943 | 14.608 |  | 52.68
[17 empty] | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.86 | 13166.602 | 186700 | 3988.081 | 13.206 |  | 47.55
[17 empty] | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.87 | 11126.391 | 155600 | 3506.183 | 12.693 |  | 45.68
[17 empty] | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.87 | 9472.6 | 127395 | 2791.078 | 13.576 |  | 48.9
[17 empty] | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.87 | 9704.444 | 128645 | 3117.41 | 12.452 |  | 44.79
[17 empty] | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.87 | 13138.3 | 187084 | 4056.419 | 12.956 |  | 46.64
[17 empty] | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.87 | 14517.875 | 206204 | 4559.46 | 12.736 |  | 45.83
[17 empty] | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.87 | 15841.149 | 227470 | 4493.075 | 14.103 |  | 50.83
[17 empty] | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.87 | 12003.916 | 162939 | 4114.258 | 11.671 |  | 41.94
[17 empty] | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.86 | 13315.447 | 183072 | 4004.459 | 13.301 |  | 47.9
[17 empty] | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.87 | 13054.155 | 177255 | 3985.281 | 13.102 |  | 47.17
[17 empty] | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.86 | 62.169 | 761 | 5093.202 | 0.049
[17 empty] | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte |  |  |  | 4939.262
[17 empty] | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.85 | 18.73 | 503 | 5233.792 | 0.014
[17 empty] | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.9 | 39.308 | 527 | 4518.646 | 0.035
[17 empty] | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.86 | 0.79 | 104 | 4595.066 | 0.001
[17 empty] | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.85 | 9.745 | 292 | 4516.958 | 0.009
[17 empty] | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte |  |  |  | 4562.208
[17 empty] | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.88 | 0.404 | 57 | 4277.345 | 0
[17 empty] | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.87 | 15.454 | 583 | 4028.417 | 0.015
[17 empty] | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.84 | 4.492 | 168 | 4201.701 | 0.004
[17 empty] | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.85 | 3.701 | 145 | 4522.662 | 0.003
[17 empty] | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte |  |  |  | 4864.867
[17 empty] | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte |  |  |  | 4130.82
[17 empty] | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.87 | 11.221 | 400 | 4509.017 | 0.01
[17 empty] | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.87 | 13.982 | 290 | 4377.103 | 0.013
[17 empty] | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte |  |  |  | 3747.828
[17 empty] | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte |  |  |  | 4038.739
[17 empty] | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.87 | 21.437 | 444 | 3789.513 | 0.023
[17 empty] | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.88 | 4.288 | 199 | 4403.589 | 0.004
[17 empty] | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.87 | 11.475 | 363 | 4598.22 | 0.01
[17 empty] | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.85 | 10.074 | 350 | 4689.85 | 0.009
[17 empty] | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.88 | 6.985 | 141 | 4132.267 | 0.007
[17 empty] | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.86 | 50.12 | 1129 | 4125.632 | 0.049
[17 empty] | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.88 | 5.404 | 192 | 4422.781 | 0.005
[17 empty] | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.85 | 8.917 | 245 | 4699.705 | 0.008
[17 empty] | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.88 | 44.112 | 881 | 4633.228 | 0.038
[17 empty] | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.88 | 9.015 | 235 | 4931.43 | 0.007
[17 empty] | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.85 | 10.279 | 210 | 4408.869 | 0.009
[17 empty] | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte |  |  |  | 4910.336
[17 empty] | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.85 | 10.55 | 300 | 4659.491 | 0.009
[17 empty] | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.83 | 9.293 | 285 | 3744.223 | 0.01
[17 empty] | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte |  |  |  | 4034.394
[17 empty] | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.87 | 19.84 | 431 | 4172.714 | 0.019
[17 empty] | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte |  |  |  | 3922.801
[17 empty] | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.85 | 29.131 | 366 | 3270.064 | 0.036
[17 empty] | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte |  |  |  | 4049.156
[17 empty] | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.87 | 38.099 | 1027 | 4646.139 | 0.033
[17 empty] | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte |  |  |  | 4670.868
[17 empty] | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.88 | 34.705 | 698 | 4760.106 | 0.029
[17 empty] | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.87 | 16.086 | 360 | 4087.066 | 0.016
[17 empty] | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.88 | 16.546 | 403 | 4565.343 | 0.014
[17 empty] | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.85 | 20.354 | 479 | 4533.528 | 0.018
[17 empty] | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte |  |  |  | 4314.995
[17 empty] | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.9 | 8.644 | 217 | 4169.22 | 0.008
[17 empty] | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.88 | 5.067 | 195 | 4343.747 | 0.005
[17 empty] | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte |  |  |  | 3994.895
[17 empty] | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte |  |  |  | 3869.739
[17 empty] | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte |  |  |  | 4094.355
[17 empty] | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.86 | 51.112 | 1346 | 4088.134 | 0.05
[17 empty] | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.82 | 8.796 | 183 | 3656.892 | 0.01
[17 empty] | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.9 | 3.336 | 179 | 4115.619 | 0.003
[17 empty] | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.89 | 5.356 | 204 | 3668.877 | 0.006
[17 empty] | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.9 | 4.782 | 169 | 3204.234 | 0.006
[17 empty] | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.84 | 3.179 | 97 | 3915.398 | 0.003
[17 empty] | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.84 | 3.194 | 60 | 3898.605 | 0.003
[17 empty] | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.88 | 14.696 | 311 | 4061.239 | 0.014
[17 empty] | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte |  |  |  | 3517.021
[17 empty] | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte |  |  |  | 3202.326
[17 empty] | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte |  |  |  | 3661.565
[17 empty] | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.85 | 3.96 | 193 | 2804.06 | 0.006
[17 empty] | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte |  |  |  | 4173.079
[17 empty] | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.89 | 7.057 | 258 | 3930.435 | 0.007
[17 empty] | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte |  |  |  | 3689.984
[17 empty] | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.9 | 4.366 | 159 | 3838.459 | 0.005
[17 empty] | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.85 | 2.995 | 159 | 3972.163 | 0.003
[17 empty] | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.9 | 0.645 | 36 | 3842.832 | 0.001
[17 empty] | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.87 | 0.965 | 116 | 3893.189 | 0.001
[17 empty] | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.9 | 5.351 | 219 | 3559.379 | 0.006
[17 empty] | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.84 | 2.901 | 108 | 4010.53 | 0.003
[17 empty] | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.87 | 6.052 | 191 | 4683.644 | 0.005
[17 empty] | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte |  |  |  | 3473.603
[17 empty] | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte |  |  |  | 3760.475
[17 empty] | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.85 | 3.359 | 200 | 3804.575 | 0.004
[17 empty] | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.85 | 0.614 | 56 | 3650.367 | 0.001
[17 empty] | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.82 | 3.82 | 148 | 4028.423 | 0.004
[17 empty] | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.88 | 8.494 | 187 | 3605.647 | 0.009
[17 empty] | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte |  |  |  | 3280.838
[17 empty] | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte |  |  |  | 3179.154
[17 empty] | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte |  |  |  | 4269.086
[17 empty] | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.9 | 0.276 | 37 | 4165.854 | 0
[17 empty] | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.87 | 3.55 | 130 | 4105.743 | 0.003
[17 empty] | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.91 | 5.588 | 229 | 4091.521 | 0.005
[17 empty] | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.84 | 5.883 | 249 | 3699.73 | 0.006
[17 empty] | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.88 | 3.841 | 109 | 3781.379 | 0.004
[17 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | 3.86 | 7.841 | 320
[17 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.88 | 1.966 | 103 | 2.377 | 3.308 |  | 11.37
[17 empty] | 20231106_HepCl_ReRun_003 | Blank | Blank |  |  |  | 0.237
[17 empty] | 20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 8.596
[17 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | 3.88 | 6.992 | 172
[17 empty] | 20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.04
[17 empty] | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.9 | 10.663 | 141 | 2.584 | 16.506 |  | 59.62
[17 empty] | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.87 | 5.334 | 155 | 0.567 | 37.63 |  | 136.84
[17 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank | 3.92 | 10.411 | 310
[17 empty] | 20231106_HepCl_ReRun_040 | Blank | Blank |  |  |  | 2.239
[17 empty] | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.9 | 1.849 | 69 | 1.263 | 5.856 |  | 20.68
[17 empty] | 20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 1.55
[17 empty] | 20231106_HepCl_ReRun_071 | Blank | Blank |  |  |  | 5.703
[17 empty] | 20231106_HepCl_ReRun_075 | Blank | Blank |  |  |  | 0.322
[17 empty] | 20231106_HepCl_ReRun_079 | Blank | Blank | 3.83 | 0.96 | 30 | 4.597 | 0.835 |  | 2.32
[17 empty] | 20231106_HepCl_ReRun_084 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | 3.86 | 1.076 | 86
[17 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | 3.89 | 6.497 | 168
[17 empty] | 20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 1.329
[17 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | 3.83 | 17.593 | 249
[17 empty] | 20231106_HepCl_ReRun_108 | Blank | Blank | 3.85 | 6.219 | 242
[17 empty] | 20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 3.311
[17 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank | 3.91 | 20.044 | 377
[17 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_131 | Blank | Blank |  |  |  | 1.14
[17 empty] | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.82 | 1.416 | 98 | 0.164 | 34.537 |  | 125.54
[17 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | 3.88 | 3.301 | 150
[17 empty] | 20231106_HepCl_ReRun_142 | Blank | Blank |  |  |  | 2.523
[17 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank |  |  |  | 5.162
[17 empty] | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.89 | 12.933 | 398 | 0.579 | 89.347 |  | 325.92
[17 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.83 | 4.698 | 88
[17 empty] | 20231106_HepCl_ReRun_166 | Blank | Blank |  |  |  | 0.413
[17 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank |  |  |  | 2.772
[17 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | 3.89 | 4.209 | 179
[17 empty] | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.84 | 7.722 | 217 | 1.668 | 18.518 |  | 66.97
[17 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | 3.87 | 10.259 | 349
[17 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank | 3.83 | 5.461 | 215
[17 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_200 | Blank | Blank |  |  |  | 0.525
[17 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | 3.87 | 9.578 | 198
[17 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank | 3.87 | 8.066 | 252
[17 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | 3.87 | 3.376 | 132
[17 empty] | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.83 | 4.137 | 166 | 4.215 | 3.926 |  | 13.62
[17 empty] | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.85 | 11.589 | 261 | 1.939 | 23.907 |  | 86.67
[17 empty] | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 0.26
[17 empty] | 20231106_HepCl_ReRun_246 | Blank | Blank | 3.87 | 0.829 | 74
[17 empty] | 20231106_HepCl_ReRun_253 | Blank | Blank | 3.84 | 5.148 | 173
[17 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | 3.86 | 7.01 | 252
[17 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.9 | 7.372 | 255
[17 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 3.82 | 2.362 | 85
[17 empty] | 20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.328
[17 empty] | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank |  |  |  | 8.997
[17 empty] | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.87 | 35.183 | 628 | 1.072 | 131.28 |  | 479.23
[17 empty] | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.9 | 2.746 | 126
[17 empty] | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 0.37

## 6 2 FTS
Human Active Cells | [8 empty] | Media Only | [7 empty] | 6:2 FTS
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.78 | 171.963 | 5263 | 3034.11 | 0.227 | 1.825 | 1.75 | -4.2
0 | Mix 2 Active T0 A | 72.57 | 7.0571 | 1161.12 | 1109.0667 | 50.2671 | 4.532% |  | 0 | Mix 2 Media T0 A | 63.69 | 6.9266 | 1019.04 | 1005.76 | 13.4429 | 1.337% | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.78 | 168.791 | 5259 | 2368.634 | 0.285 | 1.825 | 2.61 | 43
| Mix 2 Active T0 B | 69.08 | 7.0079 | 1105.28 | [5 empty] | Mix 2 Media T0 B | 62.88 | 6.9138 | 1006.08 |  |  |  | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.78 | 141.014 | 4907 | 2974.036 | 0.19 | 1.825 | 1.2 | -34.2
| Mix 2 Active T0 C | 66.3 | 6.9668 | 1060.8 | [5 empty] | Mix 2 Media T0 C | 62.01 | 6.8999 | 992.16 |  |  |  | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.78 | 177.583 | 5236 | 2930.851 | 0.242 | 1.825 | 1.98 | 8.4
15 | Mix 2 Active T15 A | 63.05 | 6.9165 | 1008.8 | 927.4133 | 115.0088 | 12.4% |  | 15 | Mix 2 Media T15 A | 59.26 | 6.8545 | 948.16 | 930.8267 | 44.6019 | 4.792% | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.78 | 157.768 | 4759 | 2982.801 | 0.212 | 1.825 | 1.52 | -16.5
| Mix 2 Active T15 B | 61.1 | 6.8851 | 977.6 | [5 empty] | Mix 2 Media T15 B | 55.01 | 6.7801 | 880.16 |  |  |  | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.78 | 156.739 | 4706 | 3064.16 | 0.205 | 1.825 | 1.42 | -22.1
| Mix 2 Active T15 C | 49.74 | 6.6794 | 795.84 | [5 empty] | Mix 2 Media T15 C | 60.26 | 6.8713 | 964.16 |  |  |  | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.78 | 145.769 | 5063 | 3262.376 | 0.179 | 1.825 | 1.04 | -43.1
30 | Mix 2 Active T30 A | 57.26 | 6.8202 | 916.16 | 1037.0667 | 112.4055 | 10.84% |  | 30 | Mix 2 Media T30 A | 59.67 | 6.8614 | 954.72 | 970.9333 | 63.18 | 6.507% | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.78 | 168.728 | 5698 | 3313.979 | 0.204 | 1.825 | 1.41 | -22.9
| Mix 2 Active T30 B | 66.04 | 6.9628 | 1056.64 | [5 empty] | Mix 2 Media T30 B | 57.34 | 6.8216 | 917.44 |  |  |  | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.78 | 238.272 | 7112 | 2993.711 | 0.318 | 2.919 | 3.1 | 6.3
| Mix 2 Active T30 C | 71.15 | 7.0374 | 1138.4 | [5 empty] | Mix 2 Media T30 C | 65.04 | 6.9476 | 1040.64 |  |  |  | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.78 | 193.975 | 6727 | 3350.341 | 0.232 | 2.919 | 1.82 | -37.7
60 | Mix 2 Active T60 A | 74.63 | 7.0851 | 1194.08 | 1068.96 | 109.8029 | 10.27% |  | 60 | Mix 2 Media T60 A | 53.74 | 6.7567 | 859.84 | 873.3333 | 71.7579 | 8.217% | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.78 | 347.814 | 11862 | 3101.963 | 0.449 | 4.669 | 5.03 | 7.6
| Mix 2 Active T60 B | 64.01 | 6.9316 | 1024.16 | [5 empty] | Mix 2 Media T60 B | 50.58 | 6.6961 | 809.28 |  |  |  | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.78 | 335.471 | 11180 | 2541.916 | 0.528 | 4.669 | 6.2 | 32.8
| Mix 2 Active T60 C | 61.79 | 6.8963 | 988.64 | [5 empty] | Mix 2 Media T60 C | 59.43 | 6.8574 | 950.88 |  |  |  | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.78 | 584.512 | 18989 | 3138.186 | 0.745 | 7.45 | 9.41 | 26.3
120 | Mix 2 Active T120 A | 60.17 | 6.8698 | 962.72 | 989.92 | 43.5588 | 4.4% |  | 120 | Mix 2 Media T120 A | 66.02 | 6.9625 | 1056.32 | 1003.5733 | 79.872 | 7.959% | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.78 | 445.528 | 14562 | 3151.321 | 0.566 | 7.45 | 6.75 | -9.3
| Mix 2 Active T120 B | 65.01 | 6.9471 | 1040.16 | [5 empty] | Mix 2 Media T120 B | 65.17 | 6.9496 | 1042.72 |  |  |  | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.78 | 761.471 | 26224 | 3163.276 | 0.963 | 11.875 | 12.63 | 6.3
| Mix 2 Active T120 C | 60.43 | 6.8741 | 966.88 | [5 empty] | Mix 2 Media T120 C | 56.98 | 6.8153 | 911.68 |  |  |  | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.78 | 725.711 | 25169 | 3307.122 | 0.878 | 11.875 | 11.37 | -4.3
240 | Mix 2 Active T240 A | 53.9 | 6.7597 | 862.4 | 870.9867 | 23.5262 | 2.701% |  | 240 | Mix 2 Media T240 A | 55.01 | 6.7801 | 880.16 | 934.9867 | 57.3475 | 6.134% | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.78 | 1328.398 | 45730 | 3229.261 | 1.645 | 19.125 | 22.71 | 18.8
| Mix 2 Active T240 B | 56.1 | 6.7997 | 897.6 | [5 empty] | Mix 2 Media T240 B | 62.16 | 6.9023 | 994.56 |  |  |  | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.78 | 1252.312 | 42857 | 3103.845 | 1.614 | 19.125 | 22.25 | 16.3
| Mix 2 Active T240 C | 53.31 | 6.7487 | 852.96 | [5 empty] | Mix 2 Media T240 C | 58.14 | 6.8354 | 930.24 |  |  |  | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.77 | 1914.911 | 65185 | 2767.938 | 2.767 | 30.5 | 39.29 | 28.8
[17 empty] | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.78 | 2141.835 | 73548 | 3365.248 | 2.546 | 30.5 | 36.02 | 18.1
Human Inactive Cells | [16 empty] | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.78 | 3139.108 | 108007 | 3439.733 | 3.65 | 48.75 | 52.34 | 7.4
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.78 | 3087.364 | 107125 | 3673.549 | 3.362 | 48.75 | 48.08 | -1.4
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.78 | 4940.688 | 170489 | 3919.615 | 5.042 | 78.125 | 72.91 | -6.7
0 | Mix 2 Inactive T0 A | 57.36 | 6.8219 | 917.76 | 988.32 | 129.6268 | 13.12% | [9 empty] | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.78 | 5185.91 | 177930 | 4066.942 | 5.101 | 78.125 | 73.77 | -5.6
| Mix 2 Inactive T0 B | 56.83 | 6.8127 | 909.28 | [12 empty] | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.78 | 7579.503 | 259401 | 3888.294 | 7.797 | 125 | 113.62 | -9.1
| Mix 2 Inactive T0 C | 71.12 | 7.037 | 1137.92 | [12 empty] | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.78 | 8270.189 | 287977 | 3980.426 | 8.311 | 125 | 121.21 | -3
240 | Mix 2 Inactive T240 A | 58.92 | 6.8488 | 942.72 | 920.08 | 32.0178 | 3.48% | [9 empty] | 20231106_HepCl_ReRun_042 | QC B | QC | 3.78 | 191.108 | 5682 | 3127.347 | 0.244 | 2.506 | 2.01 | -19.8
| Mix 2 Inactive T240 B | 56.09 | 6.7995 | 897.44 | [12 empty] | 20231106_HepCl_ReRun_081 | QC B | QC | 3.78 | 213.334 | 7274 | 2887.97 | 0.295 | 2.506 | 2.76 | 10.3
| Mix 2 Inactive T240 C | 60.42 | 6.8739 | 966.72 | [12 empty] | 20231106_HepCl_ReRun_043 | QC C | QC | 3.78 | 449.471 | 14963 | 3371.759 | 0.533 | 6.269 | 6.28 | 0.1
[17 empty] | 20231106_HepCl_ReRun_082 | QC C | QC | 3.78 | 390.65 | 13088 | 3610.88 | 0.433 | 6.269 | 4.79 | -23.5
| Calibration Curve | [15 empty] | 20231106_HepCl_ReRun_044 | QC D | QC | 3.78 | 1519.89 | 52572 | 3225.131 | 1.885 | 25 | 26.25 | 5
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20231106_HepCl_ReRun_083 | QC D | QC | 3.78 | 1626.875 | 56148 | 3571.287 | 1.822 | 25 | 25.32 | 1.3
| 20231106_HepCl_ReRun_013 | CC06 | 1.825 | 1.75 | 1.6162 | 0.49762 | 0.30789 | 1.564 | -0.11438 | [7 empty] | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.78 | 205.262 | 6517 | 4776.272 | 0.172 |  | 0.94
| 20231106_HepCl_ReRun_031 | CC06 | 1.825 | 2.61 | [12 empty] | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.78 | 292.673 | 9773 | 3195.141 | 0.366 |  | 3.81
| 20231106_HepCl_ReRun_035 | CC06 | 1.825 | 1.2 | [12 empty] | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.78 | 210.794 | 7030 | 4462.737 | 0.189 |  | 1.19
| 20231106_HepCl_ReRun_039 | CC06 | 1.825 | 1.98 | [12 empty] | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.78 | 144.765 | 4211 | 4447.937 | 0.13 |  | 0.32
| 20231106_HepCl_ReRun_052 | CC06 | 1.825 | 1.52 | [12 empty] | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.78 | 99.272 | 2968 | 4036.639 | 0.098
| 20231106_HepCl_ReRun_070 | CC06 | 1.825 | 1.42 | [12 empty] | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.77 | 136.658 | 4376 | 3155.708 | 0.173 |  | 0.96
| 20231106_HepCl_ReRun_074 | CC06 | 1.825 | 1.04 | [12 empty] | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.78 | 110.127 | 3275 | 4054.992 | 0.109 |  | 0
| 20231106_HepCl_ReRun_078 | CC06 | 1.825 | 1.41 | [12 empty] | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.78 | 98.905 | 2884 | 4225.446 | 0.094
[17 empty] | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.77 | 130.767 | 3718 | 4254.767 | 0.123 |  | 0.21
| Blanks should be less than… | 0.78201 |  | nM | [12 empty] | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.78 | 137.994 | 4125 | 4188.782 | 0.132 |  | 0.34
| eLOQ | 1.825 |  | nM | [12 empty] | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.78 | 140.797 | 4509 | 4496.683 | 0.125 |  | 0.25
[17 empty] | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.78 | 121.203 | 3598 | 4153.806 | 0.117 |  | 0.12
| QC FLAGS | [15 empty] | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.78 | 108.098 | 3082 | 3209.497 | 0.135 |  | 0.39
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.78 | 139.492 | 4379 | 3877.044 | 0.144 |  | 0.52
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.78 | 106.808 | 3014 | 3869.138 | 0.11 |  | 0.03
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.77 | 94.632 | 2824 | 3451.207 | 0.11 |  | 0.02
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.78 | 83.841 | 2740 | 3384.839 | 0.099
| Beloq eLoQ | [15 empty] | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.78 | 78.491 | 2813 | 3411.54 | 0.092
[17 empty] | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.78 | 372.564 | 13071 | 3946.084 | 0.378 |  | 3.98
[17 empty] | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.78 | 110.375 | 3443 | 4015.599 | 0.11 |  | 0.02
[17 empty] | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.78 | 78.054 | 2592 | 3722.044 | 0.084
[17 empty] | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.78 | 138.96 | 4052 | 3719.442 | 0.149 |  | 0.61
[17 empty] | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.78 | 126.729 | 4009 | 3898.412 | 0.13 |  | 0.32
[17 empty] | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.78 | 122.661 | 3841 | 3840.235 | 0.128 |  | 0.29
[17 empty] | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.78 | 115.056 | 3650 | 3796.952 | 0.121 |  | 0.19
[17 empty] | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.77 | 98.341 | 3529 | 3988.584 | 0.099
[17 empty] | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.78 | 149.401 | 5058 | 4089.779 | 0.146 |  | 0.56
[17 empty] | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.78 | 185.739 | 6213 | 4257.981 | 0.174 |  | 0.98
[17 empty] | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.78 | 297.333 | 9442 | 4107.876 | 0.29 |  | 2.68
[17 empty] | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.78 | 246.998 | 8342 | 3864.161 | 0.256 |  | 2.18
[17 empty] | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.78 | 86.154 | 2791 | 3543.577 | 0.097
[17 empty] | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.78 | 115.902 | 3813 | 3555.943 | 0.13 |  | 0.32
[17 empty] | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.78 | 137.6 | 4290 | 3988.081 | 0.138 |  | 0.44
[17 empty] | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.78 | 56.173 | 1957 | 3506.183 | 0.064
[17 empty] | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.77 | 80.953 | 2729 | 2791.078 | 0.116 |  | 0.11
[17 empty] | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.78 | 62.378 | 2190 | 3117.41 | 0.08
[17 empty] | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.78 | 189.653 | 5817 | 4056.419 | 0.187 |  | 1.16
[17 empty] | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.78 | 241.631 | 8170 | 4559.46 | 0.212 |  | 1.53
[17 empty] | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.78 | 290.765 | 9335 | 4493.075 | 0.259 |  | 2.22
[17 empty] | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.78 | 190.399 | 6349 | 4114.258 | 0.185 |  | 1.13
[17 empty] | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.78 | 243.066 | 7101 | 4004.459 | 0.243 |  | 1.99
[17 empty] | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.78 | 176.193 | 5558 | 3985.281 | 0.177 |  | 1.01
[17 empty] | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.78 | 6391.303 | 220071 | 5093.202 | 5.019 |  | 72.57
[17 empty] | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.78 | 5906.037 | 203264 | 4939.262 | 4.783 |  | 69.08
[17 empty] | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.78 | 6012.225 | 207327 | 5233.792 | 4.595 |  | 66.3
[17 empty] | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.78 | 4942.229 | 169810 | 4518.646 | 4.375 |  | 63.05
[17 empty] | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.78 | 4874.581 | 169161 | 4595.066 | 4.243 |  | 61.1
[17 empty] | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.78 | 3923.402 | 133112 | 4516.958 | 3.474 |  | 49.74
[17 empty] | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.78 | 4543.188 | 154753 | 4562.208 | 3.983 |  | 57.26
[17 empty] | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.78 | 4894.958 | 166429 | 4277.345 | 4.578 |  | 66.04
[17 empty] | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.78 | 4958.275 | 168396 | 4028.417 | 4.923 |  | 71.15
[17 empty] | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.78 | 5418.749 | 187049 | 4201.701 | 5.159 |  | 74.63
[17 empty] | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.78 | 5020.021 | 174032 | 4522.662 | 4.44 |  | 64.01
[17 empty] | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.78 | 5217.081 | 179277 | 4864.867 | 4.29 |  | 61.79
[17 empty] | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.78 | 4317.092 | 148947 | 4130.82 | 4.18 |  | 60.17
[17 empty] | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.78 | 5081.018 | 175617 | 4509.017 | 4.507 |  | 65.01
[17 empty] | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.78 | 4593.416 | 157865 | 4377.103 | 4.198 |  | 60.43
[17 empty] | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.78 | 3518.875 | 119857 | 3747.828 | 3.756 |  | 53.9
[17 empty] | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.78 | 3942.875 | 135332 | 4038.739 | 3.905 |  | 56.1
[17 empty] | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.78 | 3520.303 | 121928 | 3789.513 | 3.716 |  | 53.31
[17 empty] | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.78 | 4863.822 | 167123 | 4403.589 | 4.418 |  | 63.69
[17 empty] | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.78 | 5016.161 | 172989 | 4598.22 | 4.364 |  | 62.88
[17 empty] | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.78 | 5046.894 | 172993 | 4689.85 | 4.305 |  | 62.01
[17 empty] | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.78 | 4255.095 | 146051 | 4132.267 | 4.119 |  | 59.26
[17 empty] | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.78 | 3951.575 | 134733 | 4125.632 | 3.831 |  | 55.01
[17 empty] | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.78 | 4628.658 | 158950 | 4422.781 | 4.186 |  | 60.26
[17 empty] | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.78 | 4871.632 | 168775 | 4699.705 | 4.146 |  | 59.67
[17 empty] | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.78 | 4619.891 | 158711 | 4633.228 | 3.988 |  | 57.34
[17 empty] | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.78 | 5559.856 | 190760 | 4931.43 | 4.51 |  | 65.04
[17 empty] | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.78 | 4127.725 | 143210 | 4408.869 | 3.745 |  | 53.74
[17 empty] | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.78 | 4334.655 | 148483 | 4910.336 | 3.531 |  | 50.58
[17 empty] | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.78 | 4811.418 | 164882 | 4659.491 | 4.13 |  | 59.43
[17 empty] | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.78 | 4283.556 | 146551 | 3744.223 | 4.576 |  | 66.02
[17 empty] | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.78 | 4557.383 | 156709 | 4034.394 | 4.519 |  | 65.17
[17 empty] | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.78 | 4135.498 | 140199 | 4172.714 | 3.964 |  | 56.98
[17 empty] | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.78 | 3757.041 | 129864 | 3922.801 | 3.831 |  | 55.01
[17 empty] | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.78 | 3527.55 | 121269 | 3270.064 | 4.315 |  | 62.16
[17 empty] | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.78 | 4092.351 | 141067 | 4049.156 | 4.043 |  | 58.14
[17 empty] | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.78 | 4634.335 | 157508 | 4646.139 | 3.99 |  | 57.36
[17 empty] | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.78 | 4617.606 | 157457 | 4670.868 | 3.954 |  | 56.83
[17 empty] | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.78 | 5855.985 | 201916 | 4760.106 | 4.921 |  | 71.12
[17 empty] | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.78 | 4184.689 | 143212 | 4087.066 | 4.096 |  | 58.92
[17 empty] | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.78 | 4455.831 | 152479 | 4565.343 | 3.904 |  | 56.09
[17 empty] | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.78 | 4756.871 | 162191 | 4533.528 | 4.197 |  | 60.42
[17 empty] | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.78 | 65.511 | 2011 | 4314.995 | 0.061
[17 empty] | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.78 | 90.844 | 2571 | 4169.22 | 0.087
[17 empty] | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.78 | 59.138 | 1829 | 4343.747 | 0.054
[17 empty] | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.78 | 42.996 | 1394 | 3994.895 | 0.043
[17 empty] | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.78 | 39.354 | 1325 | 3869.739 | 0.041
[17 empty] | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.78 | 65.363 | 1799 | 4094.355 | 0.064
[17 empty] | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.78 | 164.218 | 5442 | 4088.134 | 0.161 |  | 0.77
[17 empty] | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.78 | 66.106 | 1966 | 3656.892 | 0.072
[17 empty] | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.77 | 67.688 | 2189 | 4115.619 | 0.066
[17 empty] | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.78 | 58.988 | 1991 | 3668.877 | 0.064
[17 empty] | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.78 | 114.769 | 3960 | 3204.234 | 0.143 |  | 0.51
[17 empty] | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.77 | 205.511 | 7108 | 3915.398 | 0.21 |  | 1.5
[17 empty] | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.78 | 229.48 | 8090 | 3898.605 | 0.235 |  | 1.88
[17 empty] | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.78 | 54.196 | 1918 | 4061.239 | 0.053
[17 empty] | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.78 | 73.433 | 2244 | 3517.021 | 0.084
[17 empty] | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.78 | 29.845 | 1106 | 3202.326 | 0.037
[17 empty] | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.78 | 51.159 | 1557 | 3661.565 | 0.056
[17 empty] | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.78 | 58.809 | 1433 | 2804.06 | 0.084
[17 empty] | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.78 | 77.001 | 2803 | 4173.079 | 0.074
[17 empty] | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.78 | 41.344 | 1442 | 3930.435 | 0.042
[17 empty] | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.77 | 39.448 | 947 | 3689.984 | 0.043
[17 empty] | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.78 | 74.386 | 1977 | 3838.459 | 0.078
[17 empty] | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.77 | 43.105 | 1522 | 3972.163 | 0.043
[17 empty] | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.78 | 58.642 | 2012 | 3842.832 | 0.061
[17 empty] | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.77 | 87.68 | 2747 | 3893.189 | 0.09
[17 empty] | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.78 | 85.112 | 3077 | 3559.379 | 0.096
[17 empty] | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.78 | 56.909 | 2038 | 4010.53 | 0.057
[17 empty] | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.78 | 6827.96 | 236140 | 4683.644 | 5.831 |  | 84.57
[17 empty] | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.78 | 217.342 | 7810 | 3473.603 | 0.25 |  | 2.1
[17 empty] | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.78 | 2704.803 | 93865 | 3760.475 | 2.877 |  | 40.91
[17 empty] | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.78 | 79.311 | 2437 | 3804.575 | 0.083
[17 empty] | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.78 | 63.602 | 1984 | 3650.367 | 0.07
[17 empty] | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.78 | 57.151 | 1930 | 4028.423 | 0.057
[17 empty] | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.78 | 34.578 | 1251 | 3605.647 | 0.038
[17 empty] | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.78 | 55.165 | 1589 | 3280.838 | 0.067
[17 empty] | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.78 | 42.574 | 1442 | 3179.154 | 0.054
[17 empty] | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.78 | 110.167 | 3794 | 4269.086 | 0.103
[17 empty] | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.78 | 102.365 | 3701 | 4165.854 | 0.098
[17 empty] | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.78 | 144.584 | 5057 | 4105.743 | 0.141 |  | 0.48
[17 empty] | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.78 | 301.788 | 9552 | 4091.521 | 0.295 |  | 2.76
[17 empty] | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.78 | 299.284 | 9875 | 3699.73 | 0.324 |  | 3.18
[17 empty] | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.78 | 306.603 | 9765 | 3781.379 | 0.324 |  | 3.19
[17 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | 3.82 | 149.304 | 2967
[17 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.81 | 70.268 | 1184 | 2.377 | 118.247 |  | 1745.78
[17 empty] | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.83 | 51.466 | 878 | 0.237 | 868.624 |  | 12834.47
[17 empty] | 20231106_HepCl_ReRun_006 | Blank | Blank | 3.81 | 46.309 | 811 | 8.596 | 21.549 |  | 316.84
[17 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | 3.81 | 31.905 | 854
[17 empty] | 20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.04
[17 empty] | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.82 | 11.18 | 423 | 2.584 | 17.307 |  | 254.14
[17 empty] | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.82 | 4.34 | 147 | 0.567 | 30.617 |  | 450.84
[17 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank | 3.82 | 14.094 | 247
[17 empty] | 20231106_HepCl_ReRun_040 | Blank | Blank |  |  |  | 2.239
[17 empty] | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 1.263
[17 empty] | 20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 1.55
[17 empty] | 20231106_HepCl_ReRun_071 | Blank | Blank |  |  |  | 5.703
[17 empty] | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.79 | 0.88 | 95 | 0.322 | 10.932 |  | 159.94
[17 empty] | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 4.597
[17 empty] | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.83 | 20.354 | 553
[17 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | 3.83 | 1.805 | 126
[17 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | 3.82 | 12.764 | 267
[17 empty] | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.83 | 17.262 | 362 | 1.329 | 51.955 |  | 766.16
[17 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | 3.82 | 6.098 | 182
[17 empty] | 20231106_HepCl_ReRun_108 | Blank | Blank | 3.82 | 2.795 | 99
[17 empty] | 20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 3.311
[17 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank | 3.82 | 10.597 | 266
[17 empty] | 20231106_HepCl_ReRun_131 | Blank | Blank |  |  |  | 1.14
[17 empty] | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.81 | 10.285 | 369 | 0.164 | 250.854 |  | 3705.38
[17 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | 3.81 | 26.678 | 650
[17 empty] | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.81 | 20.394 | 384 | 2.523 | 32.333 |  | 476.2
[17 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank |  |  |  | 5.162
[17 empty] | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.81 | 16.621 | 419 | 0.579 | 114.826 |  | 1695.23
[17 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.8 | 12.358 | 293
[17 empty] | 20231106_HepCl_ReRun_166 | Blank | Blank | 3.78 | 0.364 | 40 | 0.413 | 3.525 |  | 50.49
[17 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank | 3.82 | 10.882 | 381 | 2.772 | 15.703 |  | 230.44
[17 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | 3.82 | 11.069 | 300
[17 empty] | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.81 | 0.8 | 62 | 1.668 | 1.918 |  | 26.75
[17 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | 3.81 | 19.428 | 557
[17 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.81 | 18.713 | 586
[17 empty] | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.82 | 9.87 | 225 | 0.525 | 75.2 |  | 1109.66
[17 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank | 3.81 | 22.558 | 431
[17 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | 3.82 | 21.328 | 482
[17 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | 3.82 | 17.31 | 342
[17 empty] | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.81 | 5.699 | 139 | 4.215 | 5.408 |  | 78.32
[17 empty] | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 1.939
[17 empty] | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.82 | 13.663 | 294 | 0.26 | 210.2 |  | 3104.62
[17 empty] | 20231106_HepCl_ReRun_246 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_253 | Blank | Blank | 3.81 | 8.129 | 294
[17 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | 3.81 | 13.784 | 295
[17 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | 3.82 | 17.656 | 487
[17 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.82 | 33.317 | 899
[17 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.82 | 17.326 | 404
[17 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 3.81 | 26.536 | 537
[17 empty] | 20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.328
[17 empty] | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.81 | 29.492 | 926 | 8.997 | 13.112 |  | 192.16
[17 empty] | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.83 | 19.347 | 590 | 1.072 | 72.19 |  | 1065.19
[17 empty] | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.81 | 16.706 | 473
[17 empty] | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.81 | 13.005 | 297 | 0.37 | 140.595 |  | 2076.03

## 9H-PFNA
Human Active Cells | [8 empty] | Media Only | [7 empty] | 9H-PFNA
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.8 | 11641.195 | 382511 | 68219.305 | 0.683 | 1.825 | 1.52 | -16.5
0 | Mix 3 Active T0 A | 19.09 | 5.7218 | 305.44 | 473.6533 | 149.5377 | 31.57% |  | 0 | Mix 3 Media T0 A | 22.87 | 5.9024 | 365.92 | 456.96 | 89.7109 | 19.63% | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.81 | 10848.09 | 359772 | 65001.293 | 0.668 | 1.825 | 1.48 | -18.9
| Mix 3 Active T0 B | 32.75 | 6.2615 | 524 | [5 empty] | Mix 3 Media T0 B | 28.73 | 6.1305 | 459.68 |  |  |  | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.8 | 11571.36 | 384229 | 66638.734 | 0.695 | 1.825 | 1.56 | -14.7
| Mix 3 Active T0 C | 36.97 | 6.3827 | 591.52 | [5 empty] | Mix 3 Media T0 C | 34.08 | 6.3013 | 545.28 |  |  |  | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.81 | 11606.056 | 384908 | 61013.957 | 0.761 | 1.825 | 1.75 | -4.3
15 | Mix 3 Active T15 A | 43.77 | 6.5515 | 700.32 | 526.3467 | 150.7754 | 28.65% |  | 15 | Mix 3 Media T15 A | 36.98 | 6.383 | 591.68 | 560.1067 | 40.4143 | 7.215% | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.81 | 11166.135 | 371375 | 64340.457 | 0.694 | 1.825 | 1.56 | -14.7
| Mix 3 Active T15 B | 27.82 | 6.0983 | 445.12 | [5 empty] | Mix 3 Media T15 B | 35.88 | 6.3528 | 574.08 |  |  |  | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.81 | 11005.288 | 365567 | 63387.512 | 0.694 | 1.825 | 1.56 | -14.7
| Mix 3 Active T15 C | 27.1 | 6.0721 | 433.6 | [5 empty] | Mix 3 Media T15 C | 32.16 | 6.2433 | 514.56 |  |  |  | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.81 | 10255.913 | 336728 | 62066.836 | 0.661 | 1.825 | 1.46 | -19.9
30 | Mix 3 Active T30 A | 29.85 | 6.1688 | 477.6 | 536.1067 | 51.1015 | 9.532% |  | 30 | Mix 3 Media T30 A | 34.2 | 6.3048 | 547.2 | 635.8933 | 78.6781 | 12.37% | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.8 | 11621.744 | 390664 | 61289.246 | 0.758 | 1.825 | 1.74 | -4.7
| Mix 3 Active T30 B | 35.75 | 6.3491 | 572 | [5 empty] | Mix 3 Media T30 B | 41.45 | 6.4971 | 663.2 |  |  |  | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.81 | 21097.494 | 703071 | 64042.156 | 1.318 | 2.919 | 3.33 | 13.9
| Mix 3 Active T30 C | 34.92 | 6.3256 | 558.72 | [5 empty] | Mix 3 Media T30 C | 43.58 | 6.5472 | 697.28 |  |  |  | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.8 | 20381.063 | 685582 | 67724.766 | 1.204 | 2.919 | 3 | 2.8
60 | Mix 3 Active T60 A | 34.51 | 6.3138 | 552.16 | 559.04 | 64.8344 | 11.6% |  | 60 | Mix 3 Media T60 A | 32.78 | 6.2624 | 524.48 | 490.8267 | 41.03 | 8.359% | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.81 | 29158.406 | 967248 | 65068.902 | 1.792 | 4.669 | 4.67 | 0.1
| Mix 3 Active T60 B | 31.12 | 6.2104 | 497.92 | [5 empty] | Mix 3 Media T60 B | 27.82 | 6.0983 | 445.12 |  |  |  | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.81 | 26398.207 | 865225 | 64294.883 | 1.642 | 4.669 | 4.25 | -9.1
| Mix 3 Active T60 C | 39.19 | 6.441 | 627.04 | [5 empty] | Mix 3 Media T60 C | 31.43 | 6.2204 | 502.88 |  |  |  | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.81 | 48933.699 | 1647090 | 70109.063 | 2.792 | 7.45 | 7.51 | 0.8
120 | Mix 3 Active T120 A | 30.75 | 6.1985 | 492 | 514.4533 | 24.2443 | 4.713% |  | 120 | Mix 3 Media T120 A | 20.38 | 5.7871 | 326.08 | 441.0133 | 138.7876 | 31.47% | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.81 | 46315.258 | 1565880 | 63195.191 | 2.932 | 7.45 | 7.9 | 6.1
| Mix 3 Active T120 B | 31.95 | 6.2368 | 511.2 | [5 empty] | Mix 3 Media T120 B | 37.2 | 6.3889 | 595.2 |  |  |  | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.81 | 79207.258 | 2656081 | 69572.93 | 4.554 | 11.875 | 12.5 | 5.3
| Mix 3 Active T120 C | 33.76 | 6.2919 | 540.16 | [5 empty] | Mix 3 Media T120 C | 25.11 | 5.9959 | 401.76 |  |  |  | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.81 | 75334.328 | 2526905 | 63700.488 | 4.731 | 11.875 | 13.01 | 9.5
240 | Mix 3 Active T240 A | 29.75 | 6.1654 | 476 | 423.7867 | 52.1601 | 12.31% |  | 240 | Mix 3 Media T240 A | 22.42 | 5.8825 | 358.72 | 430.5067 | 64.116 | 14.89% | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.81 | 161645.313 | 5450768 | 72849.406 | 8.876 | 19.125 | 24.76 | 29.5
| Mix 3 Active T240 B | 26.48 | 6.049 | 423.68 | [5 empty] | Mix 3 Media T240 B | 30.13 | 6.1781 | 482.08 |  |  |  | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.81 | 152744.953 | 5169925 | 68771.055 | 8.884 | 19.125 | 24.79 | 29.6
| Mix 3 Active T240 C | 23.23 | 5.918 | 371.68 | [5 empty] | Mix 3 Media T240 C | 28.17 | 6.1108 | 450.72 |  |  |  | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.8 | 217131.609 | 7364952 | 60371.941 | 14.386 | 30.5 | 40.4 | 32.4
[17 empty] | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.81 | 215742.5 | 7243975 | 69718.367 | 12.378 | 30.5 | 34.7 | 13.8
Human Inactive Cells | [16 empty] | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.81 | 279932.938 | 9373593 | 64329.785 | 17.406 | 48.75 | 48.96 | 0.4
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.8 | 275087.375 | 9279365 | 69190.055 | 15.903 | 48.75 | 44.7 | -8.3
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.81 | 510067.344 | 17057968 | 72004.906 | 28.335 | 78.125 | 79.96 | 2.4
0 | Mix 3 Inactive T0 A | 32.85 | 6.2645 | 525.6 | 481.1733 | 81.9881 | 17.04% | [9 empty] | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.81 | 494636.938 | 16853748 | 75166.273 | 26.322 | 78.125 | 74.25 | -5
| Mix 3 Inactive T0 B | 24.16 | 5.9573 | 386.56 | [12 empty] | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.81 | 653154.375 | 22047396 | 67768.188 | 38.552 | 125 | 108.94 | -12.8
| Mix 3 Inactive T0 C | 33.21 | 6.2754 | 531.36 | [12 empty] | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.81 | 687010.063 | 23100348 | 64113.281 | 42.862 | 125 | 121.17 | -3.1
240 | Mix 3 Inactive T240 A | 32.48 | 6.2532 | 519.68 | 453.84 | 93.1118 | 20.52% | [9 empty] | 20231106_HepCl_ReRun_042 | QC B | QC | 3.81 | 22068.494 | 733400 | 76308.461 | 1.157 | 2.506 | 2.87 | 14.5
| Mix 3 Inactive T240 B | 24.25 | 5.961 | 388 | [12 empty] | 20231106_HepCl_ReRun_081 | QC B | QC | 3.8 | 19008.291 | 624404 | 84922.633 | 0.895 | 2.506 | 2.13 | -15.1
| Mix 3 Inactive T240 C | 32.21 | 6.2449 | 515.36 | [12 empty] | 20231106_HepCl_ReRun_043 | QC C | QC | 3.81 | 51954.094 | 1766232 | 84343.641 | 2.464 | 6.269 | 6.58 | 4.9
[17 empty] | 20231106_HepCl_ReRun_082 | QC C | QC | 3.81 | 56246.949 | 1905280 | 79141.828 | 2.843 | 6.269 | 7.65 | 22.1
| Calibration Curve | [15 empty] | 20231106_HepCl_ReRun_044 | QC D | QC | 3.81 | 182690.922 | 6144534 | 70742.273 | 10.33 | 25 | 28.89 | 15.6
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20231106_HepCl_ReRun_083 | QC D | QC | 3.8 | 186631.609 | 6350371 | 68663.195 | 10.872 | 25 | 30.43 | 21.7
| 20231106_HepCl_ReRun_013 | CC06 | 1.825 | 1.52 | 1.5788 | 0.10934 | 0.069258 | 0.34366 | -0.13493 | [7 empty] | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.81 | 514.675 | 8646 | 127111.25 | 0.016
| 20231106_HepCl_ReRun_031 | CC06 | 1.825 | 1.48 | [12 empty] | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.81 | 993.376 | 26476 | 107706.938 | 0.037
| 20231106_HepCl_ReRun_035 | CC06 | 1.825 | 1.56 | [12 empty] | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.81 | 642.249 | 12421 | 113437.414 | 0.023
| 20231106_HepCl_ReRun_039 | CC06 | 1.825 | 1.75 | [12 empty] | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.81 | 408.254 | 5604 | 110987.766 | 0.015
| 20231106_HepCl_ReRun_052 | CC06 | 1.825 | 1.56 | [12 empty] | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.81 | 337.916 | 5499 | 99300.031 | 0.014
| 20231106_HepCl_ReRun_070 | CC06 | 1.825 | 1.56 | [12 empty] | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.81 | 462.735 | 7215 | 105545.609 | 0.018
| 20231106_HepCl_ReRun_074 | CC06 | 1.825 | 1.46 | [12 empty] | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.81 | 453.139 | 6478 | 104039.031 | 0.017
| 20231106_HepCl_ReRun_078 | CC06 | 1.825 | 1.74 | [12 empty] | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.86 | 381.047 | 4607 | 119178.984 | 0.013
[17 empty] | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.81 | 396.53 | 5932 | 113962.063 | 0.014
| Blanks should be less than… | 0.17183 |  | nM | [12 empty] | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.81 | 322.722 | 5508 | 118969.555 | 0.011
| eLOQ | 1.825 |  | nM | [12 empty] | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.85 | 377.622 | 4486 | 122033.266 | 0.012
[17 empty] | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.85 | 390.746 | 5225 | 110901.977 | 0.014
| QC FLAGS | [15 empty] | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.81 | 412.521 | 6320 | 94312.25 | 0.017
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.8 | 408.999 | 5523 | 101210.508 | 0.016
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.81 | 356.726 | 5025 | 104400.188 | 0.014
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.81 | 404.812 | 5864 | 84769.352 | 0.019
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.85 | 333.628 | 4606 | 88387.469 | 0.015
| Beloq eLoQ | [15 empty] | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.86 | 371.009 | 4792 | 75575.227 | 0.02
[17 empty] | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.81 | 1741.671 | 49424 | 98270.578 | 0.071
[17 empty] | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.85 | 359.241 | 4577 | 98628.742 | 0.015
[17 empty] | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.86 | 327.177 | 5240 | 106868.266 | 0.012
[17 empty] | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.81 | 471.128 | 8876 | 104541.266 | 0.018
[17 empty] | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.81 | 416.801 | 6653 | 112556.063 | 0.015
[17 empty] | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.81 | 276.12 | 4768 | 104859.008 | 0.011
[17 empty] | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.81 | 414.222 | 6606 | 105383.109 | 0.016
[17 empty] | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.85 | 410.053 | 5828 | 99879.406 | 0.016
[17 empty] | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.81 | 350.653 | 7143 | 94358.172 | 0.015
[17 empty] | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.8 | 772.868 | 20550 | 116692.453 | 0.026
[17 empty] | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.81 | 947.536 | 23810 | 107368.117 | 0.035
[17 empty] | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.81 | 1016.708 | 27596 | 109868.875 | 0.037
[17 empty] | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.81 | 381.625 | 5824 | 102934.406 | 0.015
[17 empty] | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.85 | 415.396 | 5471 | 95880.813 | 0.017
[17 empty] | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.81 | 350.126 | 4576 | 111986.016 | 0.013
[17 empty] | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte |  |  |  | 84980.797
[17 empty] | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.85 | 292.296 | 3730 | 64520.02 | 0.018
[17 empty] | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.85 | 367.758 | 5330 | 77287.742 | 0.019
[17 empty] | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.81 | 457.152 | 6935 | 96807.188 | 0.019
[17 empty] | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.81 | 475.469 | 8669 | 110827.492 | 0.017
[17 empty] | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.81 | 554.069 | 8920 | 106620.406 | 0.021
[17 empty] | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.81 | 469.105 | 6650 | 86723.508 | 0.022
[17 empty] | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte |  |  |  | 97677.023
[17 empty] | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.86 | 673.771 | 8906 | 100329.82 | 0.027
[17 empty] | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.81 | 1968.833 | 59138 | 120155.836 | 0.066
[17 empty] | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.81 | 1659.804 | 49492 | 112531.641 | 0.059
[17 empty] | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.81 | 1071.008 | 26239 | 118805.617 | 0.036
[17 empty] | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.8 | 1369.747 | 38750 | 102610.383 | 0.053
[17 empty] | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.85 | 313.797 | 3798 | 102214.219 | 0.012
[17 empty] | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.81 | 374.854 | 4879 | 99535.336 | 0.015
[17 empty] | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.8 | 544.265 | 9791 | 103807.086 | 0.021
[17 empty] | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.8 | 482.408 | 8308 | 99959.609 | 0.019
[17 empty] | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.81 | 1167.067 | 33017 | 100739.109 | 0.046
[17 empty] | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.81 | 260.508 | 9475 | 95184.141 | 0.011
[17 empty] | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.81 | 3652.715 | 113771 | 106147.617 | 0.138
[17 empty] | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.81 | 1195.365 | 32755 | 98089.266 | 0.049
[17 empty] | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.81 | 339.214 | 5143 | 85510.195 | 0.016
[17 empty] | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.81 | 492.186 | 8924 | 92018.813 | 0.021
[17 empty] | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.81 | 1344.97 | 39408 | 87029.469 | 0.062
[17 empty] | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.81 | 20960.096 | 702734 | 89270.031 | 0.939 |  | 2.25
[17 empty] | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.86 | 364.35 | 4413 | 85731.852 | 0.017
[17 empty] | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.8 | 304.308 | 3837 | 82351.266 | 0.015
[17 empty] | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.8 | 331.232 | 5083 | 96676.672 | 0.014
[17 empty] | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.8 | 363.727 | 6758 | 106880.938 | 0.014
[17 empty] | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.8 | 1032.07 | 27234 | 109563.195 | 0.038
[17 empty] | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.81 | 433.407 | 8770 | 99443.953 | 0.017
[17 empty] | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.86 | 279.778 | 3623 | 106413.328 | 0.011
[17 empty] | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte |  |  |  | 109031.305
[17 empty] | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.81 | 346.74 | 5135 | 100056 | 0.014
[17 empty] | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.8 | 493.649 | 10680 | 89716.656 | 0.022
[17 empty] | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.85 | 335.928 | 4351 | 109559.086 | 0.012
[17 empty] | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.81 | 556.298 | 12640 | 90430.555 | 0.025
[17 empty] | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.81 | 745.601 | 16891 | 102643.82 | 0.029
[17 empty] | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.81 | 662.704 | 14954 | 96604.727 | 0.027
[17 empty] | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.8 | 629.296 | 15328 | 105342.867 | 0.024
[17 empty] | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.85 | 294.01 | 4664 | 98521.195 | 0.012
[17 empty] | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.81 | 312.739 | 4898 | 87807.195 | 0.014
[17 empty] | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.86 | 289.661 | 4097 | 84955.523 | 0.014
[17 empty] | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte |  |  |  | 83049.625
[17 empty] | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.85 | 285.751 | 3659 | 75152.219 | 0.015
[17 empty] | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.81 | 370.412 | 5965 | 102671.5 | 0.014
[17 empty] | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.81 | 368.678 | 5262 | 101879.82 | 0.014
[17 empty] | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.81 | 929.424 | 25504 | 119917.711 | 0.031
[17 empty] | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.81 | 397.327 | 5254 | 97705.281 | 0.016
[17 empty] | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.81 | 355.733 | 5451 | 102602.477 | 0.014
[17 empty] | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.81 | 345.67 | 4848 | 91857.102 | 0.015
[17 empty] | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.81 | 211102.141 | 6984337 | 122806.492 | 6.876 |  | 19.09
[17 empty] | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.81 | 308503.531 | 10401536 | 105540.742 | 11.692 |  | 32.75
[17 empty] | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.8 | 356524.719 | 12119818 | 108226.414 | 13.177 |  | 36.97
[17 empty] | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.81 | 384793.406 | 13078994 | 98816.094 | 15.576 |  | 43.77
[17 empty] | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.81 | 283668.125 | 9436496 | 114010.602 | 9.952 |  | 27.82
[17 empty] | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.81 | 270666.656 | 8983324 | 111625.508 | 9.699 |  | 27.1
[17 empty] | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.8 | 298650.938 | 9971744 | 111956.539 | 10.67 |  | 29.85
[17 empty] | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.81 | 333344 | 11203420 | 104588.813 | 12.749 |  | 35.75
[17 empty] | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.8 | 274019.813 | 9070789 | 87999.492 | 12.456 |  | 34.92
[17 empty] | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.81 | 299645.75 | 10020487 | 97360.305 | 12.311 |  | 34.51
[17 empty] | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.81 | 235749.125 | 7878546 | 84831.938 | 11.116 |  | 31.12
[17 empty] | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.8 | 323462.594 | 11000032 | 92662.234 | 13.963 |  | 39.19
[17 empty] | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.81 | 283991.438 | 9699914 | 103418.438 | 10.984 |  | 30.75
[17 empty] | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.81 | 303358.25 | 10312001 | 106341.625 | 11.411 |  | 31.95
[17 empty] | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.81 | 279995.094 | 9373390 | 92977.766 | 12.046 |  | 33.76
[17 empty] | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.81 | 210909.141 | 7091518 | 79338.781 | 10.633 |  | 29.75
[17 empty] | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.81 | 219247.406 | 7365930 | 92511.57 | 9.48 |  | 26.48
[17 empty] | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.81 | 177725.781 | 5986646 | 85295.422 | 8.335 |  | 23.23
[17 empty] | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.81 | 226021.344 | 7559059 | 110140.234 | 8.208 |  | 22.87
[17 empty] | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.81 | 278933.406 | 9405930 | 108602.188 | 10.274 |  | 28.73
[17 empty] | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.81 | 276028.344 | 9283592 | 90797.094 | 12.16 |  | 34.08
[17 empty] | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.81 | 367082 | 12520329 | 111401.727 | 13.18 |  | 36.98
[17 empty] | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.81 | 353533.938 | 11858321 | 110531.492 | 12.794 |  | 35.88
[17 empty] | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.81 | 323413.156 | 10911557 | 112650.75 | 11.484 |  | 32.16
[17 empty] | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.81 | 345005.031 | 11796750 | 113099.43 | 12.202 |  | 34.2
[17 empty] | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.81 | 336113.375 | 11298503 | 91099.672 | 14.758 |  | 41.45
[17 empty] | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.81 | 369345.031 | 12524720 | 95262.484 | 15.509 |  | 43.58
[17 empty] | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.81 | 278402.094 | 9377985 | 95167.969 | 11.702 |  | 32.78
[17 empty] | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.81 | 253162.359 | 8413451 | 101755.492 | 9.952 |  | 27.82
[17 empty] | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.81 | 245721.672 | 8129133 | 87558.133 | 11.226 |  | 31.43
[17 empty] | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.81 | 216459.453 | 6929789 | 118122.609 | 7.33 |  | 20.38
[17 empty] | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.81 | 338452.781 | 11647518 | 102091.367 | 13.261 |  | 37.2
[17 empty] | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.81 | 282383.156 | 9499879 | 125547.516 | 8.997 |  | 25.11
[17 empty] | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.81 | 194233.328 | 6480172 | 96524.773 | 8.049 |  | 22.42
[17 empty] | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.81 | 237021.109 | 8063289 | 88051.555 | 10.767 |  | 30.13
[17 empty] | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.81 | 243039.016 | 8258455 | 96476.289 | 10.077 |  | 28.17
[17 empty] | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.81 | 352025.063 | 11860594 | 120083.727 | 11.726 |  | 32.85
[17 empty] | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.81 | 261868.188 | 8873026 | 120913.375 | 8.663 |  | 24.16
[17 empty] | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.81 | 358415.75 | 12138892 | 120947.359 | 11.854 |  | 33.21
[17 empty] | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.81 | 284973.594 | 9606363 | 98304.648 | 11.596 |  | 32.48
[17 empty] | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.81 | 236682.141 | 7911366 | 108873.719 | 8.696 |  | 24.25
[17 empty] | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.81 | 301589.063 | 10239744 | 104880.836 | 11.502 |  | 32.21
[17 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | 3.86 | 1311.774 | 28378
[17 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.86 | 881.125 | 16293
[17 empty] | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.85 | 721.975 | 14600 | 5.359 | 538.888 |  | 1528.2
[17 empty] | 20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 13.111
[17 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | 3.85 | 443.105 | 9101
[17 empty] | 20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.806
[17 empty] | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.85 | 265.891 | 4687 | 7.869 | 135.159 |  | 382.98
[17 empty] | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.85 | 231.762 | 4646
[17 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank | 3.85 | 224.475 | 4100
[17 empty] | 20231106_HepCl_ReRun_040 | Blank | Blank | 3.85 | 223.767 | 3692
[17 empty] | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.85 | 234.037 | 4499 | 12.536 | 74.677 |  | 211.42
[17 empty] | 20231106_HepCl_ReRun_060 | Blank | Blank | 3.85 | 243.211 | 4301
[17 empty] | 20231106_HepCl_ReRun_071 | Blank | Blank | 3.86 | 178.541 | 3006 | 7.057 | 101.199 |  | 286.65
[17 empty] | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.86 | 219.784 | 4284 | 12.113 | 72.578 |  | 205.46
[17 empty] | 20231106_HepCl_ReRun_079 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.86 | 108.338 | 2639 | 2.32 | 186.79 |  | 529.44
[17 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | 3.86 | 243.378 | 5428
[17 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank |  |  |  | 3.535
[17 empty] | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.86 | 135.857 | 3418 | 4.894 | 111.04 |  | 314.56
[17 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | 3.85 | 166.996 | 3172
[17 empty] | 20231106_HepCl_ReRun_108 | Blank | Blank |  |  |  | 3.622
[17 empty] | 20231106_HepCl_ReRun_110 | Blank | Blank | 3.85 | 202.452 | 3781 | 1.239 | 653.598 |  | 1853.58
[17 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank | 3.86 | 219.288 | 4057
[17 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank | 3.85 | 189.908 | 3607
[17 empty] | 20231106_HepCl_ReRun_131 | Blank | Blank | 3.85 | 214.173 | 4079 | 14.521 | 58.997 |  | 166.94
[17 empty] | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.85 | 208.384 | 3838 | 7.49 | 111.287 |  | 315.26
[17 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | 3.85 | 363.167 | 8123
[17 empty] | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.86 | 295.344 | 5078
[17 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank | 3.85 | 306.383 | 5571 | 2.895 | 423.327 |  | 1200.4
[17 empty] | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.86 | 275.404 | 5049
[17 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.85 | 188.909 | 3291 | 4.77 | 158.414 |  | 448.95
[17 empty] | 20231106_HepCl_ReRun_166 | Blank | Blank |  |  |  | 6.376
[17 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | 3.85 | 215.745 | 4778
[17 empty] | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.86 | 188.246 | 4310 | 14.625 | 51.486 |  | 145.63
[17 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank | 3.85 | 107.023 | 2703 | 6.924 | 61.827 |  | 174.97
[17 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.86 | 207.931 | 4316 | 3.35 | 248.276 |  | 703.85
[17 empty] | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.85 | 181.452 | 3449 | 2.811 | 258.203 |  | 732.01
[17 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | 3.85 | 12.237 | 909
[17 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | 3.85 | 209.716 | 4694
[17 empty] | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.85 | 165.117 | 3099 | 9.549 | 69.166 |  | 195.78
[17 empty] | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.86 | 148.115 | 2568 | 0.334 | 1773.832 |  | 5031.24
[17 empty] | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.85 | 184.859 | 3229
[17 empty] | 20231106_HepCl_ReRun_246 | Blank | Blank | 3.86 | 213.496 | 3757 | 1.2 | 711.653 |  | 2018.26
[17 empty] | 20231106_HepCl_ReRun_253 | Blank | Blank |  |  |  | 12.938
[17 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | 3.86 | 153.336 | 3207
[17 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | 3.85 | 181.382 | 3786
[17 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.86 | 475.916 | 8865
[17 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.86 | 244.642 | 4634 | 9.577 | 102.179 |  | 289.43
[17 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 3.86 | 250.409 | 4456 | 12.974 | 77.203 |  | 218.58
[17 empty] | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.85 | 607.958 | 11669 | 0.638 | 3811.649 |  | 10811.72
[17 empty] | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.86 | 183.34 | 3807 | 8.655 | 84.733 |  | 239.94
[17 empty] | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.86 | 175.411 | 3826 | 0.454 | 1545.471 |  | 4383.47

## 8 2 FTS
Human Active Cells | [8 empty] | Media Only | [7 empty] | 8:2 FTS
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.05 | 183.985 | 6873 | 1285.358 | 0.573 | 1.825 | 2.43 | 33
0 | Mix 2 Active T0 A | 42.83 | 6.5298 | 685.28 | 682.56 | 9.8071 | 1.437% |  | 0 | Mix 2 Media T0 A | 36.66 | 6.3743 | 586.56 | 629.9733 | 80.2341 | 12.74% | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.05 | 134.609 | 4719 | 1266.854 | 0.425 | 1.825 | 1.54 | -15.5
| Mix 2 Active T0 B | 41.98 | 6.5098 | 671.68 | [5 empty] | Mix 2 Media T0 B | 36.3 | 6.3644 | 580.8 |  |  |  | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.05 | 139.414 | 5172 | 1260.427 | 0.442 | 1.825 | 1.65 | -9.8
| Mix 2 Active T0 C | 43.17 | 6.5377 | 690.72 | [5 empty] | Mix 2 Media T0 C | 45.16 | 6.5828 | 722.56 |  |  |  | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.05 | 140.109 | 4945 | 1310.875 | 0.428 | 1.825 | 1.56 | -14.7
15 | Mix 2 Active T15 A | 37.95 | 6.4089 | 607.2 | 605.5467 | 16.9406 | 2.798% |  | 15 | Mix 2 Media T15 A | 35.46 | 6.341 | 567.36 | 571.0933 | 8.9261 | 1.563% | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.05 | 146.812 | 5980 | 1179.859 | 0.498 | 1.825 | 1.98 | 8.4
| Mix 2 Active T15 B | 36.74 | 6.3765 | 587.84 | [5 empty] | Mix 2 Media T15 B | 35.29 | 6.3362 | 564.64 |  |  |  | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.05 | 120.27 | 4640 | 1220.534 | 0.394 | 1.825 | 1.36 | -25.6
| Mix 2 Active T15 C | 38.85 | 6.4323 | 621.6 | [5 empty] | Mix 2 Media T15 C | 36.33 | 6.3652 | 581.28 |  |  |  | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.05 | 143.867 | 6016 | 1181.049 | 0.487 | 1.825 | 1.92 | 5
30 | Mix 2 Active T30 A | 38.08 | 6.4123 | 609.28 | 639.1467 | 31.6498 | 4.952% |  | 30 | Mix 2 Media T30 A | 34.41 | 6.3109 | 550.56 | 571.3067 | 49.9811 | 8.749% | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.05 | 129.094 | 5605 | 1218.418 | 0.424 | 1.825 | 1.53 | -15.9
| Mix 2 Active T30 B | 39.74 | 6.4549 | 635.84 | [5 empty] | Mix 2 Media T30 B | 39.27 | 6.443 | 628.32 |  |  |  | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.05 | 221.348 | 8806 | 1359.29 | 0.651 | 2.919 | 2.9 | -0.6
| Mix 2 Active T30 C | 42.02 | 6.5107 | 672.32 | [5 empty] | Mix 2 Media T30 C | 33.44 | 6.2823 | 535.04 |  |  |  | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.05 | 193.381 | 7888 | 1307.154 | 0.592 | 2.919 | 2.54 | -12.9
60 | Mix 2 Active T60 A | 38.6 | 6.4258 | 617.6 | 691.7867 | 83.619 | 12.09% |  | 60 | Mix 2 Media T60 A | 41.45 | 6.4971 | 663.2 | 629.76 | 32.0971 | 5.097% | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.05 | 362.115 | 14411 | 1397.361 | 1.037 | 4.669 | 5.21 | 11.6
| Mix 2 Active T60 B | 42.21 | 6.5152 | 675.36 | [5 empty] | Mix 2 Media T60 B | 39.18 | 6.4408 | 626.88 |  |  |  | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.05 | 279.907 | 10543 | 1383.121 | 0.809 | 4.669 | 3.85 | -17.6
| Mix 2 Active T60 C | 48.9 | 6.6624 | 782.4 | [5 empty] | Mix 2 Media T60 C | 37.45 | 6.3956 | 599.2 |  |  |  | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.05 | 517.63 | 19667 | 1364.509 | 1.517 | 7.45 | 8.1 | 8.7
120 | Mix 2 Active T120 A | 39.37 | 6.4456 | 629.92 | 582.88 | 54.4712 | 9.345% |  | 120 | Mix 2 Media T120 A | 36.44 | 6.3683 | 583.04 | 559.7333 | 34.064 | 6.086% | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.05 | 502.933 | 19943 | 1332.81 | 1.509 | 7.45 | 8.05 | 8
| Mix 2 Active T120 B | 37.22 | 6.3894 | 595.52 | [5 empty] | Mix 2 Media T120 B | 32.54 | 6.2551 | 520.64 |  |  |  | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.05 | 764.35 | 28612 | 1367.205 | 2.236 | 11.875 | 12.41 | 4.5
| Mix 2 Active T120 C | 32.7 | 6.26 | 523.2 | [5 empty] | Mix 2 Media T120 C | 35.97 | 6.3553 | 575.52 |  |  |  | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.05 | 682.924 | 25853 | 1340.69 | 2.038 | 11.875 | 11.22 | -5.5
240 | Mix 2 Active T240 A | 33.16 | 6.2739 | 530.56 | 551.4133 | 23.3492 | 4.234% |  | 240 | Mix 2 Media T240 A | 34.3 | 6.3077 | 548.8 | 554.7733 | 22.8337 | 4.116% | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.05 | 1354.45 | 51659 | 1491.731 | 3.632 | 19.125 | 20.78 | 8.7
| Mix 2 Active T240 B | 36.04 | 6.3572 | 576.64 | [5 empty] | Mix 2 Media T240 B | 36.25 | 6.363 | 580 |  |  |  | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.05 | 1244.492 | 46411 | 1394.592 | 3.569 | 19.125 | 20.41 | 6.7
| Mix 2 Active T240 C | 34.19 | 6.3045 | 547.04 | [5 empty] | Mix 2 Media T240 C | 33.47 | 6.2832 | 535.52 |  |  |  | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.05 | 2082.749 | 78265 | 1313.037 | 6.345 | 30.5 | 37.06 | 21.5
[17 empty] | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.05 | 2029.725 | 77604 | 1381.883 | 5.875 | 30.5 | 34.24 | 12.3
Human Inactive Cells | [16 empty] | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.05 | 2907.99 | 115954 | 1391.243 | 8.361 | 48.75 | 49.15 | 0.8
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.05 | 2793.315 | 105829 | 1344.772 | 8.309 | 48.75 | 48.84 | 0.2
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.05 | 4992.377 | 192789 | 1497.735 | 13.333 | 78.125 | 78.98 | 1.1
0 | Mix 2 Inactive T0 A | 33.46 | 6.2829 | 535.36 | 598.88 | 60.9823 | 10.18% | [9 empty] | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.05 | 4922.576 | 187116 | 1523.123 | 12.928 | 78.125 | 76.55 | -2
| Mix 2 Inactive T0 B | 37.77 | 6.4041 | 604.32 | [12 empty] | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.05 | 7282.961 | 267543 | 1422.099 | 20.485 | 125 | 121.89 | -2.5
| Mix 2 Inactive T0 C | 41.06 | 6.4876 | 656.96 | [12 empty] | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.05 | 6916.329 | 258533 | 1427.197 | 19.384 | 125 | 115.29 | -7.8
240 | Mix 2 Inactive T240 A | 36.47 | 6.3691 | 583.52 | 592.16 | 12.2188 | 2.063% | [9 empty] | 20231106_HepCl_ReRun_042 | QC B | QC | 4.05 | 231.524 | 9385 | 1697.607 | 0.546 | 2.506 | 2.27 | -9.6
| Mix 2 Inactive T240 B | 37.55 | 6.3983 | 600.8 | [12 empty] | 20231106_HepCl_ReRun_081 | QC B | QC | 4.05 | 222.404 | 8528 | 1702.264 | 0.523 | 2.506 | 2.13 | -15.1
| Mix 2 Inactive T240 C | 36.39 | 6.3669 | 582.24 | [12 empty] | 20231106_HepCl_ReRun_043 | QC C | QC | 4.05 | 564.021 | 21717 | 1749.339 | 1.29 | 6.269 | 6.73 | 7.4
[17 empty] | 20231106_HepCl_ReRun_082 | QC C | QC | 4.05 | 532.346 | 19977 | 1712.841 | 1.243 | 6.269 | 6.45 | 2.9
| Calibration Curve | [15 empty] | 20231106_HepCl_ReRun_044 | QC D | QC | 4.05 | 1745.07 | 65258 | 1516.869 | 4.602 | 25 | 26.6 | 6.4
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20231106_HepCl_ReRun_083 | QC D | QC | 4.05 | 1766.013 | 67718 | 1490.915 | 4.738 | 25 | 27.42 | 9.7
| 20231106_HepCl_ReRun_013 | CC06 | 1.825 | 2.43 | 1.7462 | 0.34533 | 0.19776 | 1.0854 | -0.043151 | [7 empty] | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.05 | 455.547 | 17598 | 2610.134 | 0.698 |  | 3.18
| 20231106_HepCl_ReRun_031 | CC06 | 1.825 | 1.54 | [12 empty] | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.05 | 404.303 | 15376 | 2373.515 | 0.681 |  | 3.08
| 20231106_HepCl_ReRun_035 | CC06 | 1.825 | 1.65 | [12 empty] | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.05 | 382.016 | 15123 | 2397.312 | 0.637 |  | 2.82
| 20231106_HepCl_ReRun_039 | CC06 | 1.825 | 1.56 | [12 empty] | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.05 | 153.025 | 6504 | 2359.211 | 0.259 |  | 0.55
| 20231106_HepCl_ReRun_052 | CC06 | 1.825 | 1.98 | [12 empty] | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.05 | 408.419 | 15643 | 2233.017 | 0.732 |  | 3.38
| 20231106_HepCl_ReRun_070 | CC06 | 1.825 | 1.36 | [12 empty] | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 164.305 | 6189 | 2234.758 | 0.294 |  | 0.76
| 20231106_HepCl_ReRun_074 | CC06 | 1.825 | 1.92 | [12 empty] | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.05 | 142.847 | 6171 | 2519.11 | 0.227 |  | 0.35
| 20231106_HepCl_ReRun_078 | CC06 | 1.825 | 1.53 | [12 empty] | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.05 | 149.053 | 5522 | 2439.5 | 0.244 |  | 0.46
[17 empty] | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.05 | 261.987 | 10292 | 2235.031 | 0.469 |  | 1.81
| Blanks should be less than… | 0.54269 |  | nM | [12 empty] | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.05 | 218.786 | 8347 | 2343.231 | 0.373 |  | 1.23
| eLOQ | 1.825 |  | nM | [12 empty] | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.05 | 139.54 | 5755 | 2400.613 | 0.233 |  | 0.39
[17 empty] | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.05 | 143.163 | 5696 | 2513.042 | 0.228 |  | 0.36
| QC FLAGS | [15 empty] | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.05 | 231.241 | 9183 | 2126.96 | 0.435 |  | 1.6
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.05 | 357.515 | 13825 | 2142.216 | 0.668 |  | 3
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.05 | 167.013 | 6154 | 2372.858 | 0.282 |  | 0.68
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.05 | 315.459 | 11990 | 1960.445 | 0.644 |  | 2.85
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.05 | 136.141 | 5084 | 1928.79 | 0.282 |  | 0.69
| Beloq eLoQ | [15 empty] | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.05 | 111.834 | 4367 | 1861.049 | 0.24 |  | 0.43
[17 empty] | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.05 | 1087.664 | 41574 | 2277.866 | 1.91 |  | 10.45
[17 empty] | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.05 | 109.222 | 3935 | 1961.049 | 0.223 |  | 0.33
[17 empty] | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.05 | 134.736 | 5715 | 2366.501 | 0.228 |  | 0.36
[17 empty] | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.05 | 200.381 | 7915 | 2257.612 | 0.355 |  | 1.12
[17 empty] | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 271.229 | 10772 | 2389.278 | 0.454 |  | 1.72
[17 empty] | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.05 | 147.119 | 5598 | 2282.505 | 0.258 |  | 0.54
[17 empty] | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.05 | 167.121 | 6732 | 2296.952 | 0.291 |  | 0.74
[17 empty] | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.05 | 147.658 | 5726 | 2078.18 | 0.284 |  | 0.7
[17 empty] | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.05 | 219.107 | 8462 | 2319.462 | 0.378 |  | 1.26
[17 empty] | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.05 | 691.918 | 26391 | 2365.194 | 1.17 |  | 6.01
[17 empty] | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.05 | 669.348 | 25473 | 2378.302 | 1.126 |  | 5.75
[17 empty] | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.05 | 469.471 | 17560 | 2149.042 | 0.874 |  | 4.23
[17 empty] | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.05 | 152.881 | 6501 | 2177.948 | 0.281 |  | 0.68
[17 empty] | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.05 | 356.513 | 14002 | 1966.663 | 0.725 |  | 3.34
[17 empty] | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.05 | 221.256 | 8132 | 2330.105 | 0.38 |  | 1.27
[17 empty] | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.05 | 118.698 | 4868 | 1788.846 | 0.265 |  | 0.58
[17 empty] | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.05 | 89.453 | 3749 | 1405.372 | 0.255 |  | 0.52
[17 empty] | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.05 | 130.351 | 5363 | 1826.592 | 0.285 |  | 0.7
[17 empty] | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.05 | 207.095 | 7876 | 2059.017 | 0.402 |  | 1.41
[17 empty] | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.05 | 353.006 | 13268 | 2482.213 | 0.569 |  | 2.41
[17 empty] | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 232.939 | 8411 | 2293.903 | 0.406 |  | 1.43
[17 empty] | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.05 | 220.973 | 8680 | 2212.084 | 0.4 |  | 1.39
[17 empty] | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.05 | 205.306 | 8231 | 2292.081 | 0.358 |  | 1.14
[17 empty] | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.05 | 193.423 | 7058 | 2233.479 | 0.346 |  | 1.07
[17 empty] | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.05 | 4630.355 | 168087 | 2534.763 | 7.307 |  | 42.83
[17 empty] | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.05 | 4343.149 | 162971 | 2424.725 | 7.165 |  | 41.98
[17 empty] | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.05 | 4442.118 | 170191 | 2412.854 | 7.364 |  | 43.17
[17 empty] | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.05 | 3689.881 | 140619 | 2272.698 | 6.494 |  | 37.95
[17 empty] | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.05 | 3383.043 | 128491 | 2150.592 | 6.292 |  | 36.74
[17 empty] | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.05 | 3382.356 | 128897 | 2036.234 | 6.644 |  | 38.85
[17 empty] | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.05 | 3519.728 | 133013 | 2160.931 | 6.515 |  | 38.08
[17 empty] | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.05 | 3685.783 | 137360 | 2170.902 | 6.791 |  | 39.74
[17 empty] | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.05 | 3722.396 | 140864 | 2075.98 | 7.172 |  | 42.02
[17 empty] | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.05 | 3417.414 | 128123 | 2070.364 | 6.603 |  | 38.6
[17 empty] | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.05 | 3877.35 | 146603 | 2153.034 | 7.204 |  | 42.21
[17 empty] | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.05 | 4665.503 | 182306 | 2243.382 | 8.319 |  | 48.9
[17 empty] | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.05 | 3521.64 | 135865 | 2093.272 | 6.729 |  | 39.37
[17 empty] | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.05 | 3449.08 | 131631 | 2165.184 | 6.372 |  | 37.22
[17 empty] | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.05 | 3022.028 | 109144 | 2151.335 | 5.619 |  | 32.7
[17 empty] | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.05 | 2831.834 | 110285 | 1989.004 | 5.695 |  | 33.16
[17 empty] | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.05 | 2885.134 | 111975 | 1868.877 | 6.175 |  | 36.04
[17 empty] | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.05 | 2826.412 | 111737 | 1927.043 | 5.867 |  | 34.19
[17 empty] | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.05 | 3666.114 | 142234 | 2335.834 | 6.278 |  | 36.66
[17 empty] | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.05 | 3464.063 | 129174 | 2227.993 | 6.219 |  | 36.3
[17 empty] | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.05 | 4187.168 | 160122 | 2176.578 | 7.695 |  | 45.16
[17 empty] | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.05 | 3326.948 | 127478 | 2189.358 | 6.078 |  | 35.46
[17 empty] | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.05 | 3317.284 | 126490 | 2193.385 | 6.05 |  | 35.29
[17 empty] | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.05 | 3435.218 | 132714 | 2207.604 | 6.224 |  | 36.33
[17 empty] | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.05 | 3227.542 | 125488 | 2186.832 | 5.904 |  | 34.41
[17 empty] | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.05 | 3479.264 | 131384 | 2072.819 | 6.714 |  | 39.27
[17 empty] | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.05 | 3257.239 | 117830 | 2269.267 | 5.741 |  | 33.44
[17 empty] | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.05 | 3758.575 | 140254 | 2124.282 | 7.077 |  | 41.45
[17 empty] | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.05 | 3681.078 | 139784 | 2198.032 | 6.699 |  | 39.18
[17 empty] | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.05 | 3579.895 | 132544 | 2234.033 | 6.41 |  | 37.45
[17 empty] | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.05 | 3289.579 | 127687 | 2107.887 | 6.242 |  | 36.44
[17 empty] | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.05 | 2823.59 | 105276 | 2019.512 | 5.593 |  | 32.54
[17 empty] | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.05 | 3351.329 | 126030 | 2174.987 | 6.163 |  | 35.97
[17 empty] | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.05 | 2774.165 | 104724 | 1885.64 | 5.885 |  | 34.3
[17 empty] | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.05 | 2413.183 | 95545 | 1554.347 | 6.21 |  | 36.25
[17 empty] | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.05 | 2684.16 | 102251 | 1868.361 | 5.747 |  | 33.47
[17 empty] | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.05 | 3195.274 | 116858 | 2224.932 | 5.744 |  | 33.46
[17 empty] | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.05 | 3959.748 | 152075 | 2450.528 | 6.464 |  | 37.77
[17 empty] | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.05 | 4245.902 | 161078 | 2422.206 | 7.012 |  | 41.06
[17 empty] | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.05 | 3195.347 | 124660 | 2046.279 | 6.246 |  | 36.47
[17 empty] | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.05 | 3478.008 | 137993 | 2164.459 | 6.427 |  | 37.55
[17 empty] | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.05 | 3191.466 | 119308 | 2047.973 | 6.233 |  | 36.39
[17 empty] | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.05 | 444.989 | 16010 | 2325.502 | 0.765 |  | 3.58
[17 empty] | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.05 | 169.976 | 6440 | 2318.157 | 0.293 |  | 0.75
[17 empty] | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.05 | 77.56 | 3248 | 2446.042 | 0.127
[17 empty] | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.05 | 65.056 | 3165 | 2371.027 | 0.11
[17 empty] | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.05 | 59.29 | 2368 | 2403.863 | 0.099
[17 empty] | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.05 | 51.517 | 2109 | 2416.499 | 0.085
[17 empty] | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.05 | 134.341 | 5716 | 2279.891 | 0.236 |  | 0.41
[17 empty] | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.05 | 120.986 | 4204 | 2295.882 | 0.211 |  | 0.26
[17 empty] | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.05 | 66.306 | 2725 | 2274.982 | 0.117
[17 empty] | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.05 | 85.625 | 3468 | 1978.307 | 0.173 |  | 0.03
[17 empty] | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.05 | 98.728 | 3906 | 1924.4 | 0.205 |  | 0.22
[17 empty] | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.05 | 162.785 | 6639 | 1966.173 | 0.331 |  | 0.98
[17 empty] | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.05 | 71.583 | 3014 | 2244.717 | 0.128
[17 empty] | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.05 | 45.632 | 2059 | 2253.734 | 0.081
[17 empty] | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.05 | 68.436 | 2773 | 2225.482 | 0.123
[17 empty] | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.05 | 68.165 | 2481 | 1927.321 | 0.141
[17 empty] | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.05 | 68.579 | 2800 | 1956.9 | 0.14
[17 empty] | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.05 | 44.038 | 1395 | 1911.096 | 0.092
[17 empty] | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.05 | 86.69 | 3498 | 2183.897 | 0.159
[17 empty] | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.05 | 47.479 | 2095 | 2102.246 | 0.09
[17 empty] | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.05 | 49.284 | 2089 | 2034.973 | 0.097
[17 empty] | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.05 | 49.724 | 2144 | 2392.802 | 0.083
[17 empty] | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.05 | 54.339 | 2333 | 2360.479 | 0.092
[17 empty] | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.05 | 71.881 | 2863 | 2468 | 0.117
[17 empty] | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.05 | 64.293 | 2544 | 2278.214 | 0.113
[17 empty] | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.05 | 73.579 | 3336 | 2180.041 | 0.135
[17 empty] | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.05 | 48.467 | 1868 | 2026.516 | 0.096
[17 empty] | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.05 | 4290.784 | 170039 | 1951.082 | 8.797 |  | 51.77
[17 empty] | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.05 | 187.424 | 7368 | 1931.021 | 0.388 |  | 1.32
[17 empty] | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.05 | 1550.709 | 60551 | 1544.025 | 4.017 |  | 23.09
[17 empty] | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.05 | 63.795 | 2704 | 2285.38 | 0.112
[17 empty] | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.05 | 58.385 | 2570 | 2275.792 | 0.103
[17 empty] | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.05 | 60.019 | 2065 | 2367.659 | 0.101
[17 empty] | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.05 | 43.619 | 1630 | 2073.876 | 0.084
[17 empty] | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.05 | 49.147 | 2150 | 1820.366 | 0.108
[17 empty] | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.05 | 49.617 | 1955 | 1823.648 | 0.109
[17 empty] | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.05 | 189.058 | 7842 | 2702.162 | 0.28 |  | 0.67
[17 empty] | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.05 | 105.691 | 4814 | 2421.035 | 0.175 |  | 0.04
[17 empty] | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.05 | 149.801 | 6328 | 2636.327 | 0.227 |  | 0.36
[17 empty] | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.05 | 311.631 | 12419 | 2181.149 | 0.571 |  | 2.42
[17 empty] | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.05 | 312.361 | 11999 | 2069.581 | 0.604 |  | 2.61
[17 empty] | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.05 | 359.079 | 14109 | 2232.48 | 0.643 |  | 2.85
[17 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank |  |  |  | 2.685
[17 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.06 | 1.228 | 71 | 3.534 | 1.39 |  | 7.33
[17 empty] | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.04 | 4.574 | 189
[17 empty] | 20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 0.473
[17 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.05 | 0.796 | 47 | 1.479 | 2.153 |  | 11.91
[17 empty] | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.06 | 0.207 | 12 | 0.151 | 5.483 |  | 31.89
[17 empty] | 20231106_HepCl_ReRun_021 | Blank | Blank |  |  |  | 3.142
[17 empty] | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 0.372
[17 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank | 4.06 | 1.437 | 84 | 0.326 | 17.632 |  | 104.77
[17 empty] | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.05 | 0.854 | 50
[17 empty] | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 0.376
[17 empty] | 20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 0.214
[17 empty] | 20231106_HepCl_ReRun_071 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.06 | 0.203 | 12
[17 empty] | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 0.232
[17 empty] | 20231106_HepCl_ReRun_084 | Blank | Blank |  |  |  | 0.156
[17 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 1.338
[17 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.06 | 1.23 | 72
[17 empty] | 20231106_HepCl_ReRun_108 | Blank | Blank |  |  |  | 0.289
[17 empty] | 20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 0.254
[17 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank |  |  |  | 0.204
[17 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank |  |  |  | 0.326
[17 empty] | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.05 | 0.165 | 30 | 0.336 | 1.964 |  | 10.78
[17 empty] | 20231106_HepCl_ReRun_133 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank |  |  |  | 0.47
[17 empty] | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.02 | 0.165 | 18 | 0.21 | 3.143 |  | 17.85
[17 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.05 | 0.254 | 15
[17 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.08 | 0.303 | 18
[17 empty] | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.05 | 0.883 | 52 | 0.343 | 10.297 |  | 60.77
[17 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_182 | Blank | Blank | 4.02 | 0.146 | 9 | 0.546 | 1.07 |  | 5.41
[17 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.09 | 0.168 | 9 | 0.191 | 3.518 |  | 20.1
[17 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_200 | Blank | Blank |  |  |  | 0.632
[17 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.05 | 0.506 | 30 | 1.208 | 1.675 |  | 9.04
[17 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.04 | 0.323 | 19
[17 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank |  |  |  | 0.465
[17 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_230 | Blank | Blank |  |  |  | 0.182
[17 empty] | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 0.132
[17 empty] | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 0.507
[17 empty] | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.03 | 0.616 | 36 | 0.149 | 16.537 |  | 98.2
[17 empty] | 20231106_HepCl_ReRun_253 | Blank | Blank |  |  |  | 0.066
[17 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank |  |  |  | 0.647
[17 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.06 | 0.326 | 19 | 0.572 | 2.28 |  | 12.67
[17 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.08 | 0.433 | 25
[17 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.02 | 0.155 | 15 | 0.087 | 7.126 |  | 41.75
[17 empty] | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.05 | 0.073 | 18 | 0.348 | 0.839 |  | 4.03
[17 empty] | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.08 | 1.594 | 79 | 0.86 | 7.414 |  | 43.47
[17 empty] | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank |  |  |  | 0.598
[17 empty] | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 0.278
[17 empty] | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 3.149

## 11H-PFUnDA
Human Active Cells | [8 empty] | Media Only | [7 empty] | 11H-PFUnDA
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.06 | 1062.268 | 36328 | 40008.617 | 0.106 | 1.825 | 1.14 | -37.4
0 | Mix 3 Active T0 A | 57.55 | 6.8252 | 920.8 | 930.56 | 19.3084 | 2.075% |  | 0 | Mix 3 Media T0 A | 80.6 | 7.1621 | 1289.6 | 1164.2667 | 191.3612 | 16.44% | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.07 | 1106.86 | 42580 | 33664.371 | 0.132 | 1.825 | 1.57 | -14.2
| Mix 3 Active T0 B | 57.38 | 6.8223 | 918.08 | [5 empty] | Mix 3 Media T0 B | 59 | 6.8501 | 944 |  |  |  | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.06 | 1030.858 | 39085 | 31861.572 | 0.129 | 1.825 | 1.53 | -16.1
| Mix 3 Active T0 C | 59.55 | 6.8594 | 952.8 | [5 empty] | Mix 3 Media T0 C | 78.7 | 7.1382 | 1259.2 |  |  |  | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.07 | 1048.295 | 38562 | 30194.342 | 0.139 | 1.825 | 1.69 | -7.5
15 | Mix 3 Active T15 A | 52.44 | 6.7323 | 839.04 | 858.24 | 62.1072 | 7.237% |  | 15 | Mix 3 Media T15 A | 64.63 | 6.9413 | 1034.08 | 939.84 | 105.8727 | 11.26% | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.06 | 1061.033 | 39434 | 27641.32 | 0.154 | 1.825 | 1.93 | 5.9
| Mix 3 Active T15 B | 50.5 | 6.6946 | 808 | [5 empty] | Mix 3 Media T15 B | 51.58 | 6.7157 | 825.28 |  |  |  | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.06 | 1028.637 | 39775 | 26664.004 | 0.154 | 1.825 | 1.95 | 6.6
| Mix 3 Active T15 C | 57.98 | 6.8327 | 927.68 | [5 empty] | Mix 3 Media T15 C | 60.01 | 6.8671 | 960.16 |  |  |  | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.06 | 967.527 | 35145 | 27006.176 | 0.143 | 1.825 | 1.76 | -3.4
30 | Mix 3 Active T30 A | 46.99 | 6.6225 | 751.84 | 932 | 174.3268 | 18.7% |  | 30 | Mix 3 Media T30 A | 58.84 | 6.8474 | 941.44 | 1028.16 | 187.5347 | 18.24% | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.06 | 993.691 | 37872 | 24890.996 | 0.16 | 1.825 | 2.04 | 11.6
| Mix 3 Active T30 B | 59.02 | 6.8505 | 944.32 | [5 empty] | Mix 3 Media T30 B | 56.23 | 6.802 | 899.68 |  |  |  | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.06 | 1882.456 | 72625 | 42977.938 | 0.175 | 2.919 | 2.3 | -21.4
| Mix 3 Active T30 C | 68.74 | 7.0029 | 1099.84 | [5 empty] | Mix 3 Media T30 C | 77.71 | 7.1256 | 1243.36 |  |  |  | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.06 | 1728.86 | 67243 | 30703.615 | 0.225 | 2.919 | 3.13 | 7.3
60 | Mix 3 Active T60 A | 63.08 | 6.917 | 1009.28 | 1042.9867 | 29.4423 | 2.823% |  | 60 | Mix 3 Media T60 A | 55.17 | 6.783 | 882.72 | 1559.0933 | 591.1511 | 37.92% | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.06 | 3141.384 | 120411 | 44416.688 | 0.283 | 4.669 | 4.09 | -12.3
| Mix 3 Active T60 B | 66.48 | 6.9695 | 1063.68 | [5 empty] | Mix 3 Media T60 B | 113.6 | 7.5053 | 1817.6 |  |  |  | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.06 | 2842.049 | 102842 | 30452.596 | 0.373 | 4.669 | 5.6 | 20
| Mix 3 Active T60 C | 66 | 6.9622 | 1056 | [5 empty] | Mix 3 Media T60 C | 123.56 | 7.5893 | 1976.96 |  |  |  | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.07 | 4509.686 | 158142 | 41387.938 | 0.436 | 7.45 | 6.65 | -10.8
120 | Mix 3 Active T120 A | 58.33 | 6.8387 | 933.28 | 901.1733 | 28.8041 | 3.196% |  | 120 | Mix 3 Media T120 A | 54.62 | 6.773 | 873.92 | 926.2933 | 49.9654 | 5.394% | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.06 | 4361.453 | 168909 | 29372.057 | 0.594 | 7.45 | 9.29 | 24.7
| Mix 3 Active T120 B | 55.79 | 6.7942 | 892.64 | [5 empty] | Mix 3 Media T120 B | 60.84 | 6.8808 | 973.44 |  |  |  | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.06 | 7758.193 | 294767 | 43287.93 | 0.717 | 11.875 | 11.34 | -4.5
| Mix 3 Active T120 C | 54.85 | 6.7772 | 877.6 | [5 empty] | Mix 3 Media T120 C | 58.22 | 6.8368 | 931.52 |  |  |  | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.07 | 6099.027 | 213476 | 28786.756 | 0.847 | 11.875 | 13.52 | 13.9
240 | Mix 3 Active T240 A | 53.24 | 6.7474 | 851.84 | 854.88 | 45.5961 | 5.334% |  | 240 | Mix 3 Media T240 A | 47.5 | 6.6333 | 760 | 1013.8133 | 219.8309 | 21.68% | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.07 | 13921.18 | 525530 | 43650.934 | 1.276 | 19.125 | 20.67 | 8.1
| Mix 3 Active T240 B | 56.37 | 6.8045 | 901.92 | [5 empty] | Mix 3 Media T240 B | 71.49 | 7.0421 | 1143.84 |  |  |  | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.06 | 12133.325 | 451289 | 31747.754 | 1.529 | 19.125 | 24.9 | 30.2
| Mix 3 Active T240 C | 50.68 | 6.6981 | 810.88 | [5 empty] | Mix 3 Media T240 C | 71.1 | 7.0367 | 1137.6 |  |  |  | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.06 | 18356.199 | 719052 | 43166.441 | 1.701 | 30.5 | 27.78 | -8.9
[17 empty] | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.06 | 19046.547 | 718133 | 31395.316 | 2.427 | 30.5 | 39.89 | 30.8
Human Inactive Cells | [16 empty] | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.06 | 26155.953 | 903097 | 44292.91 | 2.362 | 48.75 | 38.82 | -20.4
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.06 | 28121.049 | 1057346 | 34492.879 | 3.261 | 48.75 | 53.83 | 10.4
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.06 | 50974.664 | 1864530 | 50112.5 | 4.069 | 78.125 | 67.32 | -13.8
0 | Mix 3 Inactive T0 A | 35.82 | 6.3511 | 573.12 | 709.1733 | 145.9672 | 20.58% | [9 empty] | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.06 | 49863.824 | 1864208 | 41397.715 | 4.818 | 78.125 | 79.83 | 2.2
| Mix 3 Inactive T0 B | 43.19 | 6.5382 | 691.04 | [12 empty] | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.07 | 71695.906 | 2683730 | 38008.148 | 7.545 | 125 | 125.37 | 0.3
| Mix 3 Inactive T0 C | 53.96 | 6.7608 | 863.36 | [12 empty] | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.06 | 70819.211 | 2617717 | 38112.547 | 7.433 | 125 | 123.49 | -1.2
240 | Mix 3 Inactive T240 A | 49.19 | 6.6683 | 787.04 | 788.72 | 2.3759 | 0.3012% | [9 empty] | 20231106_HepCl_ReRun_042 | QC B | QC | 4.06 | 1949.468 | 71825 | 51642.734 | 0.151 | 2.506 | 1.89 | -24.6
| Mix 3 Inactive T240 B | 49.4 | 6.6725 | 790.4 | [12 empty] | 20231106_HepCl_ReRun_081 | QC B | QC | 4.06 | 2080.625 | 81819 | 37154.508 | 0.224 | 2.506 | 3.11 | 24.1
| Mix 3 Inactive T240 C | 39.58 | 6.4509 | 633.28 | [12 empty] | 20231106_HepCl_ReRun_043 | QC C | QC | 4.06 | 5233.741 | 199717 | 53094.781 | 0.394 | 6.269 | 5.95 | -5
[17 empty] | 20231106_HepCl_ReRun_082 | QC C | QC | 4.06 | 4570.148 | 205965 | 41341.777 | 0.442 | 6.269 | 6.75 | 7.7
| Calibration Curve | [15 empty] | 20231106_HepCl_ReRun_044 | QC D | QC | 4.07 | 18617.66 | 686901 | 46975.297 | 1.585 | 25 | 25.84 | 3.4
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20231106_HepCl_ReRun_083 | QC D | QC | 4.06 | 15100.404 | 616116 | 32931.387 | 1.834 | 25 | 30 | 20
| 20231106_HepCl_ReRun_013 | CC06 | 1.825 | 1.14 | 1.7012 | 0.29132 | 0.17124 | 0.91563 | -0.067808 | [7 empty] | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte |  |  |  | 98002.758
| 20231106_HepCl_ReRun_031 | CC06 | 1.825 | 1.57 | [12 empty] | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte |  |  |  | 88635.797
| 20231106_HepCl_ReRun_035 | CC06 | 1.825 | 1.53 | [12 empty] | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte |  |  |  | 89190.406
| 20231106_HepCl_ReRun_039 | CC06 | 1.825 | 1.69 | [12 empty] | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte |  |  |  | 89103.305
| 20231106_HepCl_ReRun_052 | CC06 | 1.825 | 1.93 | [12 empty] | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.06 | 18.349 | 722 | 83877.867 | 0.001
| 20231106_HepCl_ReRun_070 | CC06 | 1.825 | 1.95 | [12 empty] | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 7.83 | 256 | 86799.516 | 0
| 20231106_HepCl_ReRun_074 | CC06 | 1.825 | 1.76 | [12 empty] | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.07 | 3.867 | 251 | 93797.602 | 0
| 20231106_HepCl_ReRun_078 | CC06 | 1.825 | 2.04 | [12 empty] | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.07 | 16.547 | 987 | 95630.289 | 0.001
[17 empty] | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.06 | 8.845 | 502 | 88288.648 | 0
| Blanks should be less than… | 0.45782 |  | nM | [12 empty] | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.07 | 7.013 | 311 | 88574.773 | 0
| eLOQ | 1.825 |  | nM | [12 empty] | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.07 | 9.717 | 489 | 95312.484 | 0
[17 empty] | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte |  |  |  | 94841.609
| QC FLAGS | [15 empty] | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.06 | 11.272 | 433 | 76076.914 | 0.001
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.07 | 7.606 | 361 | 76258 | 0
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.06 | 8.773 | 389 | 80176.891 | 0
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.06 | 0.389 | 70 | 65362.707 | 0
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.06 | 4.735 | 316 | 66185.563 | 0
| Beloq eLoQ | [15 empty] | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.06 | 18.365 | 455 | 64907.961 | 0.001
[17 empty] | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.06 | 64.871 | 2246 | 78984.391 | 0.003
[17 empty] | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte |  |  |  | 54687.828
[17 empty] | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.07 | 6.07 | 347 | 63527.227 | 0
[17 empty] | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.07 | 6.733 | 385 | 81925.484 | 0
[17 empty] | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 12.484 | 452 | 77283.281 | 0.001
[17 empty] | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.06 | 2.375 | 131 | 77620.695 | 0
[17 empty] | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.06 | 4.094 | 252 | 86903.203 | 0
[17 empty] | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte |  |  |  | 64245.207
[17 empty] | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.06 | 20.189 | 1028 | 84458.141 | 0.001
[17 empty] | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.06 | 68.218 | 3260 | 87830.766 | 0.003
[17 empty] | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.07 | 17.151 | 687 | 82120.75 | 0.001
[17 empty] | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.07 | 17.723 | 611 | 76499.594 | 0.001
[17 empty] | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte |  |  |  | 73346.039
[17 empty] | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.06 | 14.537 | 793 | 68280.641 | 0.001
[17 empty] | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.06 | 11.696 | 573 | 73499.289 | 0.001
[17 empty] | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte |  |  |  | 55749.684
[17 empty] | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte |  |  |  | 30788.695
[17 empty] | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.07 | 7.596 | 468 | 54407.445 | 0.001
[17 empty] | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte |  |  |  | 76163.773
[17 empty] | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.07 | 1.967 | 217 | 92251.773 | 0
[17 empty] | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 4.014 | 281 | 83986.945 | 0
[17 empty] | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.06 | 10.213 | 486 | 86446.32 | 0
[17 empty] | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte |  |  |  | 89926.844
[17 empty] | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.09 | 16.609 | 464 | 91128.219 | 0.001
[17 empty] | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.06 | 81.691 | 3339 | 107375.094 | 0.003
[17 empty] | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.06 | 71.185 | 2151 | 104266.414 | 0.003
[17 empty] | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.06 | 105.631 | 3989 | 105538.766 | 0.004
[17 empty] | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.07 | 57.227 | 2282 | 84474.836 | 0.003
[17 empty] | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.07 | 6.926 | 306 | 87116.578 | 0
[17 empty] | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.06 | 3.727 | 259 | 83468.359 | 0
[17 empty] | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.06 | 12.387 | 330 | 89927.414 | 0.001
[17 empty] | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.06 | 10.963 | 199 | 86421.602 | 0.001
[17 empty] | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.07 | 13.604 | 430 | 91154.789 | 0.001
[17 empty] | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.03 | 16.884 | 380 | 89026.438 | 0.001
[17 empty] | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.06 | 13.921 | 612 | 89150.375 | 0.001
[17 empty] | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.07 | 29.097 | 1028 | 89808.383 | 0.001
[17 empty] | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.06 | 5.17 | 309 | 78967.586 | 0
[17 empty] | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.03 | 0.403 | 62 | 83951.203 | 0
[17 empty] | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.06 | 68.997 | 2377 | 82722.727 | 0.003
[17 empty] | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.06 | 0.798 | 93 | 80214.852 | 0
[17 empty] | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte |  |  |  | 77124.359
[17 empty] | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.07 | 2.654 | 165 | 74714.813 | 0
[17 empty] | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.06 | 13.18 | 449 | 99319.828 | 0.001
[17 empty] | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte |  |  |  | 94711.563
[17 empty] | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.06 | 7.088 | 331 | 90822.211 | 0
[17 empty] | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.07 | 5.175 | 302 | 84107.078 | 0
[17 empty] | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte |  |  |  | 82892.469
[17 empty] | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte |  |  |  | 89697.148
[17 empty] | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.07 | 8.217 | 261 | 84959.609 | 0
[17 empty] | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte |  |  |  | 86283.852
[17 empty] | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.07 | 13.263 | 417 | 100046.359 | 0.001
[17 empty] | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.07 | 7.463 | 338 | 95674.383 | 0
[17 empty] | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte |  |  |  | 93804.273
[17 empty] | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.11 | 1.75 | 98 | 75229.672 | 0
[17 empty] | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.06 | 38.406 | 1610 | 87124.805 | 0.002
[17 empty] | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.02 | 2.931 | 191 | 77005.422 | 0
[17 empty] | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.08 | 9.604 | 240 | 88369.641 | 0
[17 empty] | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.06 | 0.168 | 31 | 76210.016 | 0
[17 empty] | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.09 | 0.105 | 18 | 52721.07 | 0
[17 empty] | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.07 | 9.736 | 352 | 55817.023 | 0.001
[17 empty] | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte |  |  |  | 99776.836
[17 empty] | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.06 | 55.627 | 2127 | 101979.219 | 0.002
[17 empty] | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.07 | 3.523 | 218 | 114534.242 | 0
[17 empty] | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.06 | 8.106 | 373 | 79844.398 | 0
[17 empty] | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.06 | 9.177 | 373 | 86026.953 | 0
[17 empty] | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.06 | 3.263 | 202 | 90566.922 | 0
[17 empty] | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.07 | 84765.813 | 3149459 | 97325.922 | 3.484 |  | 57.55
[17 empty] | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.06 | 89158.578 | 3440956 | 102662.727 | 3.474 |  | 57.38
[17 empty] | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.06 | 89658.922 | 3423202 | 99524.109 | 3.604 |  | 59.55
[17 empty] | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.06 | 83136.398 | 3056431 | 104638.984 | 3.178 |  | 52.44
[17 empty] | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.06 | 73273.953 | 2781368 | 95722.523 | 3.062 |  | 50.5
[17 empty] | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.07 | 91674.398 | 3524289 | 104475.281 | 3.51 |  | 57.98
[17 empty] | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.06 | 69076.906 | 2644574 | 96895.539 | 2.852 |  | 46.99
[17 empty] | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.06 | 80887.133 | 3075320 | 90586.703 | 3.572 |  | 59.02
[17 empty] | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.06 | 88501.672 | 3452261 | 85220.258 | 4.154 |  | 68.74
[17 empty] | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.06 | 83890.813 | 3278919 | 87958.188 | 3.815 |  | 63.08
[17 empty] | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.06 | 74294.75 | 2865533 | 73946.117 | 4.019 |  | 66.48
[17 empty] | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.06 | 88411.766 | 3434985 | 88629.414 | 3.99 |  | 66
[17 empty] | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.06 | 82879.844 | 2973760 | 93904.141 | 3.53 |  | 58.33
[17 empty] | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.06 | 83693.422 | 3138311 | 99092.422 | 3.378 |  | 55.79
[17 empty] | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.07 | 85406.219 | 3482632 | 102824.828 | 3.322 |  | 54.85
[17 empty] | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.07 | 60181.426 | 2252242 | 74625.852 | 3.226 |  | 53.24
[17 empty] | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.07 | 66792.984 | 2494126 | 78271.125 | 3.413 |  | 56.37
[17 empty] | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.07 | 63163.129 | 2425594 | 82232.203 | 3.072 |  | 50.68
[17 empty] | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.06 | 83874.047 | 3258557 | 68976.289 | 4.864 |  | 80.6
[17 empty] | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.07 | 72234.156 | 2659945 | 80912.125 | 3.571 |  | 59
[17 empty] | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.06 | 76960.344 | 2921204 | 64799.211 | 4.751 |  | 78.7
[17 empty] | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.07 | 93146.758 | 3627092 | 95344.883 | 3.908 |  | 64.63
[17 empty] | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.07 | 77323.891 | 2759422 | 98936.914 | 3.126 |  | 51.58
[17 empty] | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.06 | 88055.625 | 3283423 | 96992.188 | 3.631 |  | 60.01
[17 empty] | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.07 | 96175.375 | 3901113 | 108030.414 | 3.561 |  | 58.84
[17 empty] | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.06 | 85629.828 | 3304683 | 100591.141 | 3.405 |  | 56.23
[17 empty] | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.07 | 81606.148 | 3017306 | 69586.211 | 4.691 |  | 77.71
[17 empty] | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.06 | 87272.586 | 3324550 | 104467.281 | 3.342 |  | 55.17
[17 empty] | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.06 | 83227.961 | 3148725 | 48669.789 | 6.84 |  | 113.6
[17 empty] | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.06 | 73634.305 | 2678980 | 39606.652 | 7.437 |  | 123.56
[17 empty] | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.06 | 82758.57 | 3205700 | 100054.281 | 3.309 |  | 54.62
[17 empty] | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.07 | 84048.07 | 3297586 | 91328.695 | 3.681 |  | 60.84
[17 empty] | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.07 | 90018.523 | 3535425 | 102183.391 | 3.524 |  | 58.22
[17 empty] | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.07 | 68525.758 | 2553414 | 95108.883 | 2.882 |  | 47.5
[17 empty] | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.07 | 58218.891 | 2193878 | 53921.77 | 4.319 |  | 71.49
[17 empty] | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.07 | 65500.477 | 2518845 | 61000.227 | 4.295 |  | 71.1
[17 empty] | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.06 | 98320.117 | 3693365 | 180164.406 | 2.183 |  | 35.82
[17 empty] | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.07 | 84453.852 | 3030661 | 128731.063 | 2.624 |  | 43.19
[17 empty] | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.06 | 98009.695 | 3569567 | 119917.117 | 3.269 |  | 53.96
[17 empty] | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.06 | 66279.32 | 2477570 | 88866.891 | 2.983 |  | 49.19
[17 empty] | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.06 | 66022.867 | 2635104 | 88157.313 | 2.996 |  | 49.4
[17 empty] | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.06 | 59553.82 | 2256071 | 98938.063 | 2.408 |  | 39.58
[17 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.04 | 5.602 | 132
[17 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.05 | 4.978 | 128
[17 empty] | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.07 | 0.235 | 18 | 6.982 | 0.135 |  | 1.62
[17 empty] | 20231106_HepCl_ReRun_006 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.03 | 2.4 | 141 | 7.306 | 1.314 |  | 21.31
[17 empty] | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.05 | 2.265 | 131 | 0.639 | 14.178 |  | 236.14
[17 empty] | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.09 | 6.331 | 205
[17 empty] | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 4.114
[17 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank |  |  |  | 7.725
[17 empty] | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.07 | 5.768 | 153 | 5.681 | 4.061 |  | 67.19
[17 empty] | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 3.035
[17 empty] | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.06 | 11.194 | 371 | 2.979 | 15.031 |  | 250.38
[17 empty] | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.04 | 0.246 | 35 | 13.529 | 0.073 |  | 0.58
[17 empty] | 20231106_HepCl_ReRun_075 | Blank | Blank |  |  |  | 8.646
[17 empty] | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 10.156
[17 empty] | 20231106_HepCl_ReRun_084 | Blank | Blank | 4.06 | 4.429 | 252 | 4.803 | 3.689 |  | 60.97
[17 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.06 | 5.79 | 191 | 0.242 | 95.702 |  | 1597.58
[17 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | 4.09 | 9.932 | 215
[17 empty] | 20231106_HepCl_ReRun_094 | Blank | Blank | 4.08 | 2.282 | 84 | 0.119 | 76.706 |  | 1280.34
[17 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.04 | 13.821 | 319 | 1.22 | 45.315 |  | 756.11
[17 empty] | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.07 | 2.447 | 167 | 7.197 | 1.36 |  | 22.08
[17 empty] | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.06 | 3.336 | 134 | 1.133 | 11.778 |  | 196.05
[17 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.02 | 3.587 | 216 | 1.785 | 8.038 |  | 133.6
[17 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank | 4.06 | 3.813 | 215 | 8.867 | 1.72 |  | 28.09
[17 empty] | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.06 | 0.554 | 114 | 0.21 | 10.552 |  | 175.59
[17 empty] | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.09 | 2.798 | 175 | 1.004 | 11.147 |  | 185.53
[17 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.06 | 0.123 | 16 | 3.553 | 0.138 |  | 1.68
[17 empty] | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.1 | 3.614 | 178 | 4.462 | 3.24 |  | 53.47
[17 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.05 | 4.853 | 276 | 6.476 | 2.998 |  | 49.43
[17 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.02 | 4.052 | 238 | 0.75 | 21.611 |  | 360.26
[17 empty] | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.04 | 8.785 | 356 | 3.511 | 10.009 |  | 166.51
[17 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.02 | 0.679 | 63 | 6.395 | 0.425 |  | 6.46
[17 empty] | 20231106_HepCl_ReRun_182 | Blank | Blank |  |  |  | 0.43
[17 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.1 | 4.923 | 175 | 2.904 | 6.781 |  | 112.61
[17 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.04 | 7.196 | 210 | 2.254 | 12.77 |  | 212.63
[17 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank |  |  |  | 0.96
[17 empty] | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.04 | 2.128 | 95 | 1.102 | 7.724 |  | 128.36
[17 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.1 | 3.972 | 181 | 1.83 | 8.682 |  | 144.36
[17 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.07 | 12.603 | 469 | 1.23 | 40.985 |  | 683.81
[17 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.02 | 10.347 | 233 | 0.419 | 98.778 |  | 1648.94
[17 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | 4.05 | 3.697 | 187 | 5.807 | 2.547 |  | 41.9
[17 empty] | 20231106_HepCl_ReRun_230 | Blank | Blank | 4.08 | 13.86 | 401 | 2.226 | 24.906 |  | 415.29
[17 empty] | 20231106_HepCl_ReRun_237 | Blank | Blank | 4.06 | 2.31 | 145 | 8.198 | 1.127 |  | 18.19
[17 empty] | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 7.8
[17 empty] | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.08 | 0.727 | 65 | 4.036 | 0.721 |  | 11.4
[17 empty] | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.06 | 0.597 | 60 | 1.238 | 1.929 |  | 31.58
[17 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | 4.08 | 4.24 | 210 | 7.609 | 2.229 |  | 36.59
[17 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.01 | 0.367 | 74 | 3 | 0.489 |  | 7.54
[17 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.01 | 0.655 | 65 | 3.127 | 0.838 |  | 13.36
[17 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.03 | 3.303 | 150 | 4.24 | 3.116 |  | 51.41
[17 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.04 | 3.32 | 197 | 1.228 | 10.814 |  | 179.97
[17 empty] | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.03 | 3.021 | 145 | 1.469 | 8.226 |  | 136.74
[17 empty] | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.07 | 1.874 | 69 | 5.073 | 1.478 |  | 24.05
[17 empty] | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.07 | 4.099 | 215 | 16.466 | 0.996 |  | 16
[17 empty] | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 9.694
[17 empty] | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 6.959

## PFUnDA
Human Active Cells | [8 empty] | Media Only | [7 empty] | PFUnDA
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.4 | 7043 | 277350 | 40008.617 | 0.704 | 1.825 | 1.9 | 4.1
0 | Mix 1 Active T0 A | 40.4 | 6.4714 | 646.4 | 613.28 | 39.474 | 6.437% |  | 0 | Mix 1 Media T0 A | 39.62 | 6.4519 | 633.92 | 602.6133 | 27.6371 | 4.586% | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.4 | 5588.62 | 236229 | 33664.371 | 0.664 | 1.825 | 1.76 | -3.5
| Mix 1 Active T0 B | 38.99 | 6.4359 | 623.84 | [5 empty] | Mix 1 Media T0 B | 36.35 | 6.3658 | 581.6 |  |  |  | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.4 | 5433.467 | 207631 | 31861.572 | 0.682 | 1.825 | 1.82 | -0.1
| Mix 1 Active T0 C | 35.6 | 6.3449 | 569.6 | [5 empty] | Mix 1 Media T0 C | 37.02 | 6.384 | 592.32 |  |  |  | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.4 | 5458.212 | 220757 | 30194.342 | 0.723 | 1.825 | 1.97 | 7.7
15 | Mix 1 Active T15 A | 41.09 | 6.4884 | 657.44 | 637.12 | 28.8013 | 4.521% |  | 15 | Mix 1 Media T15 A | 38.83 | 6.4318 | 621.28 | 648.64 | 28.4571 | 4.387% | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.4 | 4747.757 | 201851 | 27641.32 | 0.687 | 1.825 | 1.84 | 0.9
| Mix 1 Active T15 B | 37.76 | 6.4038 | 604.16 | [5 empty] | Mix 1 Media T15 B | 42.38 | 6.5193 | 678.08 |  |  |  | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.4 | 4767.315 | 194365 | 26664.004 | 0.715 | 1.825 | 1.94 | 6.2
| Mix 1 Active T15 C | 40.61 | 6.4766 | 649.76 | [5 empty] | Mix 1 Media T15 C | 40.41 | 6.4717 | 646.56 |  |  |  | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.4 | 4664.719 | 193278 | 27006.176 | 0.691 | 1.825 | 1.85 | 1.6
30 | Mix 1 Active T30 A | 40.81 | 6.4815 | 652.96 | 661.12 | 18.5951 | 2.813% |  | 30 | Mix 1 Media T30 A | 40.62 | 6.4768 | 649.92 | 633.92 | 18.8637 | 2.976% | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.39 | 4229.273 | 162426 | 24890.996 | 0.68 | 1.825 | 1.82 | -0.5
| Mix 1 Active T30 B | 40.5 | 6.4739 | 648 | [5 empty] | Mix 1 Media T30 B | 38.32 | 6.4186 | 613.12 |  |  |  | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.4 | 10383.962 | 423591 | 42977.938 | 0.966 | 2.919 | 2.81 | -3.7
| Mix 1 Active T30 C | 42.65 | 6.5256 | 682.4 | [5 empty] | Mix 1 Media T30 C | 39.92 | 6.4595 | 638.72 |  |  |  | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.39 | 6970.027 | 276378 | 30703.615 | 0.908 | 2.919 | 2.61 | -10.7
60 | Mix 1 Active T60 A | 43.34 | 6.5417 | 693.44 | 692.48 | 1.6628 | 0.2401% |  | 60 | Mix 1 Media T60 A | 43.41 | 6.5433 | 694.56 | 711.52 | 21.8375 | 3.069% | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.4 | 16841.48 | 684362 | 44416.688 | 1.517 | 4.669 | 4.72 | 1.1
| Mix 1 Active T60 B | 43.16 | 6.5375 | 690.56 | [5 empty] | Mix 1 Media T60 B | 46.01 | 6.6014 | 736.16 |  |  |  | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.4 | 11831.235 | 485726 | 30452.596 | 1.554 | 4.669 | 4.85 | 3.9
| Mix 1 Active T60 C | 43.34 | 6.5417 | 693.44 | [5 empty] | Mix 1 Media T60 C | 43.99 | 6.5566 | 703.84 |  |  |  | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.4 | 23083.746 | 919942 | 41387.938 | 2.231 | 7.45 | 7.2 | -3.4
120 | Mix 1 Active T120 A | 37.05 | 6.3849 | 592.8 | 600.6933 | 14.0894 | 2.346% |  | 120 | Mix 1 Media T120 A | 37.13 | 6.387 | 594.08 | 615.4133 | 18.6744 | 3.034% | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.4 | 16286.906 | 675195 | 29372.057 | 2.218 | 7.45 | 7.16 | -4
| Mix 1 Active T120 B | 37.02 | 6.384 | 592.32 | [5 empty] | Mix 1 Media T120 B | 39.3 | 6.4438 | 628.8 |  |  |  | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.4 | 36668.535 | 1499343 | 43287.93 | 3.388 | 11.875 | 11.22 | -5.5
| Mix 1 Active T120 C | 38.56 | 6.4248 | 616.96 | [5 empty] | Mix 1 Media T120 C | 38.96 | 6.4351 | 623.36 |  |  |  | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.4 | 23608.922 | 980457 | 28786.756 | 3.281 | 11.875 | 10.84 | -8.7
240 | Mix 1 Active T240 A | 39.77 | 6.4557 | 636.32 | 631.2533 | 7.563 | 1.198% |  | 240 | Mix 1 Media T240 A | 40.22 | 6.467 | 643.52 | 615.4133 | 32.8099 | 5.331% | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.4 | 63299.324 | 2670024 | 43650.934 | 5.801 | 19.125 | 19.59 | 2.4
| Mix 1 Active T240 B | 38.91 | 6.4338 | 622.56 | [5 empty] | Mix 1 Media T240 B | 38.96 | 6.4351 | 623.36 |  |  |  | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.4 | 45632.473 | 1810515 | 31747.754 | 5.749 | 19.125 | 19.41 | 1.5
| Mix 1 Active T240 C | 39.68 | 6.4534 | 634.88 | [5 empty] | Mix 1 Media T240 C | 36.21 | 6.3619 | 579.36 |  |  |  | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.39 | 110621.594 | 4408315 | 43166.441 | 10.251 | 30.5 | 35.04 | 14.9
[17 empty] | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.4 | 81080.563 | 3271567 | 31395.316 | 10.33 | 30.5 | 35.31 | 15.8
Human Inactive Cells | [16 empty] | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.4 | 142737.063 | 5848686 | 44292.91 | 12.89 | 48.75 | 44.2 | -9.3
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.4 | 110454.477 | 4597438 | 34492.879 | 12.809 | 48.75 | 43.92 | -9.9
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.4 | 284172 | 11155492 | 50112.5 | 22.683 | 78.125 | 78.19 | 0.1
0 | Mix 1 Inactive T0 A | 35.57 | 6.3441 | 569.12 | 570.5067 | 2.6837 | 0.4704% | [9 empty] | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.4 | 222048.391 | 9055133 | 41397.715 | 21.455 | 78.125 | 73.93 | -5.4
| Mix 1 Inactive T0 B | 35.85 | 6.3519 | 573.6 | [12 empty] | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.4 | 376289.313 | 15172176 | 38008.148 | 39.601 | 125 | 136.92 | 9.5
| Mix 1 Inactive T0 C | 35.55 | 6.3435 | 568.8 | [12 empty] | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.4 | 327021.594 | 13200246 | 38112.547 | 34.322 | 125 | 118.6 | -5.1
240 | Mix 1 Inactive T240 A | 37.73 | 6.403 | 603.68 | 622.32 | 26.3609 | 4.236% | [9 empty] | 20231106_HepCl_ReRun_042 | QC B | QC | 4.4 | 10822.062 | 445443 | 51642.734 | 0.838 | 2.506 | 2.37 | -5.6
| Mix 1 Inactive T240 B | 40.06 | 6.463 | 640.96 | [12 empty] | 20231106_HepCl_ReRun_081 | QC B | QC | 4.4 | 7522.85 | 298162 | 37154.508 | 0.81 | 2.506 | 2.27 | -9.5
| Mix 1 Inactive T240 C | 40.06 | 6.463 | 640.96 | [12 empty] | 20231106_HepCl_ReRun_043 | QC C | QC | 4.4 | 26934.248 | 1136498 | 53094.781 | 2.029 | 6.269 | 6.5 | 3.7
[17 empty] | 20231106_HepCl_ReRun_082 | QC C | QC | 4.4 | 21356.891 | 869820 | 41341.777 | 2.066 | 6.269 | 6.63 | 5.7
| Calibration Curve | [15 empty] | 20231106_HepCl_ReRun_044 | QC D | QC | 4.4 | 85835.969 | 3549459 | 46975.297 | 7.309 | 25 | 24.83 | -0.7
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20231106_HepCl_ReRun_083 | QC D | QC | 4.4 | 60284.25 | 2577048 | 32931.387 | 7.322 | 25 | 24.87 | -0.5
| 20231106_HepCl_ReRun_013 | CC06 | 1.825 | 1.9 | 1.8625 | 0.069437 | 0.037281 | 0.21824 | 0.020548 | [7 empty] | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.4 | 288979.969 | 11903002 | 98002.758 | 11.795 |  | 40.4
| 20231106_HepCl_ReRun_031 | CC06 | 1.825 | 1.76 | [12 empty] | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.4 | 252387.266 | 10245247 | 88635.797 | 11.39 |  | 38.99
| 20231106_HepCl_ReRun_035 | CC06 | 1.825 | 1.82 | [12 empty] | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.4 | 232170.688 | 9274514 | 89190.406 | 10.412 |  | 35.6
| 20231106_HepCl_ReRun_039 | CC06 | 1.825 | 1.97 | [12 empty] | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.4 | 267149.969 | 10917079 | 89103.305 | 11.993 |  | 41.09
| 20231106_HepCl_ReRun_052 | CC06 | 1.825 | 1.84 | [12 empty] | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.4 | 231363.469 | 9350834 | 83877.867 | 11.033 |  | 37.76
| 20231106_HepCl_ReRun_070 | CC06 | 1.825 | 1.94 | [12 empty] | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.39 | 257262.188 | 10705604 | 86799.516 | 11.855 |  | 40.61
| 20231106_HepCl_ReRun_074 | CC06 | 1.825 | 1.85 | [12 empty] | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.4 | 279341.906 | 11378870 | 93797.602 | 11.913 |  | 40.81
| 20231106_HepCl_ReRun_078 | CC06 | 1.825 | 1.82 | [12 empty] | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.4 | 282657.531 | 11606108 | 95630.289 | 11.823 |  | 40.5
[17 empty] | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.4 | 274644.281 | 11201561 | 88288.648 | 12.443 |  | 42.65
| Blanks should be less than… | 0.10912 |  | nM | [12 empty] | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.4 | 279947.906 | 11400369 | 88574.773 | 12.642 |  | 43.34
| eLOQ | 1.825 |  | nM | [12 empty] | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.4 | 300004.281 | 12091701 | 95312.484 | 12.59 |  | 43.16
[17 empty] | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.4 | 299780.094 | 11869348 | 94841.609 | 12.643 |  | 43.34
| QC FLAGS | [15 empty] | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.4 | 205967.813 | 8300433 | 76076.914 | 10.829 |  | 37.05
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.4 | 206321.797 | 8091617 | 76258 | 10.822 |  | 37.02
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.4 | 225792.75 | 9374209 | 80176.891 | 11.265 |  | 38.56
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.4 | 189770.188 | 7714787 | 65362.707 | 11.613 |  | 39.77
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.4 | 188079.766 | 7687524 | 66185.563 | 11.367 |  | 38.91
| Beloq eLoQ | [15 empty] | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.4 | 188022.875 | 7440712 | 64907.961 | 11.587 |  | 39.68
[17 empty] | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.4 | 228470.469 | 9286333 | 78984.391 | 11.57 |  | 39.62
[17 empty] | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.4 | 145300.688 | 5942828 | 54687.828 | 10.628 |  | 36.35
[17 empty] | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.4 | 171869.766 | 6608780 | 63527.227 | 10.822 |  | 37.02
[17 empty] | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.4 | 232343.172 | 9388697 | 81925.484 | 11.344 |  | 38.83
[17 empty] | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.4 | 238910.734 | 9909917 | 77283.281 | 12.365 |  | 42.38
[17 empty] | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.4 | 228967.797 | 9634266 | 77620.695 | 11.799 |  | 40.41
[17 empty] | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.4 | 257647.172 | 10769629 | 86903.203 | 11.859 |  | 40.62
[17 empty] | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.4 | 179824.234 | 6961351 | 64245.207 | 11.196 |  | 38.32
[17 empty] | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.4 | 246133.094 | 10120721 | 84458.141 | 11.657 |  | 39.92
[17 empty] | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.4 | 278046.969 | 11063472 | 87830.766 | 12.663 |  | 43.41
[17 empty] | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.4 | 275360.031 | 10737526 | 82120.75 | 13.412 |  | 46.01
[17 empty] | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.4 | 245384.563 | 10366516 | 76499.594 | 12.831 |  | 43.99
[17 empty] | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.4 | 199021.969 | 7933975 | 73346.039 | 10.854 |  | 37.13
[17 empty] | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.39 | 195954.266 | 7834061 | 68280.641 | 11.479 |  | 39.3
[17 empty] | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.4 | 209089.75 | 8472446 | 73499.289 | 11.379 |  | 38.96
[17 empty] | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.4 | 163660.906 | 6955397 | 55749.684 | 11.743 |  | 40.22
[17 empty] | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.4 | 87596.305 | 3600304 | 30788.695 | 11.38 |  | 38.96
[17 empty] | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.4 | 144037.141 | 5669520 | 54407.445 | 10.59 |  | 36.21
[17 empty] | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.4 | 198084.219 | 8400724 | 76163.773 | 10.403 |  | 35.57
[17 empty] | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.4 | 241833.688 | 9651093 | 92251.773 | 10.486 |  | 35.85
[17 empty] | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.4 | 218327.75 | 8751073 | 83986.945 | 10.398 |  | 35.55
[17 empty] | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.4 | 238265.75 | 9592995 | 86446.32 | 11.025 |  | 37.73
[17 empty] | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.4 | 262954.25 | 10368118 | 89926.844 | 11.696 |  | 40.06
[17 empty] | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.4 | 266459.938 | 11257610 | 91128.219 | 11.696 |  | 40.06
[17 empty] | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.4 | 27795.627 | 1163048 | 107375.094 | 1.035 |  | 3.05
[17 empty] | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.4 | 4972.751 | 197217 | 104266.414 | 0.191 |  | 0.12
[17 empty] | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.4 | 5474.621 | 219701 | 105538.766 | 0.207 |  | 0.18
[17 empty] | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.4 | 2002.274 | 89927 | 84474.836 | 0.095
[17 empty] | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.4 | 2397.686 | 94151 | 87116.578 | 0.11
[17 empty] | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.4 | 2073.612 | 89065 | 83468.359 | 0.099
[17 empty] | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.4 | 2722.367 | 114256 | 89927.414 | 0.121
[17 empty] | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.4 | 2252.211 | 97347 | 86421.602 | 0.104
[17 empty] | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.39 | 3439.174 | 128733 | 91154.789 | 0.151
[17 empty] | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.4 | 2381.84 | 88870 | 89026.438 | 0.107
[17 empty] | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.4 | 6998.725 | 272006 | 89150.375 | 0.314 |  | 0.55
[17 empty] | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.4 | 14859.413 | 603578 | 89808.383 | 0.662 |  | 1.75
[17 empty] | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.4 | 1691.597 | 72046 | 78967.586 | 0.086
[17 empty] | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.4 | 2561.379 | 93938 | 83951.203 | 0.122
[17 empty] | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.39 | 4321.598 | 173179 | 82722.727 | 0.209 |  | 0.18
[17 empty] | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.4 | 2641.276 | 106059 | 80214.852 | 0.132
[17 empty] | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.4 | 2297.828 | 98469 | 77124.359 | 0.119
[17 empty] | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.4 | 1735.724 | 69695 | 74714.813 | 0.093
[17 empty] | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.4 | 4481.813 | 193533 | 99319.828 | 0.181 |  | 0.08
[17 empty] | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.4 | 7830.961 | 317929 | 94711.563 | 0.331 |  | 0.6
[17 empty] | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.4 | 9123.27 | 397031 | 90822.211 | 0.402 |  | 0.85
[17 empty] | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.4 | 2338.676 | 92326 | 84107.078 | 0.111
[17 empty] | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.4 | 1933.72 | 74930 | 82892.469 | 0.093
[17 empty] | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.4 | 2124.865 | 95781 | 89697.148 | 0.095
[17 empty] | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.4 | 2743.245 | 102340 | 84959.609 | 0.129
[17 empty] | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.4 | 3430.022 | 134885 | 86283.852 | 0.159 |  | 0.01
[17 empty] | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.4 | 2008.583 | 88863 | 100046.359 | 0.08
[17 empty] | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.4 | 14150.011 | 605117 | 95674.383 | 0.592 |  | 1.51
[17 empty] | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.4 | 6172.626 | 255818 | 93804.273 | 0.263 |  | 0.37
[17 empty] | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.4 | 5040.559 | 197915 | 75229.672 | 0.268 |  | 0.39
[17 empty] | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.4 | 2383.649 | 96799 | 87124.805 | 0.109
[17 empty] | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.4 | 1627.43 | 65057 | 77005.422 | 0.085
[17 empty] | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.4 | 3482.942 | 128910 | 88369.641 | 0.158 |  | 0
[17 empty] | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.4 | 1741.286 | 67916 | 76210.016 | 0.091
[17 empty] | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.4 | 1187.428 | 48088 | 52721.07 | 0.09
[17 empty] | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.4 | 1416.868 | 47558 | 55817.023 | 0.102
[17 empty] | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.4 | 2355.618 | 92222 | 99776.836 | 0.094
[17 empty] | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.4 | 3240.659 | 122106 | 101979.219 | 0.127
[17 empty] | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.4 | 16277.225 | 691768 | 114534.242 | 0.568 |  | 1.43
[17 empty] | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.4 | 3266.747 | 141507 | 79844.398 | 0.164 |  | 0.02
[17 empty] | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.4 | 3671.382 | 155139 | 86026.953 | 0.171 |  | 0.05
[17 empty] | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.4 | 3701.034 | 150082 | 90566.922 | 0.163 |  | 0.02
[17 empty] | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.4 | 4395.498 | 171055 | 97325.922 | 0.181 |  | 0.08
[17 empty] | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.4 | 4938.487 | 182785 | 102662.727 | 0.192 |  | 0.12
[17 empty] | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.4 | 3017.959 | 122122 | 99524.109 | 0.121
[17 empty] | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.4 | 2407.487 | 92725 | 104638.984 | 0.092
[17 empty] | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.4 | 2357.775 | 96783 | 95722.523 | 0.099
[17 empty] | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.4 | 2659.664 | 107623 | 104475.281 | 0.102
[17 empty] | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.4 | 6274.098 | 254249 | 96895.539 | 0.259 |  | 0.35
[17 empty] | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.4 | 2641.203 | 106814 | 90586.703 | 0.117
[17 empty] | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.4 | 2226.196 | 79075 | 85220.258 | 0.104
[17 empty] | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.4 | 2858.172 | 106053 | 87958.188 | 0.13
[17 empty] | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.4 | 3113.741 | 112647 | 73946.117 | 0.168 |  | 0.04
[17 empty] | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.4 | 5364.324 | 218230 | 88629.414 | 0.242 |  | 0.3
[17 empty] | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.4 | 2398.448 | 106629 | 93904.141 | 0.102
[17 empty] | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.4 | 2452.679 | 97232 | 99092.422 | 0.099
[17 empty] | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.4 | 3060.941 | 112753 | 102824.828 | 0.119
[17 empty] | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.4 | 2650.645 | 109312 | 74625.852 | 0.142
[17 empty] | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.4 | 2487.767 | 106960 | 78271.125 | 0.127
[17 empty] | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.4 | 1859.349 | 76758 | 82232.203 | 0.09
[17 empty] | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.4 | 2407.742 | 103954 | 68976.289 | 0.14
[17 empty] | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.4 | 1676.365 | 65562 | 80912.125 | 0.083
[17 empty] | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.4 | 1451.845 | 55029 | 64799.211 | 0.09
[17 empty] | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.4 | 2270.83 | 89051 | 95344.883 | 0.095
[17 empty] | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.4 | 2288.021 | 85144 | 98936.914 | 0.093
[17 empty] | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.4 | 2269.482 | 96527 | 96992.188 | 0.094
[17 empty] | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.4 | 2530.248 | 100776 | 108030.414 | 0.094
[17 empty] | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.4 | 3093.332 | 115462 | 100591.141 | 0.123
[17 empty] | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.4 | 1622.044 | 59654 | 69586.211 | 0.093
[17 empty] | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.4 | 121838.508 | 4179704 | 104467.281 | 4.665 |  | 15.65
[17 empty] | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.4 | 4144.071 | 183465 | 48669.789 | 0.341 |  | 0.64
[17 empty] | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.4 | 34674.074 | 1378169 | 39606.652 | 3.502 |  | 11.61
[17 empty] | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.4 | 2658.295 | 114635 | 100054.281 | 0.106
[17 empty] | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.4 | 2295.008 | 85333 | 91328.695 | 0.101
[17 empty] | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.4 | 2382.871 | 91499 | 102183.391 | 0.093
[17 empty] | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.4 | 2309.284 | 84392 | 95108.883 | 0.097
[17 empty] | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.4 | 1318.323 | 47264 | 53921.77 | 0.098
[17 empty] | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.4 | 1408.505 | 61221 | 61000.227 | 0.092
[17 empty] | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.4 | 6702.993 | 251062 | 180164.406 | 0.149
[17 empty] | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.4 | 3526.294 | 137227 | 128731.063 | 0.11
[17 empty] | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.4 | 3763.688 | 148501 | 119917.117 | 0.126
[17 empty] | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.4 | 9193.139 | 358467 | 88866.891 | 0.414 |  | 0.89
[17 empty] | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.4 | 9416.714 | 361946 | 88157.313 | 0.427 |  | 0.94
[17 empty] | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.4 | 10473.873 | 423184 | 98938.063 | 0.423 |  | 0.93
[17 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.42 | 18.895 | 1390
[17 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.41 | 62.547 | 1907
[17 empty] | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.39 | 1.965 | 356 | 6.982 | 1.126 |  | 3.36
[17 empty] | 20231106_HepCl_ReRun_006 | Blank | Blank | 4.39 | 12.852 | 1235
[17 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.35 | 36.801 | 1623 | 7.306 | 20.148 |  | 69.4
[17 empty] | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.42 | 1.451 | 372 | 0.639 | 9.083 |  | 30.99
[17 empty] | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.43 | 18.559 | 1477
[17 empty] | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 4.114
[17 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank |  |  |  | 7.725
[17 empty] | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.44 | 65.969 | 2411 | 5.681 | 46.449 |  | 160.69
[17 empty] | 20231106_HepCl_ReRun_053 | Blank | Blank | 4.43 | 2.767 | 722 | 3.035 | 3.647 |  | 12.11
[17 empty] | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.43 | 61.475 | 1351 | 2.979 | 82.544 |  | 285.99
[17 empty] | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.45 | 47.228 | 2391 | 13.529 | 13.963 |  | 47.93
[17 empty] | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.38 | 69.248 | 2207 | 8.646 | 32.037 |  | 110.66
[17 empty] | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 10.156
[17 empty] | 20231106_HepCl_ReRun_084 | Blank | Blank |  |  |  | 4.803
[17 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.36 | 4.824 | 525 | 0.242 | 79.736 |  | 276.24
[17 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | 4.36 | 22.948 | 1505
[17 empty] | 20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 0.119
[17 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.42 | 49.545 | 2242 | 1.22 | 162.443 |  | 563.34
[17 empty] | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.44 | 120.509 | 2214 | 7.197 | 66.977 |  | 231.95
[17 empty] | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.42 | 11.766 | 1260 | 1.133 | 41.539 |  | 143.65
[17 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.4 | 73.177 | 2401 | 1.785 | 163.982 |  | 568.68
[17 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank |  |  |  | 8.867
[17 empty] | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.43 | 5.209 | 616 | 0.21 | 99.219 |  | 343.87
[17 empty] | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.39 | 20.099 | 1631 | 1.004 | 80.076 |  | 277.42
[17 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.43 | 18.612 | 2037 | 3.553 | 20.954 |  | 72.19
[17 empty] | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.41 | 27.864 | 1977 | 4.462 | 24.979 |  | 86.16
[17 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_152 | Blank | Blank |  |  |  | 6.476
[17 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank |  |  |  | 0.75
[17 empty] | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.41 | 1.558 | 302 | 3.511 | 1.775 |  | 5.62
[17 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank
[17 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.43 | 88.735 | 1658 | 6.395 | 55.503 |  | 192.12
[17 empty] | 20231106_HepCl_ReRun_182 | Blank | Blank |  |  |  | 0.43
[17 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.43 | 14.778 | 1378 | 2.904 | 20.355 |  | 70.11
[17 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.41 | 3.016 | 787 | 2.254 | 5.352 |  | 18.03
[17 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank | 4.35 | 76.1 | 1422 | 0.96 | 317.083 |  | 1100.13
[17 empty] | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.43 | 185.614 | 2410 | 1.102 | 673.735 |  | 2338.17
[17 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.4 | 24.258 | 1634 | 1.83 | 53.023 |  | 183.51
[17 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank |  |  |  | 1.23
[17 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.36 | 269.191 | 2916 | 0.419 | 2569.842 |  | 8920.05
[17 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank |  |  |  | 5.807
[17 empty] | 20231106_HepCl_ReRun_230 | Blank | Blank |  |  |  | 2.226
[17 empty] | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 8.198
[17 empty] | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 7.8
[17 empty] | 20231106_HepCl_ReRun_246 | Blank | Blank |  |  |  | 4.036
[17 empty] | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.43 | 12.221 | 1633 | 1.238 | 39.486 |  | 136.52
[17 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank |  |  |  | 7.609
[17 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.43 | 36.339 | 1486 | 3 | 48.452 |  | 167.65
[17 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.37 | 140.794 | 2229 | 3.127 | 180.101 |  | 624.63
[17 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.36 | 23.582 | 2277 | 4.24 | 22.247 |  | 76.68
[17 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank |  |  |  | 1.228
[17 empty] | 20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.469
[17 empty] | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.43 | 13.412 | 1130 | 5.073 | 10.575 |  | 36.17
[17 empty] | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.38 | 59.634 | 2149 | 16.466 | 14.487 |  | 49.74
[17 empty] | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 9.694
[17 empty] | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4.43 | 0.86 | 403 | 6.959 | 0.494 |  | 1.17

## Working Sheet
6:3 FTCA | [10 empty] | 6:2 FTS | [10 empty] | M2-6:2 FTS | [10 empty] | 9H-PFNA | [10 empty] | M9PFNA | [10 empty] | 8:2 FTS | [10 empty] | M2-8:2 FTS | [10 empty] | 11H-PFUnDA | [10 empty] | PFUnDA | [10 empty] | M7PFUdA
Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
20231106_HepCl_ReRun_013 | CC06 | Standard | 3.88 | 447.461 | 8778 | 3034.11 | 0.59 | 1.825 | 1.43 | -21.8 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.78 | 171.963 | 5263 | 3034.11 | 0.227 | 1.825 | 1.75 | -4.2 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.77 | 3034.11 | 104487 |  | 3034.11 | 4 | 3.75 | -6.2 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.8 | 11641.195 | 382511 | 68219.305 | 0.683 | 1.825 | 1.52 | -16.5 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.96 | 68219.305 | 1940316 |  | 68219.305 | 4 | 4.12 | 2.9 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.05 | 183.985 | 6873 | 1285.358 | 0.573 | 1.825 | 2.43 | 33 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.05 | 1285.358 | 48475 |  | 1285.358 | 4 | 3.82 | -4.4 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.06 | 1062.268 | 36328 | 40008.617 | 0.106 | 1.825 | 1.14 | -37.4 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.4 | 7043 | 277350 | 40008.617 | 0.704 | 1.825 | 1.9 | 4.1 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.4 | 40008.617 | 1617090 |  | 40008.617 | 4 | 4.48 | 11.9
20231106_HepCl_ReRun_031 | CC06 | Standard | 3.88 | 478.13 | 8539 | 2368.634 | 0.807 | 1.825 | 2.22 | 21.7 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.78 | 168.791 | 5259 | 2368.634 | 0.285 | 1.825 | 2.61 | 43 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.78 | 2368.634 | 80581 |  | 2368.634 | 4 | 2.93 | -26.8 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.81 | 10848.09 | 359772 | 65001.293 | 0.668 | 1.825 | 1.48 | -18.9 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.96 | 65001.293 | 1804041 |  | 65001.293 | 4 | 3.92 | -2 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.05 | 134.609 | 4719 | 1266.854 | 0.425 | 1.825 | 1.54 | -15.5 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.05 | 1266.854 | 47688 |  | 1266.854 | 4 | 3.77 | -5.8 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.07 | 1106.86 | 42580 | 33664.371 | 0.132 | 1.825 | 1.57 | -14.2 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.4 | 5588.62 | 236229 | 33664.371 | 0.664 | 1.825 | 1.76 | -3.5 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.4 | 33664.371 | 1373684 |  | 33664.371 | 4 | 3.77 | -5.9
20231106_HepCl_ReRun_035 | CC06 | Standard | 3.87 | 418.264 | 7300 | 2974.036 | 0.563 | 1.825 | 1.33 | -27.3 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.78 | 141.014 | 4907 | 2974.036 | 0.19 | 1.825 | 1.2 | -34.2 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.77 | 2974.036 | 102368 |  | 2974.036 | 4 | 3.68 | -8.1 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.8 | 11571.36 | 384229 | 66638.734 | 0.695 | 1.825 | 1.56 | -14.7 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.96 | 66638.734 | 1854201 |  | 66638.734 | 4 | 4.02 | 0.5 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.05 | 139.414 | 5172 | 1260.427 | 0.442 | 1.825 | 1.65 | -9.8 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.05 | 1260.427 | 46864 |  | 1260.427 | 4 | 3.75 | -6.3 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.06 | 1030.858 | 39085 | 31861.572 | 0.129 | 1.825 | 1.53 | -16.1 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.4 | 5433.467 | 207631 | 31861.572 | 0.682 | 1.825 | 1.82 | -0.1 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.4 | 31861.572 | 1318096 |  | 31861.572 | 4 | 3.56 | -10.9
20231106_HepCl_ReRun_039 | CC06 | Standard | 3.88 | 523.429 | 9106 | 2930.851 | 0.714 | 1.825 | 1.88 | 3.1 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.78 | 177.583 | 5236 | 2930.851 | 0.242 | 1.825 | 1.98 | 8.4 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.78 | 2930.851 | 99746 |  | 2930.851 | 4 | 3.62 | -9.4 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.81 | 11606.056 | 384908 | 61013.957 | 0.761 | 1.825 | 1.75 | -4.3 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.96 | 61013.957 | 1674807 |  | 61013.957 | 4 | 3.68 | -8 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.05 | 140.109 | 4945 | 1310.875 | 0.428 | 1.825 | 1.56 | -14.7 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.05 | 1310.875 | 49191 |  | 1310.875 | 4 | 3.9 | -2.5 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.07 | 1048.295 | 38562 | 30194.342 | 0.139 | 1.825 | 1.69 | -7.5 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.4 | 5458.212 | 220757 | 30194.342 | 0.723 | 1.825 | 1.97 | 7.7 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.4 | 30194.342 | 1255627 |  | 30194.342 | 4 | 3.38 | -15.6
20231106_HepCl_ReRun_052 | CC06 | Standard | 3.88 | 452.235 | 7649 | 2982.801 | 0.606 | 1.825 | 1.49 | -18.5 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.78 | 157.768 | 4759 | 2982.801 | 0.212 | 1.825 | 1.52 | -16.5 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.78 | 2982.801 | 103645 |  | 2982.801 | 4 | 3.69 | -7.8 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.81 | 11166.135 | 371375 | 64340.457 | 0.694 | 1.825 | 1.56 | -14.7 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.96 | 64340.457 | 1828969 |  | 64340.457 | 4 | 3.88 | -3 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.05 | 146.812 | 5980 | 1179.859 | 0.498 | 1.825 | 1.98 | 8.4 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.05 | 1179.859 | 41926 |  | 1179.859 | 4 | 3.51 | -12.3 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.06 | 1061.033 | 39434 | 27641.32 | 0.154 | 1.825 | 1.93 | 5.9 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.4 | 4747.757 | 201851 | 27641.32 | 0.687 | 1.825 | 1.84 | 0.9 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.4 | 27641.32 | 1122443 |  | 27641.32 | 4 | 3.09 | -22.7
20231106_HepCl_ReRun_070 | CC06 | Standard | 3.88 | 472.381 | 7598 | 3064.16 | 0.617 | 1.825 | 1.52 | -16.5 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.78 | 156.739 | 4706 | 3064.16 | 0.205 | 1.825 | 1.42 | -22.1 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.78 | 3064.16 | 104987 |  | 3064.16 | 4 | 3.79 | -5.3 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.81 | 11005.288 | 365567 | 63387.512 | 0.694 | 1.825 | 1.56 | -14.7 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.96 | 63387.512 | 1742378 |  | 63387.512 | 4 | 3.82 | -4.4 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.05 | 120.27 | 4640 | 1220.534 | 0.394 | 1.825 | 1.36 | -25.6 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.05 | 1220.534 | 47238 |  | 1220.534 | 4 | 3.63 | -9.2 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.06 | 1028.637 | 39775 | 26664.004 | 0.154 | 1.825 | 1.95 | 6.6 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.4 | 4767.315 | 194365 | 26664.004 | 0.715 | 1.825 | 1.94 | 6.2 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.4 | 26664.004 | 1080929 |  | 26664.004 | 4 | 2.98 | -25.4
20231106_HepCl_ReRun_074 | CC06 | Standard | 3.88 | 432.128 | 7512 | 3262.376 | 0.53 | 1.825 | 1.21 | -33.9 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.78 | 145.769 | 5063 | 3262.376 | 0.179 | 1.825 | 1.04 | -43.1 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.77 | 3262.376 | 111380 |  | 3262.376 | 4 | 4.03 | 0.8 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.81 | 10255.913 | 336728 | 62066.836 | 0.661 | 1.825 | 1.46 | -19.9 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.96 | 62066.836 | 1728402 |  | 62066.836 | 4 | 3.74 | -6.4 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.05 | 143.867 | 6016 | 1181.049 | 0.487 | 1.825 | 1.92 | 5 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.05 | 1181.049 | 43690 |  | 1181.049 | 4 | 3.51 | -12.2 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.06 | 967.527 | 35145 | 27006.176 | 0.143 | 1.825 | 1.76 | -3.4 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.4 | 4664.719 | 193278 | 27006.176 | 0.691 | 1.825 | 1.85 | 1.6 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.4 | 27006.176 | 1109378 |  | 27006.176 | 4 | 3.02 | -24.5
20231106_HepCl_ReRun_078 | CC06 | Standard | 3.88 | 465.458 | 7889 | 3313.979 | 0.562 | 1.825 | 1.32 | -27.5 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.78 | 168.728 | 5698 | 3313.979 | 0.204 | 1.825 | 1.41 | -22.9 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.77 | 3313.979 | 112946 |  | 3313.979 | 4 | 4.1 | 2.4 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.8 | 11621.744 | 390664 | 61289.246 | 0.758 | 1.825 | 1.74 | -4.7 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.96 | 61289.246 | 1692146 |  | 61289.246 | 4 | 3.7 | -7.6 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.05 | 129.094 | 5605 | 1218.418 | 0.424 | 1.825 | 1.53 | -15.9 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.05 | 1218.418 | 46861 |  | 1218.418 | 4 | 3.62 | -9.4 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.06 | 993.691 | 37872 | 24890.996 | 0.16 | 1.825 | 2.04 | 11.6 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.39 | 4229.273 | 162426 | 24890.996 | 0.68 | 1.825 | 1.82 | -0.5 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.4 | 24890.996 | 992358 |  | 24890.996 | 4 | 2.78 | -30.4
20231106_HepCl_ReRun_015 | CC07 | Standard | 3.88 | 779.956 | 11885 | 2993.711 | 1.042 | 2.919 | 3.08 | 5.5 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.78 | 238.272 | 7112 | 2993.711 | 0.318 | 2.919 | 3.1 | 6.3 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.78 | 2993.711 | 101914 |  | 2993.711 | 4 | 3.7 | -7.5 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.81 | 21097.494 | 703071 | 64042.156 | 1.318 | 2.919 | 3.33 | 13.9 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.96 | 64042.156 | 1714197 |  | 64042.156 | 4 | 3.86 | -3.4 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.05 | 221.348 | 8806 | 1359.29 | 0.651 | 2.919 | 2.9 | -0.6 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.05 | 1359.29 | 50339 |  | 1359.29 | 4 | 4.04 | 1.1 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.06 | 1882.456 | 72625 | 42977.938 | 0.175 | 2.919 | 2.3 | -21.4 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.4 | 10383.962 | 423591 | 42977.938 | 0.966 | 2.919 | 2.81 | -3.7 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.4 | 42977.938 | 1740176 |  | 42977.938 | 4 | 4.81 | 20.2
20231106_HepCl_ReRun_054 | CC07 | Standard | 3.88 | 712.941 | 14349 | 3350.341 | 0.851 | 2.919 | 2.38 | -18.4 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.78 | 193.975 | 6727 | 3350.341 | 0.232 | 2.919 | 1.82 | -37.7 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.77 | 3350.341 | 114845 |  | 3350.341 | 4 | 4.14 | 3.6 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.8 | 20381.063 | 685582 | 67724.766 | 1.204 | 2.919 | 3 | 2.8 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.96 | 67724.766 | 1893346 |  | 67724.766 | 4 | 4.09 | 2.1 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.05 | 193.381 | 7888 | 1307.154 | 0.592 | 2.919 | 2.54 | -12.9 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.05 | 1307.154 | 48444 |  | 1307.154 | 4 | 3.89 | -2.8 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.06 | 1728.86 | 67243 | 30703.615 | 0.225 | 2.919 | 3.13 | 7.3 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.39 | 6970.027 | 276378 | 30703.615 | 0.908 | 2.919 | 2.61 | -10.7 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.39 | 30703.615 | 1272896 |  | 30703.615 | 4 | 3.43 | -14.1
20231106_HepCl_ReRun_016 | CC08 | Standard | 3.88 | 1201.809 | 19563 | 3101.963 | 1.55 | 4.669 | 4.94 | 5.7 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.78 | 347.814 | 11862 | 3101.963 | 0.449 | 4.669 | 5.03 | 7.6 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.78 | 3101.963 | 107129 |  | 3101.963 | 4 | 3.84 | -4.1 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.81 | 29158.406 | 967248 | 65068.902 | 1.792 | 4.669 | 4.67 | 0.1 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.97 | 65068.902 | 1781481 |  | 65068.902 | 4 | 3.93 | -1.9 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.05 | 362.115 | 14411 | 1397.361 | 1.037 | 4.669 | 5.21 | 11.6 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.05 | 1397.361 | 52204 |  | 1397.361 | 4 | 4.16 | 3.9 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.06 | 3141.384 | 120411 | 44416.688 | 0.283 | 4.669 | 4.09 | -12.3 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.4 | 16841.48 | 684362 | 44416.688 | 1.517 | 4.669 | 4.72 | 1.1 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.4 | 44416.688 | 1823227 |  | 44416.688 | 4 | 4.97 | 24.2
20231106_HepCl_ReRun_055 | CC08 | Standard | 3.87 | 1077.497 | 18622 | 2541.916 | 1.696 | 4.669 | 5.47 | 17.1 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.78 | 335.471 | 11180 | 2541.916 | 0.528 | 4.669 | 6.2 | 32.8 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.78 | 2541.916 | 86105 |  | 2541.916 | 4 | 3.14 | -21.4 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.81 | 26398.207 | 865225 | 64294.883 | 1.642 | 4.669 | 4.25 | -9.1 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.96 | 64294.883 | 1769355 |  | 64294.883 | 4 | 3.88 | -3 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.05 | 279.907 | 10543 | 1383.121 | 0.809 | 4.669 | 3.85 | -17.6 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.05 | 1383.121 | 50458 |  | 1383.121 | 4 | 4.11 | 2.9 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.06 | 2842.049 | 102842 | 30452.596 | 0.373 | 4.669 | 5.6 | 20 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.4 | 11831.235 | 485726 | 30452.596 | 1.554 | 4.669 | 4.85 | 3.9 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.4 | 30452.596 | 1229852 |  | 30452.596 | 4 | 3.41 | -14.8
20231106_HepCl_ReRun_017 | CC09 | Standard | 3.88 | 1982.509 | 35805 | 3138.186 | 2.527 | 7.45 | 8.51 | 14.2 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.78 | 584.512 | 18989 | 3138.186 | 0.745 | 7.45 | 9.41 | 26.3 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.78 | 3138.186 | 108340 |  | 3138.186 | 4 | 3.88 | -3 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.81 | 48933.699 | 1647090 | 70109.063 | 2.792 | 7.45 | 7.51 | 0.8 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.96 | 70109.063 | 1960652 |  | 70109.063 | 4 | 4.23 | 5.7 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.05 | 517.63 | 19667 | 1364.509 | 1.517 | 7.45 | 8.1 | 8.7 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.05 | 1364.509 | 50516 |  | 1364.509 | 4 | 4.06 | 1.5 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.07 | 4509.686 | 158142 | 41387.938 | 0.436 | 7.45 | 6.65 | -10.8 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.4 | 23083.746 | 919942 | 41387.938 | 2.231 | 7.45 | 7.2 | -3.4 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.4 | 41387.938 | 1671639 |  | 41387.938 | 4 | 4.63 | 15.7
20231106_HepCl_ReRun_056 | CC09 | Standard | 3.87 | 2015.593 | 34908 | 3151.321 | 2.558 | 7.45 | 8.62 | 15.7 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.78 | 445.528 | 14562 | 3151.321 | 0.566 | 7.45 | 6.75 | -9.3 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.78 | 3151.321 | 108098 |  | 3151.321 | 4 | 3.9 | -2.6 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.81 | 46315.258 | 1565880 | 63195.191 | 2.932 | 7.45 | 7.9 | 6.1 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.96 | 63195.191 | 1738904 |  | 63195.191 | 4 | 3.81 | -4.7 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.05 | 502.933 | 19943 | 1332.81 | 1.509 | 7.45 | 8.05 | 8 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.05 | 1332.81 | 48860 |  | 1332.81 | 4 | 3.96 | -0.9 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.06 | 4361.453 | 168909 | 29372.057 | 0.594 | 7.45 | 9.29 | 24.7 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.4 | 16286.906 | 675195 | 29372.057 | 2.218 | 7.45 | 7.16 | -4 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.4 | 29372.057 | 1219939 |  | 29372.057 | 4 | 3.29 | -17.9
20231106_HepCl_ReRun_018 | CC10 | Standard | 3.87 | 2973.767 | 47181 | 3163.276 | 3.76 | 11.875 | 13.02 | 9.6 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.78 | 761.471 | 26224 | 3163.276 | 0.963 | 11.875 | 12.63 | 6.3 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.77 | 3163.276 | 108367 |  | 3163.276 | 4 | 3.91 | -2.2 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.81 | 79207.258 | 2656081 | 69572.93 | 4.554 | 11.875 | 12.5 | 5.3 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.96 | 69572.93 | 1929098 |  | 69572.93 | 4 | 4.2 | 4.9 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.05 | 764.35 | 28612 | 1367.205 | 2.236 | 11.875 | 12.41 | 4.5 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.05 | 1367.205 | 51313 |  | 1367.205 | 4 | 4.07 | 1.7 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.06 | 7758.193 | 294767 | 43287.93 | 0.717 | 11.875 | 11.34 | -4.5 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.4 | 36668.535 | 1499343 | 43287.93 | 3.388 | 11.875 | 11.22 | -5.5 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.4 | 43287.93 | 1755786 |  | 43287.93 | 4 | 4.84 | 21.1
20231106_HepCl_ReRun_057 | CC10 | Standard | 3.88 | 3366.109 | 59308 | 3307.122 | 4.071 | 11.875 | 14.15 | 19.2 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.78 | 725.711 | 25169 | 3307.122 | 0.878 | 11.875 | 11.37 | -4.3 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.78 | 3307.122 | 114196 |  | 3307.122 | 4 | 4.09 | 2.2 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.81 | 75334.328 | 2526905 | 63700.488 | 4.731 | 11.875 | 13.01 | 9.5 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.96 | 63700.488 | 1730554 |  | 63700.488 | 4 | 3.84 | -3.9 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.05 | 682.924 | 25853 | 1340.69 | 2.038 | 11.875 | 11.22 | -5.5 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.05 | 1340.69 | 49527 |  | 1340.69 | 4 | 3.99 | -0.3 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.07 | 6099.027 | 213476 | 28786.756 | 0.847 | 11.875 | 13.52 | 13.9 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.4 | 23608.922 | 980457 | 28786.756 | 3.281 | 11.875 | 10.84 | -8.7 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.4 | 28786.756 | 1209493 |  | 28786.756 | 4 | 3.22 | -19.5
20231106_HepCl_ReRun_019 | CC11 | Standard | 3.88 | 5696.95 | 94449 | 3229.261 | 7.057 | 19.125 | 25.07 | 31.1 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.78 | 1328.398 | 45730 | 3229.261 | 1.645 | 19.125 | 22.71 | 18.8 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.78 | 3229.261 | 110694 |  | 3229.261 | 4 | 3.99 | -0.2 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.81 | 161645.313 | 5450768 | 72849.406 | 8.876 | 19.125 | 24.76 | 29.5 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.96 | 72849.406 | 2016856 |  | 72849.406 | 4 | 4.39 | 9.9 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.05 | 1354.45 | 51659 | 1491.731 | 3.632 | 19.125 | 20.78 | 8.7 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.05 | 1491.731 | 55680 |  | 1491.731 | 4 | 4.44 | 10.9 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.07 | 13921.18 | 525530 | 43650.934 | 1.276 | 19.125 | 20.67 | 8.1 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.4 | 63299.324 | 2670024 | 43650.934 | 5.801 | 19.125 | 19.59 | 2.4 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.4 | 43650.934 | 1729358 |  | 43650.934 | 4 | 4.88 | 22.1
20231106_HepCl_ReRun_058 | CC11 | Standard | 3.88 | 5013.81 | 76829 | 3103.845 | 6.461 | 19.125 | 22.89 | 19.7 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.78 | 1252.312 | 42857 | 3103.845 | 1.614 | 19.125 | 22.25 | 16.3 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.78 | 3103.845 | 107523 |  | 3103.845 | 4 | 3.84 | -4.1 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.81 | 152744.953 | 5169925 | 68771.055 | 8.884 | 19.125 | 24.79 | 29.6 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.96 | 68771.055 | 1866345 |  | 68771.055 | 4 | 4.15 | 3.7 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.05 | 1244.492 | 46411 | 1394.592 | 3.569 | 19.125 | 20.41 | 6.7 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.05 | 1394.592 | 50530 |  | 1394.592 | 4 | 4.15 | 3.7 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.06 | 12133.325 | 451289 | 31747.754 | 1.529 | 19.125 | 24.9 | 30.2 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.4 | 45632.473 | 1810515 | 31747.754 | 5.749 | 19.125 | 19.41 | 1.5 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.4 | 31747.754 | 1296720 |  | 31747.754 | 4 | 3.55 | -11.2
20231106_HepCl_ReRun_020 | CC12 | Standard | 3.87 | 6181.179 | 122226 | 2767.938 | 8.933 | 30.5 | 31.93 | 4.7 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.77 | 1914.911 | 65185 | 2767.938 | 2.767 | 30.5 | 39.29 | 28.8 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.77 | 2767.938 | 94438 |  | 2767.938 | 4 | 3.42 | -14.4 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.8 | 217131.609 | 7364952 | 60371.941 | 14.386 | 30.5 | 40.4 | 32.4 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.96 | 60371.941 | 1653881 |  | 60371.941 | 4 | 3.64 | -8.9 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.05 | 2082.749 | 78265 | 1313.037 | 6.345 | 30.5 | 37.06 | 21.5 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.05 | 1313.037 | 49232 |  | 1313.037 | 4 | 3.91 | -2.4 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.06 | 18356.199 | 719052 | 43166.441 | 1.701 | 30.5 | 27.78 | -8.9 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.39 | 110621.594 | 4408315 | 43166.441 | 10.251 | 30.5 | 35.04 | 14.9 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.4 | 43166.441 | 1736996 |  | 43166.441 | 4 | 4.83 | 20.7
20231106_HepCl_ReRun_059 | CC12 | Standard | 3.88 | 8168.249 | 142368 | 3365.248 | 9.709 | 30.5 | 34.77 | 14 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.78 | 2141.835 | 73548 | 3365.248 | 2.546 | 30.5 | 36.02 | 18.1 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.77 | 3365.248 | 115566 |  | 3365.248 | 4 | 4.16 | 4 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.81 | 215742.5 | 7243975 | 69718.367 | 12.378 | 30.5 | 34.7 | 13.8 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.96 | 69718.367 | 1978759 |  | 69718.367 | 4 | 4.21 | 5.1 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.05 | 2029.725 | 77604 | 1381.883 | 5.875 | 30.5 | 34.24 | 12.3 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.05 | 1381.883 | 52583 |  | 1381.883 | 4 | 4.11 | 2.8 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.06 | 19046.547 | 718133 | 31395.316 | 2.427 | 30.5 | 39.89 | 30.8 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.4 | 81080.563 | 3271567 | 31395.316 | 10.33 | 30.5 | 35.31 | 15.8 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.4 | 31395.316 | 1232338 |  | 31395.316 | 4 | 3.51 | -12.2
20231106_HepCl_ReRun_022 | CC13 | Standard | 3.88 | 11224.8 | 178932 | 3439.733 | 13.053 | 48.75 | 46.99 | -3.6 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.78 | 3139.108 | 108007 | 3439.733 | 3.65 | 48.75 | 52.34 | 7.4 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.78 | 3439.733 | 118429 |  | 3439.733 | 4 | 4.25 | 6.3 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.81 | 279932.938 | 9373593 | 64329.785 | 17.406 | 48.75 | 48.96 | 0.4 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.96 | 64329.785 | 1726832 |  | 64329.785 | 4 | 3.88 | -3 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.05 | 2907.99 | 115954 | 1391.243 | 8.361 | 48.75 | 49.15 | 0.8 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.05 | 1391.243 | 51807 |  | 1391.243 | 4 | 4.14 | 3.5 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.06 | 26155.953 | 903097 | 44292.91 | 2.362 | 48.75 | 38.82 | -20.4 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.4 | 142737.063 | 5848686 | 44292.91 | 12.89 | 48.75 | 44.2 | -9.3 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.4 | 44292.91 | 1771998 |  | 44292.91 | 4 | 4.95 | 23.9
20231106_HepCl_ReRun_061 | CC13 | Standard | 3.88 | 12265.34 | 222118 | 3673.549 | 13.355 | 48.75 | 48.1 | -1.3 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.78 | 3087.364 | 107125 | 3673.549 | 3.362 | 48.75 | 48.08 | -1.4 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.78 | 3673.549 | 126337 |  | 3673.549 | 4 | 4.54 | 13.6 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.8 | 275087.375 | 9279365 | 69190.055 | 15.903 | 48.75 | 44.7 | -8.3 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.96 | 69190.055 | 1947788 |  | 69190.055 | 4 | 4.17 | 4.4 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.05 | 2793.315 | 105829 | 1344.772 | 8.309 | 48.75 | 48.84 | 0.2 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.05 | 1344.772 | 48392 |  | 1344.772 | 4 | 4 | 0 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.06 | 28121.049 | 1057346 | 34492.879 | 3.261 | 48.75 | 53.83 | 10.4 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.4 | 110454.477 | 4597438 | 34492.879 | 12.809 | 48.75 | 43.92 | -9.9 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.4 | 34492.879 | 1450823 |  | 34492.879 | 4 | 3.86 | -3.5
20231106_HepCl_ReRun_023 | CC14 | Standard | 3.88 | 22367.092 | 397198 | 3919.615 | 22.826 | 78.125 | 82.72 | 5.9 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.78 | 4940.688 | 170489 | 3919.615 | 5.042 | 78.125 | 72.91 | -6.7 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.78 | 3919.615 | 134613 |  | 3919.615 | 4 | 4.85 | 21.2 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.81 | 510067.344 | 17057968 | 72004.906 | 28.335 | 78.125 | 79.96 | 2.4 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.96 | 72004.906 | 1945760 |  | 72004.906 | 4 | 4.34 | 8.6 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.05 | 4992.377 | 192789 | 1497.735 | 13.333 | 78.125 | 78.98 | 1.1 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.05 | 1497.735 | 54749 |  | 1497.735 | 4 | 4.46 | 11.4 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.06 | 50974.664 | 1864530 | 50112.5 | 4.069 | 78.125 | 67.32 | -13.8 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.4 | 284172 | 11155492 | 50112.5 | 22.683 | 78.125 | 78.19 | 0.1 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.4 | 50112.5 | 1895970 |  | 50112.5 | 4 | 5.61 | 40.1
20231106_HepCl_ReRun_062 | CC14 | Standard | 3.88 | 21361.063 | 375481 | 4066.942 | 21.009 | 78.125 | 76.08 | -2.6 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.78 | 5185.91 | 177930 | 4066.942 | 5.101 | 78.125 | 73.77 | -5.6 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.78 | 4066.942 | 139930 |  | 4066.942 | 4 | 5.03 | 25.7 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.81 | 494636.938 | 16853748 | 75166.273 | 26.322 | 78.125 | 74.25 | -5 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.96 | 75166.273 | 2079953 |  | 75166.273 | 4 | 4.53 | 13.4 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.05 | 4922.576 | 187116 | 1523.123 | 12.928 | 78.125 | 76.55 | -2 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.05 | 1523.123 | 57253 |  | 1523.123 | 4 | 4.53 | 13.3 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.06 | 49863.824 | 1864208 | 41397.715 | 4.818 | 78.125 | 79.83 | 2.2 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.4 | 222048.391 | 9055133 | 41397.715 | 21.455 | 78.125 | 73.93 | -5.4 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.4 | 41397.715 | 1662759 |  | 41397.715 | 4 | 4.63 | 15.8
20231106_HepCl_ReRun_024 | CC15 | Standard | 3.88 | 30114.336 | 533010 | 3888.294 | 30.979 | 125 | 112.53 | -10 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.78 | 7579.503 | 259401 | 3888.294 | 7.797 | 125 | 113.62 | -9.1 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.78 | 3888.294 | 133357 |  | 3888.294 | 4 | 4.81 | 20.2 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.81 | 653154.375 | 22047396 | 67768.188 | 38.552 | 125 | 108.94 | -12.8 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.97 | 67768.188 | 1918844 |  | 67768.188 | 4 | 4.09 | 2.2 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.05 | 7282.961 | 267543 | 1422.099 | 20.485 | 125 | 121.89 | -2.5 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.05 | 1422.099 | 52406 |  | 1422.099 | 4 | 4.23 | 5.8 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.07 | 71695.906 | 2683730 | 38008.148 | 7.545 | 125 | 125.37 | 0.3 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.4 | 376289.313 | 15172176 | 38008.148 | 39.601 | 125 | 136.92 | 9.5 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.4 | 38008.148 | 1745132 |  | 38008.148 | 4 | 4.25 | 6.3
20231106_HepCl_ReRun_063 | CC15 | Standard | 3.88 | 32258.174 | 560369 | 3980.426 | 32.417 | 125 | 117.79 | -5.8 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.78 | 8270.189 | 287977 | 3980.426 | 8.311 | 125 | 121.21 | -3 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.78 | 3980.426 | 148683 |  | 3980.426 | 4 | 4.92 | 23 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.81 | 687010.063 | 23100348 | 64113.281 | 42.862 | 125 | 121.17 | -3.1 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.96 | 64113.281 | 1795978 |  | 64113.281 | 4 | 3.87 | -3.3 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.05 | 6916.329 | 258533 | 1427.197 | 19.384 | 125 | 115.29 | -7.8 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.05 | 1427.197 | 52880 |  | 1427.197 | 4 | 4.25 | 6.1 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.06 | 70819.211 | 2617717 | 38112.547 | 7.433 | 125 | 123.49 | -1.2 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.4 | 327021.594 | 13200246 | 38112.547 | 34.322 | 125 | 118.6 | -5.1 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.4 | 38112.547 | 1543227 |  | 38112.547 | 4 | 4.26 | 6.6
20231106_HepCl_ReRun_042 | QC B | QC | 3.87 | 710.963 | 12145 | 3127.347 | 0.909 | 2.506 | 2.59 | 3.5 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.78 | 191.108 | 5682 | 3127.347 | 0.244 | 2.506 | 2.01 | -19.8 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.78 | 3127.347 | 113458 |  | 3127.347 | 4 | 3.87 | -3.3 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.81 | 22068.494 | 733400 | 76308.461 | 1.157 | 2.506 | 2.87 | 14.5 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.96 | 76308.461 | 2025701 |  | 76308.461 | 4 | 4.6 | 15.1 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.05 | 231.524 | 9385 | 1697.607 | 0.546 | 2.506 | 2.27 | -9.6 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.05 | 1697.607 | 63218 |  | 1697.607 | 4 | 5.05 | 26.2 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.06 | 1949.468 | 71825 | 51642.734 | 0.151 | 2.506 | 1.89 | -24.6 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.4 | 10822.062 | 445443 | 51642.734 | 0.838 | 2.506 | 2.37 | -5.6 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.4 | 51642.734 | 2142571 |  | 51642.734 | 4 | 5.78 | 44.4
20231106_HepCl_ReRun_081 | QC B | QC | 3.87 | 752.553 | 12162 | 2887.97 | 1.042 | 2.506 | 3.08 | 22.9 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.78 | 213.334 | 7274 | 2887.97 | 0.295 | 2.506 | 2.76 | 10.3 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.77 | 2887.97 | 97800 |  | 2887.97 | 4 | 3.57 | -10.7 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.8 | 19008.291 | 624404 | 84922.633 | 0.895 | 2.506 | 2.13 | -15.1 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.96 | 84922.633 | 2354939 |  | 84922.633 | 4 | 5.12 | 28.1 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.05 | 222.404 | 8528 | 1702.264 | 0.523 | 2.506 | 2.13 | -15.1 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.05 | 1702.264 | 67418 |  | 1702.264 | 4 | 5.06 | 26.6 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.06 | 2080.625 | 81819 | 37154.508 | 0.224 | 2.506 | 3.11 | 24.1 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.4 | 7522.85 | 298162 | 37154.508 | 0.81 | 2.506 | 2.27 | -9.5 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.4 | 37154.508 | 1560233 |  | 37154.508 | 4 | 4.16 | 3.9
20231106_HepCl_ReRun_043 | QC C | QC | 3.88 | 1686.431 | 30672 | 3371.759 | 2.001 | 6.269 | 6.58 | 5 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.78 | 449.471 | 14963 | 3371.759 | 0.533 | 6.269 | 6.28 | 0.1 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.78 | 3371.759 | 115251 |  | 3371.759 | 4 | 4.17 | 4.2 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.81 | 51954.094 | 1766232 | 84343.641 | 2.464 | 6.269 | 6.58 | 4.9 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.96 | 84343.641 | 2347158 |  | 84343.641 | 4 | 5.09 | 27.2 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.05 | 564.021 | 21717 | 1749.339 | 1.29 | 6.269 | 6.73 | 7.4 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.05 | 1749.339 | 68129 |  | 1749.339 | 4 | 5.2 | 30.1 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.06 | 5233.741 | 199717 | 53094.781 | 0.394 | 6.269 | 5.95 | -5 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.4 | 26934.248 | 1136498 | 53094.781 | 2.029 | 6.269 | 6.5 | 3.7 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.4 | 53094.781 | 2272471 |  | 53094.781 | 4 | 5.94 | 48.5
20231106_HepCl_ReRun_082 | QC C | QC | 3.88 | 2095.343 | 37365 | 3610.88 | 2.321 | 6.269 | 7.76 | 23.7 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.78 | 390.65 | 13088 | 3610.88 | 0.433 | 6.269 | 4.79 | -23.5 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.78 | 3610.88 | 124330 |  | 3610.88 | 4 | 4.46 | 11.6 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.81 | 56246.949 | 1905280 | 79141.828 | 2.843 | 6.269 | 7.65 | 22.1 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.96 | 79141.828 | 2191526 |  | 79141.828 | 4 | 4.77 | 19.4 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.05 | 532.346 | 19977 | 1712.841 | 1.243 | 6.269 | 6.45 | 2.9 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.05 | 1712.841 | 65873 |  | 1712.841 | 4 | 5.09 | 27.4 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.06 | 4570.148 | 205965 | 41341.777 | 0.442 | 6.269 | 6.75 | 7.7 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.4 | 21356.891 | 869820 | 41341.777 | 2.066 | 6.269 | 6.63 | 5.7 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.4 | 41341.777 | 1699979 |  | 41341.777 | 4 | 4.62 | 15.6
20231106_HepCl_ReRun_044 | QC D | QC | 3.88 | 6550.212 | 126329 | 3225.131 | 8.124 | 25 | 28.97 | 15.9 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.78 | 1519.89 | 52572 | 3225.131 | 1.885 | 25 | 26.25 | 5 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.78 | 3225.131 | 110928 |  | 3225.131 | 4 | 3.99 | -0.3 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.81 | 182690.922 | 6144534 | 70742.273 | 10.33 | 25 | 28.89 | 15.6 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.96 | 70742.273 | 1886520 |  | 70742.273 | 4 | 4.27 | 6.7 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.05 | 1745.07 | 65258 | 1516.869 | 4.602 | 25 | 26.6 | 6.4 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.05 | 1516.869 | 56991 |  | 1516.869 | 4 | 4.51 | 12.8 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.07 | 18617.66 | 686901 | 46975.297 | 1.585 | 25 | 25.84 | 3.4 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.4 | 85835.969 | 3549459 | 46975.297 | 7.309 | 25 | 24.83 | -0.7 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.4 | 46975.297 | 1959518 |  | 46975.297 | 4 | 5.25 | 31.4
20231106_HepCl_ReRun_083 | QC D | QC | 3.88 | 7293.705 | 119126 | 3571.287 | 8.169 | 25 | 29.14 | 16.5 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.78 | 1626.875 | 56148 | 3571.287 | 1.822 | 25 | 25.32 | 1.3 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.78 | 3571.287 | 123058 |  | 3571.287 | 4 | 4.42 | 10.4 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.8 | 186631.609 | 6350371 | 68663.195 | 10.872 | 25 | 30.43 | 21.7 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.96 | 68663.195 | 1862578 |  | 68663.195 | 4 | 4.14 | 3.6 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.05 | 1766.013 | 67718 | 1490.915 | 4.738 | 25 | 27.42 | 9.7 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.05 | 1490.915 | 57447 |  | 1490.915 | 4 | 4.43 | 10.9 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.06 | 15100.404 | 616116 | 32931.387 | 1.834 | 25 | 30 | 20 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.4 | 60284.25 | 2577048 | 32931.387 | 7.322 | 25 | 24.87 | -0.5 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.4 | 32931.387 | 1374550 |  | 32931.387 | 4 | 3.68 | -7.9
20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.89 | 15743.503 | 247765 | 4776.272 | 13.185 |  | 47.47 |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.78 | 205.262 | 6517 | 4776.272 | 0.172 |  | 0.94 |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.78 | 4776.272 | 164687 |  | 4776.272 | 4 | 5.91 | 47.6 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.81 | 514.675 | 8646 | 127111.25 | 0.016 |  |  |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.97 | 127111.25 | 3554961 |  | 127111.25 | 4 | 7.67 | 91.7 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.05 | 455.547 | 17598 | 2610.134 | 0.698 |  | 3.18 |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.05 | 2610.134 | 101241 |  | 2610.134 | 4 | 7.76 | 94.1 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte |  |  |  | 98002.758 |  |  |  |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.4 | 288979.969 | 11903002 | 98002.758 | 11.795 |  | 40.4 |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.4 | 98002.758 | 4043580 |  | 98002.758 | 4 | 10.96 | 174.1
20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.88 | 13915.396 | 245491 | 3195.141 | 17.421 |  | 62.96 |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.78 | 292.673 | 9773 | 3195.141 | 0.366 |  | 3.81 |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.77 | 3195.141 | 107465 |  | 3195.141 | 4 | 3.95 | -1.2 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.81 | 993.376 | 26476 | 107706.938 | 0.037 |  |  |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.96 | 107706.938 | 3042009 |  | 107706.938 | 4 | 6.5 | 62.4 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.05 | 404.303 | 15376 | 2373.515 | 0.681 |  | 3.08 |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.05 | 2373.515 | 89417 |  | 2373.515 | 4 | 7.06 | 76.5 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte |  |  |  | 88635.797 |  |  |  |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.4 | 252387.266 | 10245247 | 88635.797 | 11.39 |  | 38.99 |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.4 | 88635.797 | 3597215 |  | 88635.797 | 4 | 9.92 | 147.9
20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.88 | 14715.912 | 240678 | 4462.737 | 13.19 |  | 47.49 |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.78 | 210.794 | 7030 | 4462.737 | 0.189 |  | 1.19 |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.78 | 4462.737 | 153013 |  | 4462.737 | 4 | 5.52 | 37.9 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.81 | 642.249 | 12421 | 113437.414 | 0.023 |  |  |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.96 | 113437.414 | 3224395 |  | 113437.414 | 4 | 6.84 | 71.1 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.05 | 382.016 | 15123 | 2397.312 | 0.637 |  | 2.82 |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.05 | 2397.312 | 91363 |  | 2397.312 | 4 | 7.13 | 78.3 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte |  |  |  | 89190.406 |  |  |  |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.4 | 232170.688 | 9274514 | 89190.406 | 10.412 |  | 35.6 |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.4 | 89190.406 | 3620419 |  | 89190.406 | 4 | 9.98 | 149.4
20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.88 | 12368.906 | 206371 | 4447.937 | 11.123 |  | 39.94 |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.78 | 144.765 | 4211 | 4447.937 | 0.13 |  | 0.32 |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.77 | 4447.937 | 152114 |  | 4447.937 | 4 | 5.5 | 37.5 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.81 | 408.254 | 5604 | 110987.766 | 0.015 |  |  |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.96 | 110987.766 | 3056744 |  | 110987.766 | 4 | 6.7 | 67.4 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.05 | 153.025 | 6504 | 2359.211 | 0.259 |  | 0.55 |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.05 | 2359.211 | 88771 |  | 2359.211 | 4 | 7.02 | 75.4 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte |  |  |  | 89103.305 |  |  |  |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.4 | 267149.969 | 10917079 | 89103.305 | 11.993 |  | 41.09 |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.4 | 89103.305 | 3621246 |  | 89103.305 | 4 | 9.97 | 149.2
20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.88 | 12134.313 | 203497 | 4036.639 | 12.024 |  | 43.23 |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.78 | 99.272 | 2968 | 4036.639 | 0.098 |  |  |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.78 | 4036.639 | 139623 |  | 4036.639 | 4 | 4.99 | 24.8 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.81 | 337.916 | 5499 | 99300.031 | 0.014 |  |  |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.96 | 99300.031 | 2693969 |  | 99300.031 | 4 | 5.99 | 49.8 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.05 | 408.419 | 15643 | 2233.017 | 0.732 |  | 3.38 |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.05 | 2233.017 | 83972 |  | 2233.017 | 4 | 6.64 | 66.1 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.06 | 18.349 | 722 | 83877.867 | 0.001 |  |  |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.4 | 231363.469 | 9350834 | 83877.867 | 11.033 |  | 37.76 |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.4 | 83877.867 | 3228451 |  | 83877.867 | 4 | 9.38 | 134.6
20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.88 | 11903.52 | 204453 | 3155.708 | 15.088 |  | 54.43 |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.77 | 136.658 | 4376 | 3155.708 | 0.173 |  | 0.96 |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.78 | 3155.708 | 105566 |  | 3155.708 | 4 | 3.9 | -2.5 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.81 | 462.735 | 7215 | 105545.609 | 0.018 |  |  |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.96 | 105545.609 | 2893700 |  | 105545.609 | 4 | 6.37 | 59.2 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 164.305 | 6189 | 2234.758 | 0.294 |  | 0.76 |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 2234.758 | 81479 |  | 2234.758 | 4 | 6.65 | 66.2 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 7.83 | 256 | 86799.516 | 0 |  |  |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.39 | 257262.188 | 10705604 | 86799.516 | 11.855 |  | 40.61 |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.4 | 86799.516 | 3602688 |  | 86799.516 | 4 | 9.71 | 142.7
20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.88 | 11315.929 | 169106 | 4054.992 | 11.162 |  | 40.08 |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.78 | 110.127 | 3275 | 4054.992 | 0.109 |  | 0 |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.78 | 4054.992 | 139514 |  | 4054.992 | 4 | 5.01 | 25.3 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.81 | 453.139 | 6478 | 104039.031 | 0.017 |  |  |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.96 | 104039.031 | 2808497 |  | 104039.031 | 4 | 6.28 | 56.9 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.05 | 142.847 | 6171 | 2519.11 | 0.227 |  | 0.35 |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.05 | 2519.11 | 93264 |  | 2519.11 | 4 | 7.49 | 87.3 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.07 | 3.867 | 251 | 93797.602 | 0 |  |  |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.4 | 279341.906 | 11378870 | 93797.602 | 11.913 |  | 40.81 |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.4 | 93797.602 | 3755169 |  | 93797.602 | 4 | 10.49 | 162.3
20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.87 | 12633.037 | 196143 | 4225.446 | 11.959 |  | 42.99 |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.78 | 98.905 | 2884 | 4225.446 | 0.094 |  |  |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.78 | 4225.446 | 146406 |  | 4225.446 | 4 | 5.22 | 30.6 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.86 | 381.047 | 4607 | 119178.984 | 0.013 |  |  |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.96 | 119178.984 | 3371710 |  | 119178.984 | 4 | 7.19 | 79.7 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.05 | 149.053 | 5522 | 2439.5 | 0.244 |  | 0.46 |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.05 | 2439.5 | 90275 |  | 2439.5 | 4 | 7.26 | 81.4 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.07 | 16.547 | 987 | 95630.289 | 0.001 |  |  |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.4 | 282657.531 | 11606108 | 95630.289 | 11.823 |  | 40.5 |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.4 | 95630.289 | 3956639 |  | 95630.289 | 4 | 10.7 | 167.4
20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.88 | 11869.145 | 199049 | 4254.767 | 11.158 |  | 40.06 |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.77 | 130.767 | 3718 | 4254.767 | 0.123 |  | 0.21 |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.78 | 4254.767 | 145612 |  | 4254.767 | 4 | 5.26 | 31.5 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.81 | 396.53 | 5932 | 113962.063 | 0.014 |  |  |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.96 | 113962.063 | 3231676 |  | 113962.063 | 4 | 6.87 | 71.9 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.05 | 261.987 | 10292 | 2235.031 | 0.469 |  | 1.81 |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.05 | 2235.031 | 83647 |  | 2235.031 | 4 | 6.65 | 66.2 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.06 | 8.845 | 502 | 88288.648 | 0 |  |  |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.4 | 274644.281 | 11201561 | 88288.648 | 12.443 |  | 42.65 |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.4 | 88288.648 | 3610306 |  | 88288.648 | 4 | 9.88 | 146.9
20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.88 | 11134.584 | 184254 | 4188.782 | 10.633 |  | 38.14 |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.78 | 137.994 | 4125 | 4188.782 | 0.132 |  | 0.34 |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.78 | 4188.782 | 143659 |  | 4188.782 | 4 | 5.18 | 29.5 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.81 | 322.722 | 5508 | 118969.555 | 0.011 |  |  |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.96 | 118969.555 | 3318540 |  | 118969.555 | 4 | 7.18 | 79.4 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.05 | 218.786 | 8347 | 2343.231 | 0.373 |  | 1.23 |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.05 | 2343.231 | 87432 |  | 2343.231 | 4 | 6.97 | 74.3 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.07 | 7.013 | 311 | 88574.773 | 0 |  |  |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.4 | 279947.906 | 11400369 | 88574.773 | 12.642 |  | 43.34 |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.4 | 88574.773 | 3453310 |  | 88574.773 | 4 | 9.91 | 147.7
20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.87 | 10594.069 | 161173 | 4496.683 | 9.424 |  | 33.72 |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.78 | 140.797 | 4509 | 4496.683 | 0.125 |  | 0.25 |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.78 | 4496.683 | 155205 |  | 4496.683 | 4 | 5.56 | 39 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.85 | 377.622 | 4486 | 122033.266 | 0.012 |  |  |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.96 | 122033.266 | 3442648 |  | 122033.266 | 4 | 7.36 | 84 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.05 | 139.54 | 5755 | 2400.613 | 0.233 |  | 0.39 |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.05 | 2400.613 | 90344 |  | 2400.613 | 4 | 7.14 | 78.5 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.07 | 9.717 | 489 | 95312.484 | 0 |  |  |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.4 | 300004.281 | 12091701 | 95312.484 | 12.59 |  | 43.16 |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.4 | 95312.484 | 3746577 |  | 95312.484 | 4 | 10.66 | 166.6
20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.88 | 10354.753 | 164488 | 4153.806 | 9.971 |  | 35.72 |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.78 | 121.203 | 3598 | 4153.806 | 0.117 |  | 0.12 |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.78 | 4153.806 | 139829 |  | 4153.806 | 4 | 5.14 | 28.4 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.85 | 390.746 | 5225 | 110901.977 | 0.014 |  |  |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.96 | 110901.977 | 3062770 |  | 110901.977 | 4 | 6.69 | 67.3 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.05 | 143.163 | 5696 | 2513.042 | 0.228 |  | 0.36 |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.05 | 2513.042 | 92824 |  | 2513.042 | 4 | 7.48 | 86.9 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte |  |  |  | 94841.609 |  |  |  |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.4 | 299780.094 | 11869348 | 94841.609 | 12.643 |  | 43.34 |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.4 | 94841.609 | 3775499 |  | 94841.609 | 4 | 10.61 | 165.2
20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.87 | 9074.265 | 137502 | 3209.497 | 11.309 |  | 40.62 |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.78 | 108.098 | 3082 | 3209.497 | 0.135 |  | 0.39 |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.78 | 3209.497 | 108432 |  | 3209.497 | 4 | 3.97 | -0.8 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.81 | 412.521 | 6320 | 94312.25 | 0.017 |  |  |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.96 | 94312.25 | 2578495 |  | 94312.25 | 4 | 5.69 | 42.2 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.05 | 231.241 | 9183 | 2126.96 | 0.435 |  | 1.6 |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.05 | 2126.96 | 81491 |  | 2126.96 | 4 | 6.33 | 58.2 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.06 | 11.272 | 433 | 76076.914 | 0.001 |  |  |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.4 | 205967.813 | 8300433 | 76076.914 | 10.829 |  | 37.05 |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.4 | 76076.914 | 2958757 |  | 76076.914 | 4 | 8.51 | 112.8
20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.88 | 9561.769 | 153073 | 3877.044 | 9.865 |  | 35.34 |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.78 | 139.492 | 4379 | 3877.044 | 0.144 |  | 0.52 |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.78 | 3877.044 | 131675 |  | 3877.044 | 4 | 4.79 | 19.8 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.8 | 408.999 | 5523 | 101210.508 | 0.016 |  |  |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.96 | 101210.508 | 2857656 |  | 101210.508 | 4 | 6.11 | 52.6 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.05 | 357.515 | 13825 | 2142.216 | 0.668 |  | 3 |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.05 | 2142.216 | 78827 |  | 2142.216 | 4 | 6.37 | 59.3 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.07 | 7.606 | 361 | 76258 | 0 |  |  |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.4 | 206321.797 | 8091617 | 76258 | 10.822 |  | 37.02 |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.4 | 76258 | 2982523 |  | 76258 | 4 | 8.53 | 113.3
20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.88 | 9736.915 | 150985 | 3869.138 | 10.066 |  | 36.07 |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.78 | 106.808 | 3014 | 3869.138 | 0.11 |  | 0.03 |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.78 | 3869.138 | 133521 |  | 3869.138 | 4 | 4.78 | 19.6 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.81 | 356.726 | 5025 | 104400.188 | 0.014 |  |  |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.96 | 104400.188 | 2856369 |  | 104400.188 | 4 | 6.3 | 57.5 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.05 | 167.013 | 6154 | 2372.858 | 0.282 |  | 0.68 |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.05 | 2372.858 | 87556 |  | 2372.858 | 4 | 7.06 | 76.5 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.06 | 8.773 | 389 | 80176.891 | 0 |  |  |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.4 | 225792.75 | 9374209 | 80176.891 | 11.265 |  | 38.56 |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.4 | 80176.891 | 3248475 |  | 80176.891 | 4 | 8.97 | 124.2
20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.87 | 6488.23 | 94317 | 3451.207 | 7.52 |  | 26.76 |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.77 | 94.632 | 2824 | 3451.207 | 0.11 |  | 0.02 |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.78 | 3451.207 | 117553 |  | 3451.207 | 4 | 4.27 | 6.7 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.81 | 404.812 | 5864 | 84769.352 | 0.019 |  |  |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.96 | 84769.352 | 2328900 |  | 84769.352 | 4 | 5.11 | 27.8 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.05 | 315.459 | 11990 | 1960.445 | 0.644 |  | 2.85 |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.05 | 1960.445 | 72293 |  | 1960.445 | 4 | 5.83 | 45.8 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.06 | 0.389 | 70 | 65362.707 | 0 |  |  |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.4 | 189770.188 | 7714787 | 65362.707 | 11.613 |  | 39.77 |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.4 | 65362.707 | 2663057 |  | 65362.707 | 4 | 7.31 | 82.8
20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.88 | 4897.804 | 78547 | 3384.839 | 5.788 |  | 20.43 |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.78 | 83.841 | 2740 | 3384.839 | 0.099 |  |  |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.77 | 3384.839 | 115881 |  | 3384.839 | 4 | 4.19 | 4.6 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.85 | 333.628 | 4606 | 88387.469 | 0.015 |  |  |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.96 | 88387.469 | 2483812 |  | 88387.469 | 4 | 5.33 | 33.3 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.05 | 136.141 | 5084 | 1928.79 | 0.282 |  | 0.69 |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.05 | 1928.79 | 72627 |  | 1928.79 | 4 | 5.74 | 43.4 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.06 | 4.735 | 316 | 66185.563 | 0 |  |  |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.4 | 188079.766 | 7687524 | 66185.563 | 11.367 |  | 38.91 |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.4 | 66185.563 | 2694615 |  | 66185.563 | 4 | 7.4 | 85.1
20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.88 | 7423.349 | 109685 | 3411.54 | 8.704 |  | 31.09 |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.78 | 78.491 | 2813 | 3411.54 | 0.092 |  |  |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.78 | 3411.54 | 117593 |  | 3411.54 | 4 | 4.22 | 5.5 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.86 | 371.009 | 4792 | 75575.227 | 0.02 |  |  |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.96 | 75575.227 | 2013143 |  | 75575.227 | 4 | 4.56 | 14 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.05 | 111.834 | 4367 | 1861.049 | 0.24 |  | 0.43 |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.05 | 1861.049 | 71569 |  | 1861.049 | 4 | 5.54 | 38.4 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.06 | 18.365 | 455 | 64907.961 | 0.001 |  |  |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.4 | 188022.875 | 7440712 | 64907.961 | 11.587 |  | 39.68 |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.4 | 64907.961 | 2536898 |  | 64907.961 | 4 | 7.26 | 81.5
20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.87 | 13458.945 | 179811 | 3946.084 | 13.643 |  | 49.15 |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.78 | 372.564 | 13071 | 3946.084 | 0.378 |  | 3.98 |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.78 | 3946.084 | 135784 |  | 3946.084 | 4 | 4.88 | 22 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.81 | 1741.671 | 49424 | 98270.578 | 0.071 |  |  |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.96 | 98270.578 | 2691723 |  | 98270.578 | 4 | 5.93 | 48.2 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.05 | 1087.664 | 41574 | 2277.866 | 1.91 |  | 10.45 |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.05 | 2277.866 | 87482 |  | 2277.866 | 4 | 6.78 | 69.4 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.06 | 64.871 | 2246 | 78984.391 | 0.003 |  |  |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.4 | 228470.469 | 9286333 | 78984.391 | 11.57 |  | 39.62 |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.4 | 78984.391 | 3150051 |  | 78984.391 | 4 | 8.84 | 120.9
20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.88 | 13394.592 | 210981 | 4015.599 | 13.343 |  | 48.05 |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.78 | 110.375 | 3443 | 4015.599 | 0.11 |  | 0.02 |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.78 | 4015.599 | 139237 |  | 4015.599 | 4 | 4.96 | 24.1 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.85 | 359.241 | 4577 | 98628.742 | 0.015 |  |  |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.96 | 98628.742 | 2774183 |  | 98628.742 | 4 | 5.95 | 48.7 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.05 | 109.222 | 3935 | 1961.049 | 0.223 |  | 0.33 |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.05 | 1961.049 | 74969 |  | 1961.049 | 4 | 5.83 | 45.8 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte |  |  |  | 54687.828 |  |  |  |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.4 | 145300.688 | 5942828 | 54687.828 | 10.628 |  | 36.35 |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.4 | 54687.828 | 2268070 |  | 54687.828 | 4 | 6.12 | 52.9
20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.88 | 12774.548 | 184155 | 3722.044 | 13.729 |  | 49.46 |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.78 | 78.054 | 2592 | 3722.044 | 0.084 |  |  |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.78 | 3722.044 | 127148 |  | 3722.044 | 4 | 4.6 | 15.1 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.86 | 327.177 | 5240 | 106868.266 | 0.012 |  |  |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.96 | 106868.266 | 3033461 |  | 106868.266 | 4 | 6.45 | 61.2 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.05 | 134.736 | 5715 | 2366.501 | 0.228 |  | 0.36 |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.05 | 2366.501 | 90032 |  | 2366.501 | 4 | 7.04 | 76 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.07 | 6.07 | 347 | 63527.227 | 0 |  |  |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.4 | 171869.766 | 6608780 | 63527.227 | 10.822 |  | 37.02 |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.4 | 63527.227 | 2347999 |  | 63527.227 | 4 | 7.11 | 77.7
20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.87 | 13906.821 | 198513 | 3719.442 | 14.956 |  | 53.95 |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.78 | 138.96 | 4052 | 3719.442 | 0.149 |  | 0.61 |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.78 | 3719.442 | 129467 |  | 3719.442 | 4 | 4.6 | 15 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.81 | 471.128 | 8876 | 104541.266 | 0.018 |  |  |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.96 | 104541.266 | 2937157 |  | 104541.266 | 4 | 6.31 | 57.7 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.05 | 200.381 | 7915 | 2257.612 | 0.355 |  | 1.12 |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.05 | 2257.612 | 86402 |  | 2257.612 | 4 | 6.72 | 67.9 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.07 | 6.733 | 385 | 81925.484 | 0 |  |  |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.4 | 232343.172 | 9388697 | 81925.484 | 11.344 |  | 38.83 |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.4 | 81925.484 | 3419540 |  | 81925.484 | 4 | 9.16 | 129.1
20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.88 | 12754.647 | 175426 | 3898.412 | 13.087 |  | 47.12 |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.78 | 126.729 | 4009 | 3898.412 | 0.13 |  | 0.32 |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.78 | 3898.412 | 134486 |  | 3898.412 | 4 | 4.82 | 20.5 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.81 | 416.801 | 6653 | 112556.063 | 0.015 |  |  |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.96 | 112556.063 | 3156701 |  | 112556.063 | 4 | 6.79 | 69.8 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 271.229 | 10772 | 2389.278 | 0.454 |  | 1.72 |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 2389.278 | 89739 |  | 2389.278 | 4 | 7.11 | 77.7 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 12.484 | 452 | 77283.281 | 0.001 |  |  |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.4 | 238910.734 | 9909917 | 77283.281 | 12.365 |  | 42.38 |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.4 | 77283.281 | 3151650 |  | 77283.281 | 4 | 8.65 | 116.1
20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.88 | 13496.068 | 184112 | 3840.235 | 14.058 |  | 50.66 |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.78 | 122.661 | 3841 | 3840.235 | 0.128 |  | 0.29 |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.78 | 3840.235 | 132810 |  | 3840.235 | 4 | 4.75 | 18.7 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.81 | 276.12 | 4768 | 104859.008 | 0.011 |  |  |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.96 | 104859.008 | 2910590 |  | 104859.008 | 4 | 6.33 | 58.1 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.05 | 147.119 | 5598 | 2282.505 | 0.258 |  | 0.54 |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.05 | 2282.505 | 85777 |  | 2282.505 | 4 | 6.79 | 69.7 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.06 | 2.375 | 131 | 77620.695 | 0 |  |  |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.4 | 228967.797 | 9634266 | 77620.695 | 11.799 |  | 40.41 |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.4 | 77620.695 | 3186330 |  | 77620.695 | 4 | 8.68 | 117.1
20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.88 | 14466.164 | 202071 | 3796.952 | 15.24 |  | 54.99 |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.78 | 115.056 | 3650 | 3796.952 | 0.121 |  | 0.19 |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.78 | 3796.952 | 130307 |  | 3796.952 | 4 | 4.69 | 17.4 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.81 | 414.222 | 6606 | 105383.109 | 0.016 |  |  |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.96 | 105383.109 | 2935355 |  | 105383.109 | 4 | 6.36 | 58.9 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.05 | 167.121 | 6732 | 2296.952 | 0.291 |  | 0.74 |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.05 | 2296.952 | 84946 |  | 2296.952 | 4 | 6.83 | 70.8 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.06 | 4.094 | 252 | 86903.203 | 0 |  |  |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.4 | 257647.172 | 10769629 | 86903.203 | 11.859 |  | 40.62 |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.4 | 86903.203 | 3634190 |  | 86903.203 | 4 | 9.72 | 143
20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.87 | 13676.497 | 198516 | 3988.584 | 13.716 |  | 49.41 |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.77 | 98.341 | 3529 | 3988.584 | 0.099 |  |  |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.77 | 3988.584 | 136537 |  | 3988.584 | 4 | 4.93 | 23.3 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.85 | 410.053 | 5828 | 99879.406 | 0.016 |  |  |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.96 | 99879.406 | 2791654 |  | 99879.406 | 4 | 6.03 | 50.6 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.05 | 147.658 | 5726 | 2078.18 | 0.284 |  | 0.7 |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.05 | 2078.18 | 78341 |  | 2078.18 | 4 | 6.18 | 54.5 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte |  |  |  | 64245.207 |  |  |  |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.4 | 179824.234 | 6961351 | 64245.207 | 11.196 |  | 38.32 |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.4 | 64245.207 | 2644733 |  | 64245.207 | 4 | 7.19 | 79.7
20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.87 | 13672.899 | 183759 | 4089.779 | 13.373 |  | 48.16 |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.78 | 149.401 | 5058 | 4089.779 | 0.146 |  | 0.56 |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.77 | 4089.779 | 139738 |  | 4089.779 | 4 | 5.06 | 26.4 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.81 | 350.653 | 7143 | 94358.172 | 0.015 |  |  |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.96 | 94358.172 | 2525923 |  | 94358.172 | 4 | 5.69 | 42.3 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.05 | 219.107 | 8462 | 2319.462 | 0.378 |  | 1.26 |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.05 | 2319.462 | 87625 |  | 2319.462 | 4 | 6.9 | 72.5 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.06 | 20.189 | 1028 | 84458.141 | 0.001 |  |  |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.4 | 246133.094 | 10120721 | 84458.141 | 11.657 |  | 39.92 |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.4 | 84458.141 | 3363683 |  | 84458.141 | 4 | 9.45 | 136.2
20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.87 | 13819.075 | 185596 | 4257.981 | 12.982 |  | 46.73 |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.78 | 185.739 | 6213 | 4257.981 | 0.174 |  | 0.98 |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.78 | 4257.981 | 146948 |  | 4257.981 | 4 | 5.26 | 31.6 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.8 | 772.868 | 20550 | 116692.453 | 0.026 |  |  |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.96 | 116692.453 | 3338579 |  | 116692.453 | 4 | 7.04 | 76 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.05 | 691.918 | 26391 | 2365.194 | 1.17 |  | 6.01 |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.05 | 2365.194 | 86293 |  | 2365.194 | 4 | 7.04 | 75.9 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.06 | 68.218 | 3260 | 87830.766 | 0.003 |  |  |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.4 | 278046.969 | 11063472 | 87830.766 | 12.663 |  | 43.41 |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.4 | 87830.766 | 3422934 |  | 87830.766 | 4 | 9.83 | 145.6
20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.86 | 14415.857 | 197477 | 4107.876 | 14.037 |  | 50.59 |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.78 | 297.333 | 9442 | 4107.876 | 0.29 |  | 2.68 |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.77 | 4107.876 | 139990 |  | 4107.876 | 4 | 5.08 | 27 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.81 | 947.536 | 23810 | 107368.117 | 0.035 |  |  |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.96 | 107368.117 | 3028922 |  | 107368.117 | 4 | 6.48 | 61.9 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.05 | 669.348 | 25473 | 2378.302 | 1.126 |  | 5.75 |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.05 | 2378.302 | 91198 |  | 2378.302 | 4 | 7.07 | 76.9 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.07 | 17.151 | 687 | 82120.75 | 0.001 |  |  |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.4 | 275360.031 | 10737526 | 82120.75 | 13.412 |  | 46.01 |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.4 | 82120.75 | 3182716 |  | 82120.75 | 4 | 9.19 | 129.7
20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.87 | 13589.694 | 179296 | 3864.161 | 14.067 |  | 50.7 |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.78 | 246.998 | 8342 | 3864.161 | 0.256 |  | 2.18 |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.78 | 3864.161 | 132419 |  | 3864.161 | 4 | 4.78 | 19.4 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.81 | 1016.708 | 27596 | 109868.875 | 0.037 |  |  |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.96 | 109868.875 | 3072305 |  | 109868.875 | 4 | 6.63 | 65.7 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.05 | 469.471 | 17560 | 2149.042 | 0.874 |  | 4.23 |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.05 | 2149.042 | 80393 |  | 2149.042 | 4 | 6.39 | 59.8 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.07 | 17.723 | 611 | 76499.594 | 0.001 |  |  |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.4 | 245384.563 | 10366516 | 76499.594 | 12.831 |  | 43.99 |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.4 | 76499.594 | 3151665 |  | 76499.594 | 4 | 8.56 | 113.9
20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.87 | 13366.214 | 177634 | 3543.577 | 15.088 |  | 54.43 |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.78 | 86.154 | 2791 | 3543.577 | 0.097 |  |  |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.78 | 3543.577 | 120707 |  | 3543.577 | 4 | 4.38 | 9.5 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.81 | 381.625 | 5824 | 102934.406 | 0.015 |  |  |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.96 | 102934.406 | 2925716 |  | 102934.406 | 4 | 6.21 | 55.2 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.05 | 152.881 | 6501 | 2177.948 | 0.281 |  | 0.68 |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.05 | 2177.948 | 82550 |  | 2177.948 | 4 | 6.48 | 62 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte |  |  |  | 73346.039 |  |  |  |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.4 | 199021.969 | 7933975 | 73346.039 | 10.854 |  | 37.13 |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.4 | 73346.039 | 2836304 |  | 73346.039 | 4 | 8.2 | 105.1
20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.87 | 12986.323 | 189758 | 3555.943 | 14.608 |  | 52.68 |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.78 | 115.902 | 3813 | 3555.943 | 0.13 |  | 0.32 |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.78 | 3555.943 | 121173 |  | 3555.943 | 4 | 4.4 | 9.9 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.85 | 415.396 | 5471 | 95880.813 | 0.017 |  |  |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.96 | 95880.813 | 2743191 |  | 95880.813 | 4 | 5.78 | 44.6 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.05 | 356.513 | 14002 | 1966.663 | 0.725 |  | 3.34 |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.05 | 1966.663 | 74380 |  | 1966.663 | 4 | 5.85 | 46.2 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.06 | 14.537 | 793 | 68280.641 | 0.001 |  |  |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.39 | 195954.266 | 7834061 | 68280.641 | 11.479 |  | 39.3 |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.4 | 68280.641 | 2797197 |  | 68280.641 | 4 | 7.64 | 91
20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.86 | 13166.602 | 186700 | 3988.081 | 13.206 |  | 47.55 |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.78 | 137.6 | 4290 | 3988.081 | 0.138 |  | 0.44 |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.78 | 3988.081 | 136185 |  | 3988.081 | 4 | 4.93 | 23.3 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.81 | 350.126 | 4576 | 111986.016 | 0.013 |  |  |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.96 | 111986.016 | 3190599 |  | 111986.016 | 4 | 6.76 | 68.9 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.05 | 221.256 | 8132 | 2330.105 | 0.38 |  | 1.27 |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.05 | 2330.105 | 87302 |  | 2330.105 | 4 | 6.93 | 73.3 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.06 | 11.696 | 573 | 73499.289 | 0.001 |  |  |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.4 | 209089.75 | 8472446 | 73499.289 | 11.379 |  | 38.96 |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.4 | 73499.289 | 2960621 |  | 73499.289 | 4 | 8.22 | 105.5
20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.87 | 11126.391 | 155600 | 3506.183 | 12.693 |  | 45.68 |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.78 | 56.173 | 1957 | 3506.183 | 0.064 |  |  |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.78 | 3506.183 | 119804 |  | 3506.183 | 4 | 4.34 | 8.4 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte |  |  |  | 84980.797 |  |  |  |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.96 | 84980.797 | 2345692 |  | 84980.797 | 4 | 5.13 | 28.2 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.05 | 118.698 | 4868 | 1788.846 | 0.265 |  | 0.58 |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.05 | 1788.846 | 68641 |  | 1788.846 | 4 | 5.32 | 33 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte |  |  |  | 55749.684 |  |  |  |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.4 | 163660.906 | 6955397 | 55749.684 | 11.743 |  | 40.22 |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.4 | 55749.684 | 2264068 |  | 55749.684 | 4 | 6.24 | 55.9
20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.87 | 9472.6 | 127395 | 2791.078 | 13.576 |  | 48.9 |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.77 | 80.953 | 2729 | 2791.078 | 0.116 |  | 0.11 |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.77 | 2791.078 | 94932 |  | 2791.078 | 4 | 3.45 | -13.7 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.85 | 292.296 | 3730 | 64520.02 | 0.018 |  |  |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.96 | 64520.02 | 1765655 |  | 64520.02 | 4 | 3.89 | -2.7 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.05 | 89.453 | 3749 | 1405.372 | 0.255 |  | 0.52 |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.05 | 1405.372 | 53963 |  | 1405.372 | 4 | 4.18 | 4.5 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte |  |  |  | 30788.695 |  |  |  |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.4 | 87596.305 | 3600304 | 30788.695 | 11.38 |  | 38.96 |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.4 | 30788.695 | 1220278 |  | 30788.695 | 4 | 3.44 | -13.9
20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.87 | 9704.444 | 128645 | 3117.41 | 12.452 |  | 44.79 |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.78 | 62.378 | 2190 | 3117.41 | 0.08 |  |  |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.78 | 3117.41 | 106241 |  | 3117.41 | 4 | 3.85 | -3.6 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.85 | 367.758 | 5330 | 77287.742 | 0.019 |  |  |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.96 | 77287.742 | 2053044 |  | 77287.742 | 4 | 4.66 | 16.6 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.05 | 130.351 | 5363 | 1826.592 | 0.285 |  | 0.7 |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.05 | 1826.592 | 70966 |  | 1826.592 | 4 | 5.43 | 35.8 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.07 | 7.596 | 468 | 54407.445 | 0.001 |  |  |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.4 | 144037.141 | 5669520 | 54407.445 | 10.59 |  | 36.21 |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.4 | 54407.445 | 2019079 |  | 54407.445 | 4 | 6.09 | 52.2
20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.87 | 13138.3 | 187084 | 4056.419 | 12.956 |  | 46.64 |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.78 | 189.653 | 5817 | 4056.419 | 0.187 |  | 1.16 |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.78 | 4056.419 | 140110 |  | 4056.419 | 4 | 5.02 | 25.4 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.81 | 457.152 | 6935 | 96807.188 | 0.019 |  |  |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.97 | 96807.188 | 2668441 |  | 96807.188 | 4 | 5.84 | 46 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.05 | 207.095 | 7876 | 2059.017 | 0.402 |  | 1.41 |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.05 | 2059.017 | 77275 |  | 2059.017 | 4 | 6.12 | 53.1 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte |  |  |  | 76163.773 |  |  |  |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.4 | 198084.219 | 8400724 | 76163.773 | 10.403 |  | 35.57 |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.4 | 76163.773 | 3220525 |  | 76163.773 | 4 | 8.52 | 113
20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.87 | 14517.875 | 206204 | 4559.46 | 12.736 |  | 45.83 |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.78 | 241.631 | 8170 | 4559.46 | 0.212 |  | 1.53 |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.78 | 4559.46 | 157257 |  | 4559.46 | 4 | 5.64 | 40.9 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.81 | 475.469 | 8669 | 110827.492 | 0.017 |  |  |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.96 | 110827.492 | 3028426 |  | 110827.492 | 4 | 6.69 | 67.1 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.05 | 353.006 | 13268 | 2482.213 | 0.569 |  | 2.41 |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.05 | 2482.213 | 93606 |  | 2482.213 | 4 | 7.38 | 84.6 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.07 | 1.967 | 217 | 92251.773 | 0 |  |  |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.4 | 241833.688 | 9651093 | 92251.773 | 10.486 |  | 35.85 |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.4 | 92251.773 | 3670430 |  | 92251.773 | 4 | 10.32 | 158
20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.87 | 15841.149 | 227470 | 4493.075 | 14.103 |  | 50.83 |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.78 | 290.765 | 9335 | 4493.075 | 0.259 |  | 2.22 |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.78 | 4493.075 | 154908 |  | 4493.075 | 4 | 5.56 | 38.9 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.81 | 554.069 | 8920 | 106620.406 | 0.021 |  |  |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.96 | 106620.406 | 3040845 |  | 106620.406 | 4 | 6.43 | 60.8 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 232.939 | 8411 | 2293.903 | 0.406 |  | 1.43 |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 2293.903 | 85682 |  | 2293.903 | 4 | 6.82 | 70.6 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 4.014 | 281 | 83986.945 | 0 |  |  |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.4 | 218327.75 | 8751073 | 83986.945 | 10.398 |  | 35.55 |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.4 | 83986.945 | 3315792 |  | 83986.945 | 4 | 9.4 | 134.9
20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.87 | 12003.916 | 162939 | 4114.258 | 11.671 |  | 41.94 |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.78 | 190.399 | 6349 | 4114.258 | 0.185 |  | 1.13 |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.78 | 4114.258 | 141391 |  | 4114.258 | 4 | 5.09 | 27.2 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.81 | 469.105 | 6650 | 86723.508 | 0.022 |  |  |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.96 | 86723.508 | 2258546 |  | 86723.508 | 4 | 5.23 | 30.8 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.05 | 220.973 | 8680 | 2212.084 | 0.4 |  | 1.39 |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.05 | 2212.084 | 83409 |  | 2212.084 | 4 | 6.58 | 64.5 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.06 | 10.213 | 486 | 86446.32 | 0 |  |  |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.4 | 238265.75 | 9592995 | 86446.32 | 11.025 |  | 37.73 |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.4 | 86446.32 | 3432628 |  | 86446.32 | 4 | 9.67 | 141.8
20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.86 | 13315.447 | 183072 | 4004.459 | 13.301 |  | 47.9 |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.78 | 243.066 | 7101 | 4004.459 | 0.243 |  | 1.99 |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.78 | 4004.459 | 137439 |  | 4004.459 | 4 | 4.95 | 23.8 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte |  |  |  | 97677.023 |  |  |  |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.96 | 97677.023 | 2650361 |  | 97677.023 | 4 | 5.89 | 47.3 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.05 | 205.306 | 8231 | 2292.081 | 0.358 |  | 1.14 |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.05 | 2292.081 | 84135 |  | 2292.081 | 4 | 6.82 | 70.4 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte |  |  |  | 89926.844 |  |  |  |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.4 | 262954.25 | 10368118 | 89926.844 | 11.696 |  | 40.06 |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.4 | 89926.844 | 3532382 |  | 89926.844 | 4 | 10.06 | 151.5
20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.87 | 13054.155 | 177255 | 3985.281 | 13.102 |  | 47.17 |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.78 | 176.193 | 5558 | 3985.281 | 0.177 |  | 1.01 |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.77 | 3985.281 | 135627 |  | 3985.281 | 4 | 4.93 | 23.2 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.86 | 673.771 | 8906 | 100329.82 | 0.027 |  |  |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.96 | 100329.82 | 2834738 |  | 100329.82 | 4 | 6.05 | 51.3 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.05 | 193.423 | 7058 | 2233.479 | 0.346 |  | 1.07 |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.05 | 2233.479 | 87834 |  | 2233.479 | 4 | 6.64 | 66.1 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.09 | 16.609 | 464 | 91128.219 | 0.001 |  |  |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.4 | 266459.938 | 11257610 | 91128.219 | 11.696 |  | 40.06 |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.4 | 91128.219 | 3617198 |  | 91128.219 | 4 | 10.19 | 154.9
20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.86 | 62.169 | 761 | 5093.202 | 0.049 |  |  |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.78 | 6391.303 | 220071 | 5093.202 | 5.019 |  | 72.57 |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.78 | 5093.202 | 174541 |  | 5093.202 | 4 | 6.3 | 57.4 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.81 | 1968.833 | 59138 | 120155.836 | 0.066 |  |  |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.96 | 120155.836 | 3332802 |  | 120155.836 | 4 | 7.25 | 81.2 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.05 | 4630.355 | 168087 | 2534.763 | 7.307 |  | 42.83 |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.05 | 2534.763 | 92974 |  | 2534.763 | 4 | 7.54 | 88.5 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.06 | 81.691 | 3339 | 107375.094 | 0.003 |  |  |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.4 | 27795.627 | 1163048 | 107375.094 | 1.035 |  | 3.05 |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.4 | 107375.094 | 4487870 |  | 107375.094 | 4 | 12.01 | 200.3
20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte |  |  |  | 4939.262 |  |  |  |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.78 | 5906.037 | 203264 | 4939.262 | 4.783 |  | 69.08 |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.78 | 4939.262 | 170266 |  | 4939.262 | 4 | 6.11 | 52.7 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.81 | 1659.804 | 49492 | 112531.641 | 0.059 |  |  |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.96 | 112531.641 | 3182520 |  | 112531.641 | 4 | 6.79 | 69.7 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.05 | 4343.149 | 162971 | 2424.725 | 7.165 |  | 41.98 |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.05 | 2424.725 | 87860 |  | 2424.725 | 4 | 7.21 | 80.3 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.06 | 71.185 | 2151 | 104266.414 | 0.003 |  |  |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.4 | 4972.751 | 197217 | 104266.414 | 0.191 |  | 0.12 |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.4 | 104266.414 | 4273133 |  | 104266.414 | 4 | 11.66 | 191.6
20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.85 | 18.73 | 503 | 5233.792 | 0.014 |  |  |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.78 | 6012.225 | 207327 | 5233.792 | 4.595 |  | 66.3 |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.78 | 5233.792 | 181293 |  | 5233.792 | 4 | 6.47 | 61.8 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.81 | 1071.008 | 26239 | 118805.617 | 0.036 |  |  |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.96 | 118805.617 | 3380132 |  | 118805.617 | 4 | 7.17 | 79.2 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.05 | 4442.118 | 170191 | 2412.854 | 7.364 |  | 43.17 |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.05 | 2412.854 | 89939 |  | 2412.854 | 4 | 7.18 | 79.4 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.06 | 105.631 | 3989 | 105538.766 | 0.004 |  |  |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.4 | 5474.621 | 219701 | 105538.766 | 0.207 |  | 0.18 |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.4 | 105538.766 | 4219927 |  | 105538.766 | 4 | 11.81 | 195.2
20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.9 | 39.308 | 527 | 4518.646 | 0.035 |  |  |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.78 | 4942.229 | 169810 | 4518.646 | 4.375 |  | 63.05 |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.78 | 4518.646 | 155494 |  | 4518.646 | 4 | 5.59 | 39.7 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.8 | 1369.747 | 38750 | 102610.383 | 0.053 |  |  |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.96 | 102610.383 | 2811674 |  | 102610.383 | 4 | 6.19 | 54.8 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.05 | 3689.881 | 140619 | 2272.698 | 6.494 |  | 37.95 |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.05 | 2272.698 | 84725 |  | 2272.698 | 4 | 6.76 | 69 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.07 | 57.227 | 2282 | 84474.836 | 0.003 |  |  |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.4 | 2002.274 | 89927 | 84474.836 | 0.095 |  |  |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.4 | 84474.836 | 3404028 |  | 84474.836 | 4 | 9.45 | 136.2
20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.86 | 0.79 | 104 | 4595.066 | 0.001 |  |  |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.78 | 4874.581 | 169161 | 4595.066 | 4.243 |  | 61.1 |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.78 | 4595.066 | 158691 |  | 4595.066 | 4 | 5.68 | 42 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.85 | 313.797 | 3798 | 102214.219 | 0.012 |  |  |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.96 | 102214.219 | 2882559 |  | 102214.219 | 4 | 6.17 | 54.2 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.05 | 3383.043 | 128491 | 2150.592 | 6.292 |  | 36.74 |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.05 | 2150.592 | 81038 |  | 2150.592 | 4 | 6.4 | 59.9 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.07 | 6.926 | 306 | 87116.578 | 0 |  |  |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.4 | 2397.686 | 94151 | 87116.578 | 0.11 |  |  |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.4 | 87116.578 | 3474043 |  | 87116.578 | 4 | 9.75 | 143.6
20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.85 | 9.745 | 292 | 4516.958 | 0.009 |  |  |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.78 | 3923.402 | 133112 | 4516.958 | 3.474 |  | 49.74 |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.78 | 4516.958 | 153628 |  | 4516.958 | 4 | 5.58 | 39.6 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.81 | 374.854 | 4879 | 99535.336 | 0.015 |  |  |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.96 | 99535.336 | 2741889 |  | 99535.336 | 4 | 6 | 50.1 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.05 | 3382.356 | 128897 | 2036.234 | 6.644 |  | 38.85 |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.05 | 2036.234 | 75303 |  | 2036.234 | 4 | 6.06 | 51.4 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.06 | 3.727 | 259 | 83468.359 | 0 |  |  |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.4 | 2073.612 | 89065 | 83468.359 | 0.099 |  |  |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.4 | 83468.359 | 3439578 |  | 83468.359 | 4 | 9.34 | 133.4
20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte |  |  |  | 4562.208 |  |  |  |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.78 | 4543.188 | 154753 | 4562.208 | 3.983 |  | 57.26 |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.78 | 4562.208 | 155516 |  | 4562.208 | 4 | 5.64 | 41 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.8 | 544.265 | 9791 | 103807.086 | 0.021 |  |  |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.96 | 103807.086 | 2960972 |  | 103807.086 | 4 | 6.26 | 56.6 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.05 | 3519.728 | 133013 | 2160.931 | 6.515 |  | 38.08 |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.05 | 2160.931 | 80933 |  | 2160.931 | 4 | 6.43 | 60.7 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.06 | 12.387 | 330 | 89927.414 | 0.001 |  |  |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.4 | 2722.367 | 114256 | 89927.414 | 0.121 |  |  |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.4 | 89927.414 | 3486817 |  | 89927.414 | 4 | 10.06 | 151.5
20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.88 | 0.404 | 57 | 4277.345 | 0 |  |  |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.78 | 4894.958 | 166429 | 4277.345 | 4.578 |  | 66.04 |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.78 | 4277.345 | 144960 |  | 4277.345 | 4 | 5.29 | 32.2 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.8 | 482.408 | 8308 | 99959.609 | 0.019 |  |  |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.96 | 99959.609 | 2879778 |  | 99959.609 | 4 | 6.03 | 50.8 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.05 | 3685.783 | 137360 | 2170.902 | 6.791 |  | 39.74 |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.05 | 2170.902 | 80815 |  | 2170.902 | 4 | 6.46 | 61.4 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.06 | 10.963 | 199 | 86421.602 | 0.001 |  |  |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.4 | 2252.211 | 97347 | 86421.602 | 0.104 |  |  |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.4 | 86421.602 | 3416719 |  | 86421.602 | 4 | 9.67 | 141.7
20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.87 | 15.454 | 583 | 4028.417 | 0.015 |  |  |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.78 | 4958.275 | 168396 | 4028.417 | 4.923 |  | 71.15 |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.78 | 4028.417 | 137600 |  | 4028.417 | 4 | 4.98 | 24.5 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.81 | 1167.067 | 33017 | 100739.109 | 0.046 |  |  |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.96 | 100739.109 | 2878863 |  | 100739.109 | 4 | 6.08 | 51.9 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.05 | 3722.396 | 140864 | 2075.98 | 7.172 |  | 42.02 |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.05 | 2075.98 | 76015 |  | 2075.98 | 4 | 6.18 | 54.4 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.07 | 13.604 | 430 | 91154.789 | 0.001 |  |  |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.39 | 3439.174 | 128733 | 91154.789 | 0.151 |  |  |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.4 | 91154.789 | 3585221 |  | 91154.789 | 4 | 10.2 | 154.9
20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.84 | 4.492 | 168 | 4201.701 | 0.004 |  |  |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.78 | 5418.749 | 187049 | 4201.701 | 5.159 |  | 74.63 |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.78 | 4201.701 | 142567 |  | 4201.701 | 4 | 5.2 | 29.9 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.81 | 260.508 | 9475 | 95184.141 | 0.011 |  |  |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.96 | 95184.141 | 2623040 |  | 95184.141 | 4 | 5.74 | 43.6 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.05 | 3417.414 | 128123 | 2070.364 | 6.603 |  | 38.6 |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.05 | 2070.364 | 77569 |  | 2070.364 | 4 | 6.16 | 54 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.03 | 16.884 | 380 | 89026.438 | 0.001 |  |  |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.4 | 2381.84 | 88870 | 89026.438 | 0.107 |  |  |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.4 | 89026.438 | 3511176 |  | 89026.438 | 4 | 9.96 | 149
20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.85 | 3.701 | 145 | 4522.662 | 0.003 |  |  |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.78 | 5020.021 | 174032 | 4522.662 | 4.44 |  | 64.01 |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.78 | 4522.662 | 156282 |  | 4522.662 | 4 | 5.59 | 39.8 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.81 | 3652.715 | 113771 | 106147.617 | 0.138 |  |  |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.96 | 106147.617 | 3012649 |  | 106147.617 | 4 | 6.4 | 60.1 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.05 | 3877.35 | 146603 | 2153.034 | 7.204 |  | 42.21 |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.05 | 2153.034 | 81167 |  | 2153.034 | 4 | 6.4 | 60.1 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.06 | 13.921 | 612 | 89150.375 | 0.001 |  |  |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.4 | 6998.725 | 272006 | 89150.375 | 0.314 |  | 0.55 |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.4 | 89150.375 | 3327173 |  | 89150.375 | 4 | 9.97 | 149.3
20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte |  |  |  | 4864.867 |  |  |  |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.78 | 5217.081 | 179277 | 4864.867 | 4.29 |  | 61.79 |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.78 | 4864.867 | 169224 |  | 4864.867 | 4 | 6.02 | 50.4 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.81 | 1195.365 | 32755 | 98089.266 | 0.049 |  |  |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.96 | 98089.266 | 2678359 |  | 98089.266 | 4 | 5.92 | 47.9 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.05 | 4665.503 | 182306 | 2243.382 | 8.319 |  | 48.9 |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.05 | 2243.382 | 84356 |  | 2243.382 | 4 | 6.67 | 66.8 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.07 | 29.097 | 1028 | 89808.383 | 0.001 |  |  |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.4 | 14859.413 | 603578 | 89808.383 | 0.662 |  | 1.75 |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.4 | 89808.383 | 3663026 |  | 89808.383 | 4 | 10.05 | 151.2
20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte |  |  |  | 4130.82 |  |  |  |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.78 | 4317.092 | 148947 | 4130.82 | 4.18 |  | 60.17 |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.78 | 4130.82 | 140590 |  | 4130.82 | 4 | 5.11 | 27.7 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.81 | 339.214 | 5143 | 85510.195 | 0.016 |  |  |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.96 | 85510.195 | 2336976 |  | 85510.195 | 4 | 5.16 | 29 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.05 | 3521.64 | 135865 | 2093.272 | 6.729 |  | 39.37 |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.05 | 2093.272 | 76894 |  | 2093.272 | 4 | 6.23 | 55.7 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.06 | 5.17 | 309 | 78967.586 | 0 |  |  |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.4 | 1691.597 | 72046 | 78967.586 | 0.086 |  |  |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.4 | 78967.586 | 3313229 |  | 78967.586 | 4 | 8.83 | 120.8
20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.87 | 11.221 | 400 | 4509.017 | 0.01 |  |  |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.78 | 5081.018 | 175617 | 4509.017 | 4.507 |  | 65.01 |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.78 | 4509.017 | 155305 |  | 4509.017 | 4 | 5.58 | 39.4 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.81 | 492.186 | 8924 | 92018.813 | 0.021 |  |  |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.96 | 92018.813 | 2410122 |  | 92018.813 | 4 | 5.55 | 38.8 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.05 | 3449.08 | 131631 | 2165.184 | 6.372 |  | 37.22 |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.05 | 2165.184 | 82844 |  | 2165.184 | 4 | 6.44 | 61 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.03 | 0.403 | 62 | 83951.203 | 0 |  |  |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.4 | 2561.379 | 93938 | 83951.203 | 0.122 |  |  |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.4 | 83951.203 | 3239414 |  | 83951.203 | 4 | 9.39 | 134.8
20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.87 | 13.982 | 290 | 4377.103 | 0.013 |  |  |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.78 | 4593.416 | 157865 | 4377.103 | 4.198 |  | 60.43 |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.78 | 4377.103 | 150190 |  | 4377.103 | 4 | 5.41 | 35.3 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.81 | 1344.97 | 39408 | 87029.469 | 0.062 |  |  |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.97 | 87029.469 | 2377545 |  | 87029.469 | 4 | 5.25 | 31.3 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.05 | 3022.028 | 109144 | 2151.335 | 5.619 |  | 32.7 |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.05 | 2151.335 | 79762 |  | 2151.335 | 4 | 6.4 | 60 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.06 | 68.997 | 2377 | 82722.727 | 0.003 |  |  |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.39 | 4321.598 | 173179 | 82722.727 | 0.209 |  | 0.18 |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.4 | 82722.727 | 3247101 |  | 82722.727 | 4 | 9.25 | 131.3
20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte |  |  |  | 3747.828 |  |  |  |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.78 | 3518.875 | 119857 | 3747.828 | 3.756 |  | 53.9 |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.78 | 3747.828 | 127274 |  | 3747.828 | 4 | 4.63 | 15.8 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.81 | 20960.096 | 702734 | 89270.031 | 0.939 |  | 2.25 |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.96 | 89270.031 | 2472065 |  | 89270.031 | 4 | 5.39 | 34.6 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.05 | 2831.834 | 110285 | 1989.004 | 5.695 |  | 33.16 |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.05 | 1989.004 | 74074 |  | 1989.004 | 4 | 5.92 | 47.9 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.06 | 0.798 | 93 | 80214.852 | 0 |  |  |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.4 | 2641.276 | 106059 | 80214.852 | 0.132 |  |  |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.4 | 80214.852 | 3212792 |  | 80214.852 | 4 | 8.97 | 124.3
20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte |  |  |  | 4038.739 |  |  |  |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.78 | 3942.875 | 135332 | 4038.739 | 3.905 |  | 56.1 |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.78 | 4038.739 | 139785 |  | 4038.739 | 4 | 4.99 | 24.8 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.86 | 364.35 | 4413 | 85731.852 | 0.017 |  |  |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.96 | 85731.852 | 2440697 |  | 85731.852 | 4 | 5.17 | 29.3 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.05 | 2885.134 | 111975 | 1868.877 | 6.175 |  | 36.04 |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.05 | 1868.877 | 72443 |  | 1868.877 | 4 | 5.56 | 39 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte |  |  |  | 77124.359 |  |  |  |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.4 | 2297.828 | 98469 | 77124.359 | 0.119 |  |  |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.4 | 77124.359 | 3314510 |  | 77124.359 | 4 | 8.63 | 115.7
20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.87 | 21.437 | 444 | 3789.513 | 0.023 |  |  |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.78 | 3520.303 | 121928 | 3789.513 | 3.716 |  | 53.31 |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.78 | 3789.513 | 131680 |  | 3789.513 | 4 | 4.69 | 17.1 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.8 | 304.308 | 3837 | 82351.266 | 0.015 |  |  |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.96 | 82351.266 | 2320337 |  | 82351.266 | 4 | 4.97 | 24.2 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.05 | 2826.412 | 111737 | 1927.043 | 5.867 |  | 34.19 |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.05 | 1927.043 | 72505 |  | 1927.043 | 4 | 5.73 | 43.3 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.07 | 2.654 | 165 | 74714.813 | 0 |  |  |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.4 | 1735.724 | 69695 | 74714.813 | 0.093 |  |  |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.4 | 74714.813 | 3064155 |  | 74714.813 | 4 | 8.36 | 108.9
20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.88 | 4.288 | 199 | 4403.589 | 0.004 |  |  |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.78 | 4863.822 | 167123 | 4403.589 | 4.418 |  | 63.69 |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.78 | 4403.589 | 152070 |  | 4403.589 | 4 | 5.44 | 36.1 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.8 | 331.232 | 5083 | 96676.672 | 0.014 |  |  |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.96 | 96676.672 | 2624358 |  | 96676.672 | 4 | 5.83 | 45.8 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.05 | 3666.114 | 142234 | 2335.834 | 6.278 |  | 36.66 |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.05 | 2335.834 | 89409 |  | 2335.834 | 4 | 6.95 | 73.7 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.06 | 13.18 | 449 | 99319.828 | 0.001 |  |  |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.4 | 4481.813 | 193533 | 99319.828 | 0.181 |  | 0.08 |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.4 | 99319.828 | 4043231 |  | 99319.828 | 4 | 11.11 | 177.8
20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.87 | 11.475 | 363 | 4598.22 | 0.01 |  |  |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.78 | 5016.161 | 172989 | 4598.22 | 4.364 |  | 62.88 |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.78 | 4598.22 | 158526 |  | 4598.22 | 4 | 5.69 | 42.1 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.8 | 363.727 | 6758 | 106880.938 | 0.014 |  |  |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.96 | 106880.938 | 3036519 |  | 106880.938 | 4 | 6.45 | 61.2 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.05 | 3464.063 | 129174 | 2227.993 | 6.219 |  | 36.3 |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.05 | 2227.993 | 82124 |  | 2227.993 | 4 | 6.63 | 65.7 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte |  |  |  | 94711.563 |  |  |  |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.4 | 7830.961 | 317929 | 94711.563 | 0.331 |  | 0.6 |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.4 | 94711.563 | 3750901 |  | 94711.563 | 4 | 10.59 | 164.9
20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.85 | 10.074 | 350 | 4689.85 | 0.009 |  |  |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.78 | 5046.894 | 172993 | 4689.85 | 4.305 |  | 62.01 |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.78 | 4689.85 | 160258 |  | 4689.85 | 4 | 5.8 | 45 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.8 | 1032.07 | 27234 | 109563.195 | 0.038 |  |  |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.96 | 109563.195 | 3144449 |  | 109563.195 | 4 | 6.61 | 65.2 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.05 | 4187.168 | 160122 | 2176.578 | 7.695 |  | 45.16 |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.05 | 2176.578 | 83746 |  | 2176.578 | 4 | 6.47 | 61.9 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.06 | 7.088 | 331 | 90822.211 | 0 |  |  |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.4 | 9123.27 | 397031 | 90822.211 | 0.402 |  | 0.85 |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.4 | 90822.211 | 3627143 |  | 90822.211 | 4 | 10.16 | 154
20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.88 | 6.985 | 141 | 4132.267 | 0.007 |  |  |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.78 | 4255.095 | 146051 | 4132.267 | 4.119 |  | 59.26 |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.78 | 4132.267 | 141081 |  | 4132.267 | 4 | 5.11 | 27.7 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.81 | 433.407 | 8770 | 99443.953 | 0.017 |  |  |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.97 | 99443.953 | 2784842 |  | 99443.953 | 4 | 6 | 50 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.05 | 3326.948 | 127478 | 2189.358 | 6.078 |  | 35.46 |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.05 | 2189.358 | 78834 |  | 2189.358 | 4 | 6.51 | 62.8 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.07 | 5.175 | 302 | 84107.078 | 0 |  |  |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.4 | 2338.676 | 92326 | 84107.078 | 0.111 |  |  |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.4 | 84107.078 | 3258718 |  | 84107.078 | 4 | 9.41 | 135.2
20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.86 | 50.12 | 1129 | 4125.632 | 0.049 |  |  |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.78 | 3951.575 | 134733 | 4125.632 | 3.831 |  | 55.01 |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.78 | 4125.632 | 141685 |  | 4125.632 | 4 | 5.1 | 27.5 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.86 | 279.778 | 3623 | 106413.328 | 0.011 |  |  |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.96 | 106413.328 | 3078344 |  | 106413.328 | 4 | 6.42 | 60.5 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.05 | 3317.284 | 126490 | 2193.385 | 6.05 |  | 35.29 |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.05 | 2193.385 | 83567 |  | 2193.385 | 4 | 6.52 | 63.1 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte |  |  |  | 82892.469 |  |  |  |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.4 | 1933.72 | 74930 | 82892.469 | 0.093 |  |  |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.4 | 82892.469 | 3306039 |  | 82892.469 | 4 | 9.27 | 131.8
20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.88 | 5.404 | 192 | 4422.781 | 0.005 |  |  |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.78 | 4628.658 | 158950 | 4422.781 | 4.186 |  | 60.26 |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.78 | 4422.781 | 153183 |  | 4422.781 | 4 | 5.47 | 36.7 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte |  |  |  | 109031.305 |  |  |  |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.97 | 109031.305 | 3128109 |  | 109031.305 | 4 | 6.58 | 64.4 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.05 | 3435.218 | 132714 | 2207.604 | 6.224 |  | 36.33 |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.05 | 2207.604 | 80918 |  | 2207.604 | 4 | 6.57 | 64.2 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte |  |  |  | 89697.148 |  |  |  |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.4 | 2124.865 | 95781 | 89697.148 | 0.095 |  |  |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.4 | 89697.148 | 3831743 |  | 89697.148 | 4 | 10.03 | 150.8
20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.85 | 8.917 | 245 | 4699.705 | 0.008 |  |  |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.78 | 4871.632 | 168775 | 4699.705 | 4.146 |  | 59.67 |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.78 | 4699.705 | 162117 |  | 4699.705 | 4 | 5.81 | 45.3 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.81 | 346.74 | 5135 | 100056 | 0.014 |  |  |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.96 | 100056 | 2828201 |  | 100056 | 4 | 6.04 | 50.9 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.05 | 3227.542 | 125488 | 2186.832 | 5.904 |  | 34.41 |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.05 | 2186.832 | 81592 |  | 2186.832 | 4 | 6.5 | 62.6 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.07 | 8.217 | 261 | 84959.609 | 0 |  |  |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.4 | 2743.245 | 102340 | 84959.609 | 0.129 |  |  |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.4 | 84959.609 | 3222198 |  | 84959.609 | 4 | 9.5 | 137.6
20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.88 | 44.112 | 881 | 4633.228 | 0.038 |  |  |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.78 | 4619.891 | 158711 | 4633.228 | 3.988 |  | 57.34 |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.78 | 4633.228 | 160681 |  | 4633.228 | 4 | 5.73 | 43.2 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.8 | 493.649 | 10680 | 89716.656 | 0.022 |  |  |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.96 | 89716.656 | 2419546 |  | 89716.656 | 4 | 5.41 | 35.3 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.05 | 3479.264 | 131384 | 2072.819 | 6.714 |  | 39.27 |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.05 | 2072.819 | 77608 |  | 2072.819 | 4 | 6.17 | 54.1 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte |  |  |  | 86283.852 |  |  |  |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.4 | 3430.022 | 134885 | 86283.852 | 0.159 |  | 0.01 |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.4 | 86283.852 | 3504027 |  | 86283.852 | 4 | 9.65 | 141.3
20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.88 | 9.015 | 235 | 4931.43 | 0.007 |  |  |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.78 | 5559.856 | 190760 | 4931.43 | 4.51 |  | 65.04 |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.78 | 4931.43 | 170328 |  | 4931.43 | 4 | 6.1 | 52.4 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.85 | 335.928 | 4351 | 109559.086 | 0.012 |  |  |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.97 | 109559.086 | 2959963 |  | 109559.086 | 4 | 6.61 | 65.2 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.05 | 3257.239 | 117830 | 2269.267 | 5.741 |  | 33.44 |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.05 | 2269.267 | 85774 |  | 2269.267 | 4 | 6.75 | 68.8 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.07 | 13.263 | 417 | 100046.359 | 0.001 |  |  |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.4 | 2008.583 | 88863 | 100046.359 | 0.08 |  |  |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.4 | 100046.359 | 4210013 |  | 100046.359 | 4 | 11.19 | 179.8
20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.85 | 10.279 | 210 | 4408.869 | 0.009 |  |  |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.78 | 4127.725 | 143210 | 4408.869 | 3.745 |  | 53.74 |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.78 | 4408.869 | 152599 |  | 4408.869 | 4 | 5.45 | 36.3 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.81 | 556.298 | 12640 | 90430.555 | 0.025 |  |  |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.96 | 90430.555 | 2440787 |  | 90430.555 | 4 | 5.46 | 36.4 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.05 | 3758.575 | 140254 | 2124.282 | 7.077 |  | 41.45 |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.05 | 2124.282 | 80191 |  | 2124.282 | 4 | 6.32 | 58 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.07 | 7.463 | 338 | 95674.383 | 0 |  |  |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.4 | 14150.011 | 605117 | 95674.383 | 0.592 |  | 1.51 |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.4 | 95674.383 | 4084776 |  | 95674.383 | 4 | 10.7 | 167.6
20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte |  |  |  | 4910.336 |  |  |  |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.78 | 4334.655 | 148483 | 4910.336 | 3.531 |  | 50.58 |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.78 | 4910.336 | 169194 |  | 4910.336 | 4 | 6.07 | 51.8 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.81 | 745.601 | 16891 | 102643.82 | 0.029 |  |  |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.97 | 102643.82 | 2885297 |  | 102643.82 | 4 | 6.19 | 54.8 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.05 | 3681.078 | 139784 | 2198.032 | 6.699 |  | 39.18 |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.05 | 2198.032 | 83709 |  | 2198.032 | 4 | 6.54 | 63.5 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte |  |  |  | 93804.273 |  |  |  |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.4 | 6172.626 | 255818 | 93804.273 | 0.263 |  | 0.37 |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.4 | 93804.273 | 3910266 |  | 93804.273 | 4 | 10.49 | 162.3
20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.85 | 10.55 | 300 | 4659.491 | 0.009 |  |  |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.78 | 4811.418 | 164882 | 4659.491 | 4.13 |  | 59.43 |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.78 | 4659.491 | 160745 |  | 4659.491 | 4 | 5.76 | 44 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.81 | 662.704 | 14954 | 96604.727 | 0.027 |  |  |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.96 | 96604.727 | 2642798 |  | 96604.727 | 4 | 5.83 | 45.7 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.05 | 3579.895 | 132544 | 2234.033 | 6.41 |  | 37.45 |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.05 | 2234.033 | 83528 |  | 2234.033 | 4 | 6.65 | 66.1 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.11 | 1.75 | 98 | 75229.672 | 0 |  |  |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.4 | 5040.559 | 197915 | 75229.672 | 0.268 |  | 0.39 |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.4 | 75229.672 | 2819570 |  | 75229.672 | 4 | 8.42 | 110.4
20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.83 | 9.293 | 285 | 3744.223 | 0.01 |  |  |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.78 | 4283.556 | 146551 | 3744.223 | 4.576 |  | 66.02 |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.78 | 3744.223 | 127381 |  | 3744.223 | 4 | 4.63 | 15.7 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.8 | 629.296 | 15328 | 105342.867 | 0.024 |  |  |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.96 | 105342.867 | 3033331 |  | 105342.867 | 4 | 6.35 | 58.9 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.05 | 3289.579 | 127687 | 2107.887 | 6.242 |  | 36.44 |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.05 | 2107.887 | 80161 |  | 2107.887 | 4 | 6.27 | 56.8 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.06 | 38.406 | 1610 | 87124.805 | 0.002 |  |  |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.4 | 2383.649 | 96799 | 87124.805 | 0.109 |  |  |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.4 | 87124.805 | 3374102 |  | 87124.805 | 4 | 9.75 | 143.7
20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte |  |  |  | 4034.394 |  |  |  |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.78 | 4557.383 | 156709 | 4034.394 | 4.519 |  | 65.17 |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.78 | 4034.394 | 138969 |  | 4034.394 | 4 | 4.99 | 24.7 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.85 | 294.01 | 4664 | 98521.195 | 0.012 |  |  |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.97 | 98521.195 | 2867227 |  | 98521.195 | 4 | 5.94 | 48.6 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.05 | 2823.59 | 105276 | 2019.512 | 5.593 |  | 32.54 |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.05 | 2019.512 | 74692 |  | 2019.512 | 4 | 6.01 | 50.2 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.02 | 2.931 | 191 | 77005.422 | 0 |  |  |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.4 | 1627.43 | 65057 | 77005.422 | 0.085 |  |  |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.4 | 77005.422 | 3313557 |  | 77005.422 | 4 | 8.61 | 115.4
20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.87 | 19.84 | 431 | 4172.714 | 0.019 |  |  |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.78 | 4135.498 | 140199 | 4172.714 | 3.964 |  | 56.98 |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.78 | 4172.714 | 143367 |  | 4172.714 | 4 | 5.16 | 29 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.81 | 312.739 | 4898 | 87807.195 | 0.014 |  |  |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.97 | 87807.195 | 2354495 |  | 87807.195 | 4 | 5.3 | 32.4 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.05 | 3351.329 | 126030 | 2174.987 | 6.163 |  | 35.97 |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.05 | 2174.987 | 82892 |  | 2174.987 | 4 | 6.47 | 61.7 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.08 | 9.604 | 240 | 88369.641 | 0 |  |  |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.4 | 3482.942 | 128910 | 88369.641 | 0.158 |  | 0 |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.4 | 88369.641 | 3280962 |  | 88369.641 | 4 | 9.89 | 147.1
20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte |  |  |  | 3922.801 |  |  |  |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.78 | 3757.041 | 129864 | 3922.801 | 3.831 |  | 55.01 |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.78 | 3922.801 | 137534 |  | 3922.801 | 4 | 4.85 | 21.3 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.86 | 289.661 | 4097 | 84955.523 | 0.014 |  |  |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.97 | 84955.523 | 2334746 |  | 84955.523 | 4 | 5.13 | 28.1 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.05 | 2774.165 | 104724 | 1885.64 | 5.885 |  | 34.3 |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.05 | 1885.64 | 69423 |  | 1885.64 | 4 | 5.61 | 40.2 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.06 | 0.168 | 31 | 76210.016 | 0 |  |  |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.4 | 1741.286 | 67916 | 76210.016 | 0.091 |  |  |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.4 | 76210.016 | 3155056 |  | 76210.016 | 4 | 8.53 | 113.1
20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.85 | 29.131 | 366 | 3270.064 | 0.036 |  |  |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.78 | 3527.55 | 121269 | 3270.064 | 4.315 |  | 62.16 |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.78 | 3270.064 | 113843 |  | 3270.064 | 4 | 4.04 | 1.1 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte |  |  |  | 83049.625 |  |  |  |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.96 | 83049.625 | 2385833 |  | 83049.625 | 4 | 5.01 | 25.3 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.05 | 2413.183 | 95545 | 1554.347 | 6.21 |  | 36.25 |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.05 | 1554.347 | 59618 |  | 1554.347 | 4 | 4.62 | 15.6 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.09 | 0.105 | 18 | 52721.07 | 0 |  |  |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.4 | 1187.428 | 48088 | 52721.07 | 0.09 |  |  |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.4 | 52721.07 | 2267876 |  | 52721.07 | 4 | 5.9 | 47.4
20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte |  |  |  | 4049.156 |  |  |  |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.78 | 4092.351 | 141067 | 4049.156 | 4.043 |  | 58.14 |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.78 | 4049.156 | 140243 |  | 4049.156 | 4 | 5.01 | 25.2 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.85 | 285.751 | 3659 | 75152.219 | 0.015 |  |  |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.97 | 75152.219 | 2040597 |  | 75152.219 | 4 | 4.53 | 13.3 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.05 | 2684.16 | 102251 | 1868.361 | 5.747 |  | 33.47 |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.05 | 1868.361 | 70814 |  | 1868.361 | 4 | 5.56 | 38.9 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.07 | 9.736 | 352 | 55817.023 | 0.001 |  |  |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.4 | 1416.868 | 47558 | 55817.023 | 0.102 |  |  |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.4 | 55817.023 | 2277350 |  | 55817.023 | 4 | 6.24 | 56.1
20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.87 | 38.099 | 1027 | 4646.139 | 0.033 |  |  |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.78 | 4634.335 | 157508 | 4646.139 | 3.99 |  | 57.36 |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.78 | 4646.139 | 159487 |  | 4646.139 | 4 | 5.74 | 43.6 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.81 | 370.412 | 5965 | 102671.5 | 0.014 |  |  |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.96 | 102671.5 | 2832084 |  | 102671.5 | 4 | 6.19 | 54.8 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.05 | 3195.274 | 116858 | 2224.932 | 5.744 |  | 33.46 |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.05 | 2224.932 | 81925 |  | 2224.932 | 4 | 6.62 | 65.5 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte |  |  |  | 99776.836 |  |  |  |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.4 | 2355.618 | 92222 | 99776.836 | 0.094 |  |  |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.4 | 99776.836 | 4067719 |  | 99776.836 | 4 | 11.16 | 179
20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte |  |  |  | 4670.868 |  |  |  |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.78 | 4617.606 | 157457 | 4670.868 | 3.954 |  | 56.83 |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.78 | 4670.868 | 159954 |  | 4670.868 | 4 | 5.78 | 44.4 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.81 | 368.678 | 5262 | 101879.82 | 0.014 |  |  |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.96 | 101879.82 | 2750035 |  | 101879.82 | 4 | 6.15 | 53.7 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.05 | 3959.748 | 152075 | 2450.528 | 6.464 |  | 37.77 |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.05 | 2450.528 | 94042 |  | 2450.528 | 4 | 7.29 | 82.2 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.06 | 55.627 | 2127 | 101979.219 | 0.002 |  |  |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.4 | 3240.659 | 122106 | 101979.219 | 0.127 |  |  |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.4 | 101979.219 | 4124754 |  | 101979.219 | 4 | 11.41 | 185.2
20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.88 | 34.705 | 698 | 4760.106 | 0.029 |  |  |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.78 | 5855.985 | 201916 | 4760.106 | 4.921 |  | 71.12 |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.78 | 4760.106 | 162646 |  | 4760.106 | 4 | 5.89 | 47.1 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.81 | 929.424 | 25504 | 119917.711 | 0.031 |  |  |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.97 | 119917.711 | 3374345 |  | 119917.711 | 4 | 7.23 | 80.9 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.05 | 4245.902 | 161078 | 2422.206 | 7.012 |  | 41.06 |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.05 | 2422.206 | 92834 |  | 2422.206 | 4 | 7.21 | 80.1 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.07 | 3.523 | 218 | 114534.242 | 0 |  |  |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.4 | 16277.225 | 691768 | 114534.242 | 0.568 |  | 1.43 |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.4 | 114534.242 | 4796443 |  | 114534.242 | 4 | 12.81 | 220.3
20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.87 | 16.086 | 360 | 4087.066 | 0.016 |  |  |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.78 | 4184.689 | 143212 | 4087.066 | 4.096 |  | 58.92 |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.78 | 4087.066 | 139402 |  | 4087.066 | 4 | 5.05 | 26.3 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.81 | 397.327 | 5254 | 97705.281 | 0.016 |  |  |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.96 | 97705.281 | 2789053 |  | 97705.281 | 4 | 5.89 | 47.4 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.05 | 3195.347 | 124660 | 2046.279 | 6.246 |  | 36.47 |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.05 | 2046.279 | 77503 |  | 2046.279 | 4 | 6.09 | 52.2 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.06 | 8.106 | 373 | 79844.398 | 0 |  |  |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.4 | 3266.747 | 141507 | 79844.398 | 0.164 |  | 0.02 |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.4 | 79844.398 | 3250808 |  | 79844.398 | 4 | 8.93 | 123.3
20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.88 | 16.546 | 403 | 4565.343 | 0.014 |  |  |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.78 | 4455.831 | 152479 | 4565.343 | 3.904 |  | 56.09 |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.78 | 4565.343 | 156691 |  | 4565.343 | 4 | 5.64 | 41.1 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.81 | 355.733 | 5451 | 102602.477 | 0.014 |  |  |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.96 | 102602.477 | 2908390 |  | 102602.477 | 4 | 6.19 | 54.7 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.05 | 3478.008 | 137993 | 2164.459 | 6.427 |  | 37.55 |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.05 | 2164.459 | 83044 |  | 2164.459 | 4 | 6.44 | 61 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.06 | 9.177 | 373 | 86026.953 | 0 |  |  |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.4 | 3671.382 | 155139 | 86026.953 | 0.171 |  | 0.05 |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.4 | 86026.953 | 3398443 |  | 86026.953 | 4 | 9.62 | 140.6
20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.85 | 20.354 | 479 | 4533.528 | 0.018 |  |  |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.78 | 4756.871 | 162191 | 4533.528 | 4.197 |  | 60.42 |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.78 | 4533.528 | 154364 |  | 4533.528 | 4 | 5.61 | 40.1 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.81 | 345.67 | 4848 | 91857.102 | 0.015 |  |  |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.96 | 91857.102 | 2594836 |  | 91857.102 | 4 | 5.54 | 38.5 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.05 | 3191.466 | 119308 | 2047.973 | 6.233 |  | 36.39 |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.05 | 2047.973 | 77307 |  | 2047.973 | 4 | 6.09 | 52.3 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.06 | 3.263 | 202 | 90566.922 | 0 |  |  |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.4 | 3701.034 | 150082 | 90566.922 | 0.163 |  | 0.02 |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.4 | 90566.922 | 3678126 |  | 90566.922 | 4 | 10.13 | 153.3
20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte |  |  |  | 4314.995 |  |  |  |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.78 | 65.511 | 2011 | 4314.995 | 0.061 |  |  |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.78 | 4314.995 | 149357 |  | 4314.995 | 4 | 5.34 | 33.4 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.81 | 211102.141 | 6984337 | 122806.492 | 6.876 |  | 19.09 |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.97 | 122806.492 | 3541354 |  | 122806.492 | 4 | 7.41 | 85.2 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.05 | 444.989 | 16010 | 2325.502 | 0.765 |  | 3.58 |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.05 | 2325.502 | 83837 |  | 2325.502 | 4 | 6.92 | 72.9 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.07 | 84765.813 | 3149459 | 97325.922 | 3.484 |  | 57.55 |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.4 | 4395.498 | 171055 | 97325.922 | 0.181 |  | 0.08 |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.4 | 97325.922 | 4022685 |  | 97325.922 | 4 | 10.89 | 172.2
20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.9 | 8.644 | 217 | 4169.22 | 0.008 |  |  |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.78 | 90.844 | 2571 | 4169.22 | 0.087 |  |  |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.78 | 4169.22 | 142828 |  | 4169.22 | 4 | 5.15 | 28.9 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.81 | 308503.531 | 10401536 | 105540.742 | 11.692 |  | 32.75 |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.96 | 105540.742 | 2950542 |  | 105540.742 | 4 | 6.37 | 59.2 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.05 | 169.976 | 6440 | 2318.157 | 0.293 |  | 0.75 |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.05 | 2318.157 | 89454 |  | 2318.157 | 4 | 6.9 | 72.4 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.06 | 89158.578 | 3440956 | 102662.727 | 3.474 |  | 57.38 |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.4 | 4938.487 | 182785 | 102662.727 | 0.192 |  | 0.12 |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.4 | 102662.727 | 3738008 |  | 102662.727 | 4 | 11.48 | 187.1
20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.88 | 5.067 | 195 | 4343.747 | 0.005 |  |  |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.78 | 59.138 | 1829 | 4343.747 | 0.054 |  |  |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.78 | 4343.747 | 149701 |  | 4343.747 | 4 | 5.37 | 34.3 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.8 | 356524.719 | 12119818 | 108226.414 | 13.177 |  | 36.97 |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.96 | 108226.414 | 2970949 |  | 108226.414 | 4 | 6.53 | 63.2 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.05 | 77.56 | 3248 | 2446.042 | 0.127 |  |  |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.05 | 2446.042 | 96028 |  | 2446.042 | 4 | 7.28 | 81.9 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.06 | 89658.922 | 3423202 | 99524.109 | 3.604 |  | 59.55 |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.4 | 3017.959 | 122122 | 99524.109 | 0.121 |  |  |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.4 | 99524.109 | 3978316 |  | 99524.109 | 4 | 11.13 | 178.3
20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte |  |  |  | 3994.895 |  |  |  |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.78 | 42.996 | 1394 | 3994.895 | 0.043 |  |  |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.77 | 3994.895 | 137003 |  | 3994.895 | 4 | 4.94 | 23.5 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.81 | 384793.406 | 13078994 | 98816.094 | 15.576 |  | 43.77 |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.96 | 98816.094 | 2714665 |  | 98816.094 | 4 | 5.96 | 49 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.05 | 65.056 | 3165 | 2371.027 | 0.11 |  |  |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.05 | 2371.027 | 93883 |  | 2371.027 | 4 | 7.05 | 76.3 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.06 | 83136.398 | 3056431 | 104638.984 | 3.178 |  | 52.44 |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.4 | 2407.487 | 92725 | 104638.984 | 0.092 |  |  |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.4 | 104638.984 | 4058627 |  | 104638.984 | 4 | 11.71 | 192.6
20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte |  |  |  | 3869.739 |  |  |  |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.78 | 39.354 | 1325 | 3869.739 | 0.041 |  |  |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.78 | 3869.739 | 132127 |  | 3869.739 | 4 | 4.78 | 19.6 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.81 | 283668.125 | 9436496 | 114010.602 | 9.952 |  | 27.82 |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.96 | 114010.602 | 3199621 |  | 114010.602 | 4 | 6.88 | 71.9 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.05 | 59.29 | 2368 | 2403.863 | 0.099 |  |  |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.05 | 2403.863 | 91589 |  | 2403.863 | 4 | 7.15 | 78.8 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.06 | 73273.953 | 2781368 | 95722.523 | 3.062 |  | 50.5 |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.4 | 2357.775 | 96783 | 95722.523 | 0.099 |  |  |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.4 | 95722.523 | 3923375 |  | 95722.523 | 4 | 10.71 | 167.7
20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte |  |  |  | 4094.355 |  |  |  |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.78 | 65.363 | 1799 | 4094.355 | 0.064 |  |  |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.78 | 4094.355 | 141335 |  | 4094.355 | 4 | 5.06 | 26.6 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.81 | 270666.656 | 8983324 | 111625.508 | 9.699 |  | 27.1 |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.97 | 111625.508 | 3109118 |  | 111625.508 | 4 | 6.73 | 68.3 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.05 | 51.517 | 2109 | 2416.499 | 0.085 |  |  |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.05 | 2416.499 | 92311 |  | 2416.499 | 4 | 7.19 | 79.7 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.07 | 91674.398 | 3524289 | 104475.281 | 3.51 |  | 57.98 |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.4 | 2659.664 | 107623 | 104475.281 | 0.102 |  |  |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.4 | 104475.281 | 4375782 |  | 104475.281 | 4 | 11.69 | 192.2
20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.86 | 51.112 | 1346 | 4088.134 | 0.05 |  |  |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.78 | 164.218 | 5442 | 4088.134 | 0.161 |  | 0.77 |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.78 | 4088.134 | 141052 |  | 4088.134 | 4 | 5.05 | 26.4 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.8 | 298650.938 | 9971744 | 111956.539 | 10.67 |  | 29.85 |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.96 | 111956.539 | 3172459 |  | 111956.539 | 4 | 6.75 | 68.8 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.05 | 134.341 | 5716 | 2279.891 | 0.236 |  | 0.41 |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.05 | 2279.891 | 87427 |  | 2279.891 | 4 | 6.78 | 69.5 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.06 | 69076.906 | 2644574 | 96895.539 | 2.852 |  | 46.99 |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.4 | 6274.098 | 254249 | 96895.539 | 0.259 |  | 0.35 |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.4 | 96895.539 | 3608709 |  | 96895.539 | 4 | 10.84 | 171
20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.82 | 8.796 | 183 | 3656.892 | 0.01 |  |  |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.78 | 66.106 | 1966 | 3656.892 | 0.072 |  |  |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.78 | 3656.892 | 126601 |  | 3656.892 | 4 | 4.52 | 13 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.81 | 333344 | 11203420 | 104588.813 | 12.749 |  | 35.75 |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.96 | 104588.813 | 2995732 |  | 104588.813 | 4 | 6.31 | 57.7 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.05 | 120.986 | 4204 | 2295.882 | 0.211 |  | 0.26 |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.05 | 2295.882 | 91146 |  | 2295.882 | 4 | 6.83 | 70.7 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.06 | 80887.133 | 3075320 | 90586.703 | 3.572 |  | 59.02 |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.4 | 2641.203 | 106814 | 90586.703 | 0.117 |  |  |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.4 | 90586.703 | 3675424 |  | 90586.703 | 4 | 10.13 | 153.3
20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.9 | 3.336 | 179 | 4115.619 | 0.003 |  |  |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.77 | 67.688 | 2189 | 4115.619 | 0.066 |  |  |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.78 | 4115.619 | 141480 |  | 4115.619 | 4 | 5.09 | 27.2 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.8 | 274019.813 | 9070789 | 87999.492 | 12.456 |  | 34.92 |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.96 | 87999.492 | 2312692 |  | 87999.492 | 4 | 5.31 | 32.7 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.05 | 66.306 | 2725 | 2274.982 | 0.117 |  |  |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.05 | 2274.982 | 87382 |  | 2274.982 | 4 | 6.77 | 69.2 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.06 | 88501.672 | 3452261 | 85220.258 | 4.154 |  | 68.74 |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.4 | 2226.196 | 79075 | 85220.258 | 0.104 |  |  |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.4 | 85220.258 | 3006224 |  | 85220.258 | 4 | 9.53 | 138.3
20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.89 | 5.356 | 204 | 3668.877 | 0.006 |  |  |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.78 | 58.988 | 1991 | 3668.877 | 0.064 |  |  |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.78 | 3668.877 | 125748 |  | 3668.877 | 4 | 4.54 | 13.4 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.81 | 299645.75 | 10020487 | 97360.305 | 12.311 |  | 34.51 |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.97 | 97360.305 | 2799244 |  | 97360.305 | 4 | 5.87 | 46.8 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.05 | 85.625 | 3468 | 1978.307 | 0.173 |  | 0.03 |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.05 | 1978.307 | 76393 |  | 1978.307 | 4 | 5.88 | 47.1 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.06 | 83890.813 | 3278919 | 87958.188 | 3.815 |  | 63.08 |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.4 | 2858.172 | 106053 | 87958.188 | 0.13 |  |  |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.4 | 87958.188 | 3322711 |  | 87958.188 | 4 | 9.84 | 146
20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.9 | 4.782 | 169 | 3204.234 | 0.006 |  |  |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.78 | 114.769 | 3960 | 3204.234 | 0.143 |  | 0.51 |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.78 | 3204.234 | 110651 |  | 3204.234 | 4 | 3.96 | -1 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.81 | 235749.125 | 7878546 | 84831.938 | 11.116 |  | 31.12 |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.96 | 84831.938 | 2370620 |  | 84831.938 | 4 | 5.12 | 27.9 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.05 | 98.728 | 3906 | 1924.4 | 0.205 |  | 0.22 |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.05 | 1924.4 | 73973 |  | 1924.4 | 4 | 5.72 | 43.1 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.06 | 74294.75 | 2865533 | 73946.117 | 4.019 |  | 66.48 |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.4 | 3113.741 | 112647 | 73946.117 | 0.168 |  | 0.04 |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.4 | 73946.117 | 2831281 |  | 73946.117 | 4 | 8.27 | 106.8
20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.84 | 3.179 | 97 | 3915.398 | 0.003 |  |  |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.77 | 205.511 | 7108 | 3915.398 | 0.21 |  | 1.5 |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.78 | 3915.398 | 135372 |  | 3915.398 | 4 | 4.84 | 21 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.8 | 323462.594 | 11000032 | 92662.234 | 13.963 |  | 39.19 |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.96 | 92662.234 | 2609777 |  | 92662.234 | 4 | 5.59 | 39.7 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.05 | 162.785 | 6639 | 1966.173 | 0.331 |  | 0.98 |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.05 | 1966.173 | 76220 |  | 1966.173 | 4 | 5.85 | 46.2 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.06 | 88411.766 | 3434985 | 88629.414 | 3.99 |  | 66 |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.4 | 5364.324 | 218230 | 88629.414 | 0.242 |  | 0.3 |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.4 | 88629.414 | 3533247 |  | 88629.414 | 4 | 9.91 | 147.9
20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.84 | 3.194 | 60 | 3898.605 | 0.003 |  |  |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.78 | 229.48 | 8090 | 3898.605 | 0.235 |  | 1.88 |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.78 | 3898.605 | 133591 |  | 3898.605 | 4 | 4.82 | 20.5 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.81 | 283991.438 | 9699914 | 103418.438 | 10.984 |  | 30.75 |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.96 | 103418.438 | 3004115 |  | 103418.438 | 4 | 6.24 | 56 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.05 | 71.583 | 3014 | 2244.717 | 0.128 |  |  |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.05 | 2244.717 | 87807 |  | 2244.717 | 4 | 6.68 | 66.9 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.06 | 82879.844 | 2973760 | 93904.141 | 3.53 |  | 58.33 |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.4 | 2398.448 | 106629 | 93904.141 | 0.102 |  |  |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.4 | 93904.141 | 4091944 |  | 93904.141 | 4 | 10.5 | 162.6
20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.88 | 14.696 | 311 | 4061.239 | 0.014 |  |  |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.78 | 54.196 | 1918 | 4061.239 | 0.053 |  |  |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.78 | 4061.239 | 140995 |  | 4061.239 | 4 | 5.02 | 25.5 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.81 | 303358.25 | 10312001 | 106341.625 | 11.411 |  | 31.95 |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.97 | 106341.625 | 3036697 |  | 106341.625 | 4 | 6.42 | 60.4 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.05 | 45.632 | 2059 | 2253.734 | 0.081 |  |  |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.05 | 2253.734 | 90122 |  | 2253.734 | 4 | 6.7 | 67.6 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.06 | 83693.422 | 3138311 | 99092.422 | 3.378 |  | 55.79 |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.4 | 2452.679 | 97232 | 99092.422 | 0.099 |  |  |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.4 | 99092.422 | 3960975 |  | 99092.422 | 4 | 11.08 | 177.1
20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte |  |  |  | 3517.021 |  |  |  |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.78 | 73.433 | 2244 | 3517.021 | 0.084 |  |  |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.78 | 3517.021 | 121705 |  | 3517.021 | 4 | 4.35 | 8.7 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.81 | 279995.094 | 9373390 | 92977.766 | 12.046 |  | 33.76 |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.97 | 92977.766 | 2623362 |  | 92977.766 | 4 | 5.61 | 40.2 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.05 | 68.436 | 2773 | 2225.482 | 0.123 |  |  |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.05 | 2225.482 | 85507 |  | 2225.482 | 4 | 6.62 | 65.5 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.07 | 85406.219 | 3482632 | 102824.828 | 3.322 |  | 54.85 |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.4 | 3060.941 | 112753 | 102824.828 | 0.119 |  |  |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.4 | 102824.828 | 3514491 |  | 102824.828 | 4 | 11.5 | 187.6
20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte |  |  |  | 3202.326 |  |  |  |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.78 | 29.845 | 1106 | 3202.326 | 0.037 |  |  |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.78 | 3202.326 | 110024 |  | 3202.326 | 4 | 3.96 | -1 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.81 | 210909.141 | 7091518 | 79338.781 | 10.633 |  | 29.75 |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.97 | 79338.781 | 2209920 |  | 79338.781 | 4 | 4.79 | 19.7 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.05 | 68.165 | 2481 | 1927.321 | 0.141 |  |  |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.05 | 1927.321 | 76090 |  | 1927.321 | 4 | 5.73 | 43.3 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.07 | 60181.426 | 2252242 | 74625.852 | 3.226 |  | 53.24 |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.4 | 2650.645 | 109312 | 74625.852 | 0.142 |  |  |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.4 | 74625.852 | 2961373 |  | 74625.852 | 4 | 8.35 | 108.7
20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte |  |  |  | 3661.565 |  |  |  |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.78 | 51.159 | 1557 | 3661.565 | 0.056 |  |  |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.78 | 3661.565 | 126574 |  | 3661.565 | 4 | 4.53 | 13.2 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.81 | 219247.406 | 7365930 | 92511.57 | 9.48 |  | 26.48 |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.97 | 92511.57 | 2666977 |  | 92511.57 | 4 | 5.58 | 39.5 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.05 | 68.579 | 2800 | 1956.9 | 0.14 |  |  |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.05 | 1956.9 | 75761 |  | 1956.9 | 4 | 5.82 | 45.5 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.07 | 66792.984 | 2494126 | 78271.125 | 3.413 |  | 56.37 |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.4 | 2487.767 | 106960 | 78271.125 | 0.127 |  |  |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.4 | 78271.125 | 3317273 |  | 78271.125 | 4 | 8.76 | 118.9
20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.85 | 3.96 | 193 | 2804.06 | 0.006 |  |  |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.78 | 58.809 | 1433 | 2804.06 | 0.084 |  |  |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.78 | 2804.06 | 95771 |  | 2804.06 | 4 | 3.47 | -13.3 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.81 | 177725.781 | 5986646 | 85295.422 | 8.335 |  | 23.23 |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.97 | 85295.422 | 2390381 |  | 85295.422 | 4 | 5.15 | 28.6 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.05 | 44.038 | 1395 | 1911.096 | 0.092 |  |  |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.05 | 1911.096 | 76727 |  | 1911.096 | 4 | 5.68 | 42.1 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.07 | 63163.129 | 2425594 | 82232.203 | 3.072 |  | 50.68 |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.4 | 1859.349 | 76758 | 82232.203 | 0.09 |  |  |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.4 | 82232.203 | 3383647 |  | 82232.203 | 4 | 9.2 | 130
20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte |  |  |  | 4173.079 |  |  |  |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.78 | 77.001 | 2803 | 4173.079 | 0.074 |  |  |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.78 | 4173.079 | 145040 |  | 4173.079 | 4 | 5.16 | 29 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.81 | 226021.344 | 7559059 | 110140.234 | 8.208 |  | 22.87 |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.96 | 110140.234 | 3188337 |  | 110140.234 | 4 | 6.64 | 66.1 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.05 | 86.69 | 3498 | 2183.897 | 0.159 |  |  |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.05 | 2183.897 | 87172 |  | 2183.897 | 4 | 6.5 | 62.4 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.06 | 83874.047 | 3258557 | 68976.289 | 4.864 |  | 80.6 |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.4 | 2407.742 | 103954 | 68976.289 | 0.14 |  |  |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.4 | 68976.289 | 2929824 |  | 68976.289 | 4 | 7.72 | 92.9
20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.89 | 7.057 | 258 | 3930.435 | 0.007 |  |  |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.78 | 41.344 | 1442 | 3930.435 | 0.042 |  |  |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.78 | 3930.435 | 136738 |  | 3930.435 | 4 | 4.86 | 21.5 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.81 | 278933.406 | 9405930 | 108602.188 | 10.274 |  | 28.73 |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.97 | 108602.188 | 3147145 |  | 108602.188 | 4 | 6.55 | 63.8 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.05 | 47.479 | 2095 | 2102.246 | 0.09 |  |  |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.05 | 2102.246 | 79293 |  | 2102.246 | 4 | 6.25 | 56.3 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.07 | 72234.156 | 2659945 | 80912.125 | 3.571 |  | 59 |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.4 | 1676.365 | 65562 | 80912.125 | 0.083 |  |  |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.4 | 80912.125 | 3194553 |  | 80912.125 | 4 | 9.05 | 126.3
20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte |  |  |  | 3689.984 |  |  |  |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.77 | 39.448 | 947 | 3689.984 | 0.043 |  |  |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.78 | 3689.984 | 126097 |  | 3689.984 | 4 | 4.56 | 14.1 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.81 | 276028.344 | 9283592 | 90797.094 | 12.16 |  | 34.08 |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.96 | 90797.094 | 2563937 |  | 90797.094 | 4 | 5.48 | 36.9 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.05 | 49.284 | 2089 | 2034.973 | 0.097 |  |  |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.05 | 2034.973 | 79528 |  | 2034.973 | 4 | 6.05 | 51.3 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.06 | 76960.344 | 2921204 | 64799.211 | 4.751 |  | 78.7 |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.4 | 1451.845 | 55029 | 64799.211 | 0.09 |  |  |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.4 | 64799.211 | 2501513 |  | 64799.211 | 4 | 7.25 | 81.2
20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.9 | 4.366 | 159 | 3838.459 | 0.005 |  |  |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.78 | 74.386 | 1977 | 3838.459 | 0.078 |  |  |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.78 | 3838.459 | 133137 |  | 3838.459 | 4 | 4.75 | 18.6 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.81 | 367082 | 12520329 | 111401.727 | 13.18 |  | 36.98 |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.97 | 111401.727 | 3155190 |  | 111401.727 | 4 | 6.72 | 68 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.05 | 49.724 | 2144 | 2392.802 | 0.083 |  |  |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.05 | 2392.802 | 92804 |  | 2392.802 | 4 | 7.12 | 77.9 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.07 | 93146.758 | 3627092 | 95344.883 | 3.908 |  | 64.63 |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.4 | 2270.83 | 89051 | 95344.883 | 0.095 |  |  |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.4 | 95344.883 | 3645451 |  | 95344.883 | 4 | 10.67 | 166.6
20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.85 | 2.995 | 159 | 3972.163 | 0.003 |  |  |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.77 | 43.105 | 1522 | 3972.163 | 0.043 |  |  |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.78 | 3972.163 | 136332 |  | 3972.163 | 4 | 4.91 | 22.8 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.81 | 353533.938 | 11858321 | 110531.492 | 12.794 |  | 35.88 |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.96 | 110531.492 | 3088762 |  | 110531.492 | 4 | 6.67 | 66.7 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.05 | 54.339 | 2333 | 2360.479 | 0.092 |  |  |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.05 | 2360.479 | 94249 |  | 2360.479 | 4 | 7.02 | 75.5 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.07 | 77323.891 | 2759422 | 98936.914 | 3.126 |  | 51.58 |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.4 | 2288.021 | 85144 | 98936.914 | 0.093 |  |  |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.4 | 98936.914 | 3873731 |  | 98936.914 | 4 | 11.07 | 176.7
20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.9 | 0.645 | 36 | 3842.832 | 0.001 |  |  |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.78 | 58.642 | 2012 | 3842.832 | 0.061 |  |  |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.78 | 3842.832 | 133074 |  | 3842.832 | 4 | 4.75 | 18.8 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.81 | 323413.156 | 10911557 | 112650.75 | 11.484 |  | 32.16 |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.97 | 112650.75 | 3159743 |  | 112650.75 | 4 | 6.8 | 69.9 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.05 | 71.881 | 2863 | 2468 | 0.117 |  |  |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.05 | 2468 | 99081 |  | 2468 | 4 | 7.34 | 83.5 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.06 | 88055.625 | 3283423 | 96992.188 | 3.631 |  | 60.01 |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.4 | 2269.482 | 96527 | 96992.188 | 0.094 |  |  |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.4 | 96992.188 | 3737935 |  | 96992.188 | 4 | 10.85 | 171.3
20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.87 | 0.965 | 116 | 3893.189 | 0.001 |  |  |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.77 | 87.68 | 2747 | 3893.189 | 0.09 |  |  |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.78 | 3893.189 | 135899 |  | 3893.189 | 4 | 4.81 | 20.3 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.81 | 345005.031 | 11796750 | 113099.43 | 12.202 |  | 34.2 |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.97 | 113099.43 | 3265188 |  | 113099.43 | 4 | 6.82 | 70.6 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.05 | 64.293 | 2544 | 2278.214 | 0.113 |  |  |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.05 | 2278.214 | 91468 |  | 2278.214 | 4 | 6.78 | 69.4 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.07 | 96175.375 | 3901113 | 108030.414 | 3.561 |  | 58.84 |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.4 | 2530.248 | 100776 | 108030.414 | 0.094 |  |  |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.4 | 108030.414 | 4106710 |  | 108030.414 | 4 | 12.08 | 202.1
20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.9 | 5.351 | 219 | 3559.379 | 0.006 |  |  |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.78 | 85.112 | 3077 | 3559.379 | 0.096 |  |  |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.78 | 3559.379 | 121498 |  | 3559.379 | 4 | 4.4 | 10 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.81 | 336113.375 | 11298503 | 91099.672 | 14.758 |  | 41.45 |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.96 | 91099.672 | 2589675 |  | 91099.672 | 4 | 5.5 | 37.4 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.05 | 73.579 | 3336 | 2180.041 | 0.135 |  |  |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.05 | 2180.041 | 83876 |  | 2180.041 | 4 | 6.48 | 62.1 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.06 | 85629.828 | 3304683 | 100591.141 | 3.405 |  | 56.23 |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.4 | 3093.332 | 115462 | 100591.141 | 0.123 |  |  |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.4 | 100591.141 | 3646289 |  | 100591.141 | 4 | 11.25 | 181.3
20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.84 | 2.901 | 108 | 4010.53 | 0.003 |  |  |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.78 | 56.909 | 2038 | 4010.53 | 0.057 |  |  |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.78 | 4010.53 | 139009 |  | 4010.53 | 4 | 4.96 | 24 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.81 | 369345.031 | 12524720 | 95262.484 | 15.509 |  | 43.58 |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.97 | 95262.484 | 2696060 |  | 95262.484 | 4 | 5.75 | 43.7 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.05 | 48.467 | 1868 | 2026.516 | 0.096 |  |  |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.05 | 2026.516 | 78913 |  | 2026.516 | 4 | 6.03 | 50.7 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.07 | 81606.148 | 3017306 | 69586.211 | 4.691 |  | 77.71 |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.4 | 1622.044 | 59654 | 69586.211 | 0.093 |  |  |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.4 | 69586.211 | 2536447 |  | 69586.211 | 4 | 7.78 | 94.6
20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.87 | 6.052 | 191 | 4683.644 | 0.005 |  |  |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.78 | 6827.96 | 236140 | 4683.644 | 5.831 |  | 84.57 |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.78 | 4683.644 | 161988 |  | 4683.644 | 4 | 5.79 | 44.8 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.81 | 278402.094 | 9377985 | 95167.969 | 11.702 |  | 32.78 |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.96 | 95167.969 | 2738875 |  | 95167.969 | 4 | 5.74 | 43.5 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.05 | 4290.784 | 170039 | 1951.082 | 8.797 |  | 51.77 |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.05 | 1951.082 | 75634 |  | 1951.082 | 4 | 5.8 | 45.1 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.06 | 87272.586 | 3324550 | 104467.281 | 3.342 |  | 55.17 |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.4 | 121838.508 | 4179704 | 104467.281 | 4.665 |  | 15.65 |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.4 | 104467.281 | 3438160 |  | 104467.281 | 4 | 11.69 | 192.2
20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte |  |  |  | 3473.603 |  |  |  |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.78 | 217.342 | 7810 | 3473.603 | 0.25 |  | 2.1 |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.78 | 3473.603 | 119313 |  | 3473.603 | 4 | 4.29 | 7.4 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.81 | 253162.359 | 8413451 | 101755.492 | 9.952 |  | 27.82 |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.97 | 101755.492 | 2993356 |  | 101755.492 | 4 | 6.14 | 53.5 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.05 | 187.424 | 7368 | 1931.021 | 0.388 |  | 1.32 |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.05 | 1931.021 | 76802 |  | 1931.021 | 4 | 5.74 | 43.6 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.06 | 83227.961 | 3148725 | 48669.789 | 6.84 |  | 113.6 |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.4 | 4144.071 | 183465 | 48669.789 | 0.341 |  | 0.64 |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.4 | 48669.789 | 2086193 |  | 48669.789 | 4 | 5.44 | 36.1
20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte |  |  |  | 3760.475 |  |  |  |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.78 | 2704.803 | 93865 | 3760.475 | 2.877 |  | 40.91 |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.78 | 3760.475 | 129538 |  | 3760.475 | 4 | 4.65 | 16.2 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.81 | 245721.672 | 8129133 | 87558.133 | 11.226 |  | 31.43 |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.97 | 87558.133 | 2520853 |  | 87558.133 | 4 | 5.28 | 32.1 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.05 | 1550.709 | 60551 | 1544.025 | 4.017 |  | 23.09 |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.05 | 1544.025 | 60954 |  | 1544.025 | 4 | 4.59 | 14.8 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.06 | 73634.305 | 2678980 | 39606.652 | 7.437 |  | 123.56 |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.4 | 34674.074 | 1378169 | 39606.652 | 3.502 |  | 11.61 |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.4 | 39606.652 | 1617206 |  | 39606.652 | 4 | 4.43 | 10.8
20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.85 | 3.359 | 200 | 3804.575 | 0.004 |  |  |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.78 | 79.311 | 2437 | 3804.575 | 0.083 |  |  |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.77 | 3804.575 | 130640 |  | 3804.575 | 4 | 4.7 | 17.6 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.81 | 216459.453 | 6929789 | 118122.609 | 7.33 |  | 20.38 |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.97 | 118122.609 | 3489290 |  | 118122.609 | 4 | 7.13 | 78.1 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.05 | 63.795 | 2704 | 2285.38 | 0.112 |  |  |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.05 | 2285.38 | 88611 |  | 2285.38 | 4 | 6.8 | 70 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.06 | 82758.57 | 3205700 | 100054.281 | 3.309 |  | 54.62 |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.4 | 2658.295 | 114635 | 100054.281 | 0.106 |  |  |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.4 | 100054.281 | 4122043 |  | 100054.281 | 4 | 11.19 | 179.8
20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.85 | 0.614 | 56 | 3650.367 | 0.001 |  |  |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.78 | 63.602 | 1984 | 3650.367 | 0.07 |  |  |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.78 | 3650.367 | 127037 |  | 3650.367 | 4 | 4.51 | 12.8 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.81 | 338452.781 | 11647518 | 102091.367 | 13.261 |  | 37.2 |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.97 | 102091.367 | 2883097 |  | 102091.367 | 4 | 6.16 | 54 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.05 | 58.385 | 2570 | 2275.792 | 0.103 |  |  |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.05 | 2275.792 | 87791 |  | 2275.792 | 4 | 6.77 | 69.2 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.07 | 84048.07 | 3297586 | 91328.695 | 3.681 |  | 60.84 |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.4 | 2295.008 | 85333 | 91328.695 | 0.101 |  |  |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.4 | 91328.695 | 3601070 |  | 91328.695 | 4 | 10.22 | 155.4
20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.82 | 3.82 | 148 | 4028.423 | 0.004 |  |  |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.78 | 57.151 | 1930 | 4028.423 | 0.057 |  |  |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.78 | 4028.423 | 139652 |  | 4028.423 | 4 | 4.98 | 24.5 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.81 | 282383.156 | 9499879 | 125547.516 | 8.997 |  | 25.11 |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.97 | 125547.516 | 3727889 |  | 125547.516 | 4 | 7.57 | 89.3 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.05 | 60.019 | 2065 | 2367.659 | 0.101 |  |  |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.05 | 2367.659 | 92560 |  | 2367.659 | 4 | 7.04 | 76.1 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.07 | 90018.523 | 3535425 | 102183.391 | 3.524 |  | 58.22 |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.4 | 2382.871 | 91499 | 102183.391 | 0.093 |  |  |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.4 | 102183.391 | 3971211 |  | 102183.391 | 4 | 11.43 | 185.8
20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.88 | 8.494 | 187 | 3605.647 | 0.009 |  |  |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.78 | 34.578 | 1251 | 3605.647 | 0.038 |  |  |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.78 | 3605.647 | 124797 |  | 3605.647 | 4 | 4.46 | 11.5 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.81 | 194233.328 | 6480172 | 96524.773 | 8.049 |  | 22.42 |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.97 | 96524.773 | 2796460 |  | 96524.773 | 4 | 5.82 | 45.6 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.05 | 43.619 | 1630 | 2073.876 | 0.084 |  |  |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.05 | 2073.876 | 84779 |  | 2073.876 | 4 | 6.17 | 54.2 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.07 | 68525.758 | 2553414 | 95108.883 | 2.882 |  | 47.5 |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.4 | 2309.284 | 84392 | 95108.883 | 0.097 |  |  |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.4 | 95108.883 | 3551584 |  | 95108.883 | 4 | 10.64 | 166
20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte |  |  |  | 3280.838 |  |  |  |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.78 | 55.165 | 1589 | 3280.838 | 0.067 |  |  |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.78 | 3280.838 | 113487 |  | 3280.838 | 4 | 4.06 | 1.4 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.81 | 237021.109 | 8063289 | 88051.555 | 10.767 |  | 30.13 |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.97 | 88051.555 | 2541873 |  | 88051.555 | 4 | 5.31 | 32.8 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.05 | 49.147 | 2150 | 1820.366 | 0.108 |  |  |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.05 | 1820.366 | 71864 |  | 1820.366 | 4 | 5.41 | 35.4 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.07 | 58218.891 | 2193878 | 53921.77 | 4.319 |  | 71.49 |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.4 | 1318.323 | 47264 | 53921.77 | 0.098 |  |  |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.4 | 53921.77 | 2013740 |  | 53921.77 | 4 | 6.03 | 50.8
20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte |  |  |  | 3179.154 |  |  |  |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.78 | 42.574 | 1442 | 3179.154 | 0.054 |  |  |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.78 | 3179.154 | 110380 |  | 3179.154 | 4 | 3.93 | -1.7 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.81 | 243039.016 | 8258455 | 96476.289 | 10.077 |  | 28.17 |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.97 | 96476.289 | 2806186 |  | 96476.289 | 4 | 5.82 | 45.5 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.05 | 49.617 | 1955 | 1823.648 | 0.109 |  |  |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.05 | 1823.648 | 69186 |  | 1823.648 | 4 | 5.42 | 35.6 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.07 | 65500.477 | 2518845 | 61000.227 | 4.295 |  | 71.1 |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.4 | 1408.505 | 61221 | 61000.227 | 0.092 |  |  |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.4 | 61000.227 | 2395311 |  | 61000.227 | 4 | 6.82 | 70.6
20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte |  |  |  | 4269.086 |  |  |  |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.78 | 110.167 | 3794 | 4269.086 | 0.103 |  |  |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.78 | 4269.086 | 147221 |  | 4269.086 | 4 | 5.28 | 32 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.81 | 352025.063 | 11860594 | 120083.727 | 11.726 |  | 32.85 |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.97 | 120083.727 | 3464551 |  | 120083.727 | 4 | 7.24 | 81.1 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.05 | 189.058 | 7842 | 2702.162 | 0.28 |  | 0.67 |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.05 | 2702.162 | 105538 |  | 2702.162 | 4 | 8.04 | 100.9 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.06 | 98320.117 | 3693365 | 180164.406 | 2.183 |  | 35.82 |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.4 | 6702.993 | 251062 | 180164.406 | 0.149 |  |  |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.4 | 180164.406 | 6972059 |  | 180164.406 | 4 | 20.15 | 403.9
20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.9 | 0.276 | 37 | 4165.854 | 0 |  |  |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.78 | 102.365 | 3701 | 4165.854 | 0.098 |  |  |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.78 | 4165.854 | 143861 |  | 4165.854 | 4 | 5.15 | 28.8 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.81 | 261868.188 | 8873026 | 120913.375 | 8.663 |  | 24.16 |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.97 | 120913.375 | 3529813 |  | 120913.375 | 4 | 7.29 | 82.4 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.05 | 105.691 | 4814 | 2421.035 | 0.175 |  | 0.04 |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.05 | 2421.035 | 92806 |  | 2421.035 | 4 | 7.2 | 80 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.07 | 84453.852 | 3030661 | 128731.063 | 2.624 |  | 43.19 |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.4 | 3526.294 | 137227 | 128731.063 | 0.11 |  |  |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.4 | 128731.063 | 5090614 |  | 128731.063 | 4 | 14.4 | 260
20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.87 | 3.55 | 130 | 4105.743 | 0.003 |  |  |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.78 | 144.584 | 5057 | 4105.743 | 0.141 |  | 0.48 |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.78 | 4105.743 | 139697 |  | 4105.743 | 4 | 5.08 | 26.9 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.81 | 358415.75 | 12138892 | 120947.359 | 11.854 |  | 33.21 |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.97 | 120947.359 | 3381191 |  | 120947.359 | 4 | 7.3 | 82.4 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.05 | 149.801 | 6328 | 2636.327 | 0.227 |  | 0.36 |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.05 | 2636.327 | 101850 |  | 2636.327 | 4 | 7.84 | 96 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.06 | 98009.695 | 3569567 | 119917.117 | 3.269 |  | 53.96 |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.4 | 3763.688 | 148501 | 119917.117 | 0.126 |  |  |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.4 | 119917.117 | 4419736 |  | 119917.117 | 4 | 13.41 | 235.4
20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.91 | 5.588 | 229 | 4091.521 | 0.005 |  |  |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.78 | 301.788 | 9552 | 4091.521 | 0.295 |  | 2.76 |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.78 | 4091.521 | 141866 |  | 4091.521 | 4 | 5.06 | 26.5 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.81 | 284973.594 | 9606363 | 98304.648 | 11.596 |  | 32.48 |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.97 | 98304.648 | 2728535 |  | 98304.648 | 4 | 5.93 | 48.3 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.05 | 311.631 | 12419 | 2181.149 | 0.571 |  | 2.42 |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.05 | 2181.149 | 82489 |  | 2181.149 | 4 | 6.49 | 62.2 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.06 | 66279.32 | 2477570 | 88866.891 | 2.983 |  | 49.19 |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.4 | 9193.139 | 358467 | 88866.891 | 0.414 |  | 0.89 |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.4 | 88866.891 | 3528454 |  | 88866.891 | 4 | 9.94 | 148.5
20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.84 | 5.883 | 249 | 3699.73 | 0.006 |  |  |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.78 | 299.284 | 9875 | 3699.73 | 0.324 |  | 3.18 |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.78 | 3699.73 | 127092 |  | 3699.73 | 4 | 4.57 | 14.4 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.81 | 236682.141 | 7911366 | 108873.719 | 8.696 |  | 24.25 |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.97 | 108873.719 | 3186634 |  | 108873.719 | 4 | 6.57 | 64.2 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.05 | 312.361 | 11999 | 2069.581 | 0.604 |  | 2.61 |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.05 | 2069.581 | 81333 |  | 2069.581 | 4 | 6.16 | 53.9 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.06 | 66022.867 | 2635104 | 88157.313 | 2.996 |  | 49.4 |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.4 | 9416.714 | 361946 | 88157.313 | 0.427 |  | 0.94 |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.4 | 88157.313 | 3428768 |  | 88157.313 | 4 | 9.86 | 146.5
20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.88 | 3.841 | 109 | 3781.379 | 0.004 |  |  |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.78 | 306.603 | 9765 | 3781.379 | 0.324 |  | 3.19 |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.78 | 3781.379 | 130088 |  | 3781.379 | 4 | 4.68 | 16.9 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.81 | 301589.063 | 10239744 | 104880.836 | 11.502 |  | 32.21 |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.97 | 104880.836 | 3105835 |  | 104880.836 | 4 | 6.33 | 58.2 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.05 | 359.079 | 14109 | 2232.48 | 0.643 |  | 2.85 |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.05 | 2232.48 | 87859 |  | 2232.48 | 4 | 6.64 | 66 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.06 | 59553.82 | 2256071 | 98938.063 | 2.408 |  | 39.58 |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.4 | 10473.873 | 423184 | 98938.063 | 0.423 |  | 0.93 |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.4 | 98938.063 | 4006047 |  | 98938.063 | 4 | 11.07 | 176.7
20231106_HepCl_ReRun_001 | Blank | Blank | 3.86 | 7.841 | 320 | [5 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | 3.82 | 149.304 | 2967 | [5 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_001 | Blank | Blank | 3.86 | 1311.774 | 28378 | [5 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_001 | Blank | Blank |  |  |  | 2.685 |  |  |  |  | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.07 | 2.685 | 157 |  | 2.685 | 4 | 0.01 | -99.8 | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.04 | 5.602 | 132 | [5 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.42 | 18.895 | 1390 | [5 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_002 | Blank | Blank | 3.88 | 1.966 | 103 | 2.377 | 3.308 |  | 11.37 |  | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.81 | 70.268 | 1184 | 2.377 | 118.247 |  | 1745.78 |  | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.79 | 2.377 | 87 |  | 2.377 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.86 | 881.125 | 16293 | [5 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.06 | 1.228 | 71 | 3.534 | 1.39 |  | 7.33 |  | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.04 | 3.534 | 118 |  | 3.534 | 4 | 0.01 | -99.7 | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.05 | 4.978 | 128 | [5 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.41 | 62.547 | 1907 | [5 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_003 | Blank | Blank |  |  |  | 0.237 |  |  |  |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.83 | 51.466 | 878 | 0.237 | 868.624 |  | 12834.47 |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.81 | 0.237 | 33 |  | 0.237 | 4 | 0 | -100 | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.85 | 721.975 | 14600 | 5.359 | 538.888 |  | 1528.2 |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.95 | 5.359 | 177 |  | 5.359 | 4 | 0 | -100 | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.04 | 4.574 | 189 | [5 empty] | 20231106_HepCl_ReRun_003 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.07 | 0.235 | 18 | 6.982 | 0.135 |  | 1.62 |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.39 | 1.965 | 356 | 6.982 | 1.126 |  | 3.36 |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.39 | 6.982 | 235 |  | 6.982 | 4 | 0 | -100
20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 8.596 |  |  |  |  | 20231106_HepCl_ReRun_006 | Blank | Blank | 3.81 | 46.309 | 811 | 8.596 | 21.549 |  | 316.84 |  | 20231106_HepCl_ReRun_006 | Blank | Blank | 3.82 | 8.596 | 316 |  | 8.596 | 4 | 0.01 | -99.7 | 20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 13.111 |  |  |  |  | 20231106_HepCl_ReRun_006 | Blank | Blank | 3.96 | 13.111 | 178 |  | 13.111 | 4 | 0 | -100 | 20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 0.473 |  |  |  |  | 20231106_HepCl_ReRun_006 | Blank | Blank | 4.04 | 0.473 | 28 |  | 0.473 | 4 | 0 | -100 | 20231106_HepCl_ReRun_006 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_006 | Blank | Blank | 4.39 | 12.852 | 1235 | [5 empty] | 20231106_HepCl_ReRun_006 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_007 | Blank | Blank | 3.88 | 6.992 | 172 | [5 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | 3.81 | 31.905 | 854 | [5 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_007 | Blank | Blank | 3.85 | 443.105 | 9101 | [5 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.05 | 0.796 | 47 | 1.479 | 2.153 |  | 11.91 |  | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.09 | 1.479 | 80 |  | 1.479 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.03 | 2.4 | 141 | 7.306 | 1.314 |  | 21.31 |  | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.35 | 36.801 | 1623 | 7.306 | 20.148 |  | 69.4 |  | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.4 | 7.306 | 320 |  | 7.306 | 4 | 0 | -100
20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.04 |  |  |  |  | 20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.04 |  |  |  |  | 20231106_HepCl_ReRun_014 | Blank | Blank | 3.74 | 3.04 | 110 |  | 3.04 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.806 |  |  |  |  | 20231106_HepCl_ReRun_014 | Blank | Blank | 3.96 | 3.806 | 133 |  | 3.806 | 4 | 0 | -100 | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.06 | 0.207 | 12 | 0.151 | 5.483 |  | 31.89 |  | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.01 | 0.151 | 15 |  | 0.151 | 4 | 0 | -100 | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.05 | 2.265 | 131 | 0.639 | 14.178 |  | 236.14 |  | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.42 | 1.451 | 372 | 0.639 | 9.083 |  | 30.99 |  | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.37 | 0.639 | 71 |  | 0.639 | 4 | 0 | -100
20231106_HepCl_ReRun_021 | Blank | Blank | 3.9 | 10.663 | 141 | 2.584 | 16.506 |  | 59.62 |  | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.82 | 11.18 | 423 | 2.584 | 17.307 |  | 254.14 |  | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.82 | 2.584 | 93 |  | 2.584 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.85 | 265.891 | 4687 | 7.869 | 135.159 |  | 382.98 |  | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.99 | 7.869 | 293 |  | 7.869 | 4 | 0 | -100 | 20231106_HepCl_ReRun_021 | Blank | Blank |  |  |  | 3.142 |  |  |  |  | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.01 | 3.142 | 118 |  | 3.142 | 4 | 0.01 | -99.8 | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.09 | 6.331 | 205 | [5 empty] | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.43 | 18.559 | 1477 | [5 empty] | 20231106_HepCl_ReRun_021 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_032 | Blank | Blank | 3.87 | 5.334 | 155 | 0.567 | 37.63 |  | 136.84 |  | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.82 | 4.34 | 147 | 0.567 | 30.617 |  | 450.84 |  | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.8 | 0.567 | 13 |  | 0.567 | 4 | 0 | -100 | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.85 | 231.762 | 4646 | [5 empty] | 20231106_HepCl_ReRun_032 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 0.372 |  |  |  |  | 20231106_HepCl_ReRun_032 | Blank | Blank | 4.01 | 0.372 | 22 |  | 0.372 | 4 | 0 | -100 | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 4.114 |  |  |  |  | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 4.114 |  |  |  |  | 20231106_HepCl_ReRun_032 | Blank | Blank | 4.35 | 4.114 | 195 |  | 4.114 | 4 | 0 | -100
20231106_HepCl_ReRun_036 | Blank | Blank | 3.92 | 10.411 | 310 | [5 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank | 3.82 | 14.094 | 247 | [5 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_036 | Blank | Blank | 3.85 | 224.475 | 4100 | [5 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_036 | Blank | Blank | 4.06 | 1.437 | 84 | 0.326 | 17.632 |  | 104.77 |  | 20231106_HepCl_ReRun_036 | Blank | Blank | 4.07 | 0.326 | 19 |  | 0.326 | 4 | 0 | -100 | 20231106_HepCl_ReRun_036 | Blank | Blank |  |  |  | 7.725 |  |  |  |  | 20231106_HepCl_ReRun_036 | Blank | Blank |  |  |  | 7.725 |  |  |  |  | 20231106_HepCl_ReRun_036 | Blank | Blank | 4.41 | 7.725 | 173 |  | 7.725 | 4 | 0 | -100
20231106_HepCl_ReRun_040 | Blank | Blank |  |  |  | 2.239 |  |  |  |  | 20231106_HepCl_ReRun_040 | Blank | Blank |  |  |  | 2.239 |  |  |  |  | 20231106_HepCl_ReRun_040 | Blank | Blank | 3.75 | 2.239 | 81 |  | 2.239 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_040 | Blank | Blank | 3.85 | 223.767 | 3692 | [5 empty] | 20231106_HepCl_ReRun_040 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.05 | 0.854 | 50 | [5 empty] | 20231106_HepCl_ReRun_040 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.07 | 5.768 | 153 | 5.681 | 4.061 |  | 67.19 |  | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.44 | 65.969 | 2411 | 5.681 | 46.449 |  | 160.69 |  | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.4 | 5.681 | 201 |  | 5.681 | 4 | 0 | -100
20231106_HepCl_ReRun_053 | Blank | Blank | 3.9 | 1.849 | 69 | 1.263 | 5.856 |  | 20.68 |  | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 1.263 |  |  |  |  | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.79 | 1.263 | 40 |  | 1.263 | 4 | 0 | -100 | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.85 | 234.037 | 4499 | 12.536 | 74.677 |  | 211.42 |  | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.97 | 12.536 | 315 |  | 12.536 | 4 | 0 | -100 | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 0.376 |  |  |  |  | 20231106_HepCl_ReRun_053 | Blank | Blank | 4.09 | 0.376 | 22 |  | 0.376 | 4 | 0 | -100 | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 3.035 |  |  |  |  | 20231106_HepCl_ReRun_053 | Blank | Blank | 4.43 | 2.767 | 722 | 3.035 | 3.647 |  | 12.11 |  | 20231106_HepCl_ReRun_053 | Blank | Blank | 4.39 | 3.035 | 208 |  | 3.035 | 4 | 0 | -100
20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 1.55 |  |  |  |  | 20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 1.55 |  |  |  |  | 20231106_HepCl_ReRun_060 | Blank | Blank | 3.76 | 1.55 | 56 |  | 1.55 | 4 | 0 | -100 | 20231106_HepCl_ReRun_060 | Blank | Blank | 3.85 | 243.211 | 4301 | [5 empty] | 20231106_HepCl_ReRun_060 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 0.214 |  |  |  |  | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.05 | 0.214 | 13 |  | 0.214 | 4 | 0 | -100 | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.06 | 11.194 | 371 | 2.979 | 15.031 |  | 250.38 |  | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.43 | 61.475 | 1351 | 2.979 | 82.544 |  | 285.99 |  | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.41 | 2.979 | 184 |  | 2.979 | 4 | 0 | -100
20231106_HepCl_ReRun_071 | Blank | Blank |  |  |  | 5.703 |  |  |  |  | 20231106_HepCl_ReRun_071 | Blank | Blank |  |  |  | 5.703 |  |  |  |  | 20231106_HepCl_ReRun_071 | Blank | Blank | 3.79 | 5.703 | 133 |  | 5.703 | 4 | 0.01 | -99.8 | 20231106_HepCl_ReRun_071 | Blank | Blank | 3.86 | 178.541 | 3006 | 7.057 | 101.199 |  | 286.65 |  | 20231106_HepCl_ReRun_071 | Blank | Blank | 3.97 | 7.057 | 248 |  | 7.057 | 4 | 0 | -100 | 20231106_HepCl_ReRun_071 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_071 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.04 | 0.246 | 35 | 13.529 | 0.073 |  | 0.58 |  | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.45 | 47.228 | 2391 | 13.529 | 13.963 |  | 47.93 |  | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.37 | 13.529 | 246 |  | 13.529 | 4 | 0 | -100
20231106_HepCl_ReRun_075 | Blank | Blank |  |  |  | 0.322 |  |  |  |  | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.79 | 0.88 | 95 | 0.322 | 10.932 |  | 159.94 |  | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.81 | 0.322 | 12 |  | 0.322 | 4 | 0 | -100 | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.86 | 219.784 | 4284 | 12.113 | 72.578 |  | 205.46 |  | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.97 | 12.113 | 359 |  | 12.113 | 4 | 0 | -100 | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.06 | 0.203 | 12 | [5 empty] | 20231106_HepCl_ReRun_075 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_075 | Blank | Blank |  |  |  | 8.646 |  |  |  |  | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.38 | 69.248 | 2207 | 8.646 | 32.037 |  | 110.66 |  | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.4 | 8.646 | 340 |  | 8.646 | 4 | 0 | -100
20231106_HepCl_ReRun_079 | Blank | Blank | 3.83 | 0.96 | 30 | 4.597 | 0.835 |  | 2.32 |  | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 4.597 |  |  |  |  | 20231106_HepCl_ReRun_079 | Blank | Blank | 3.81 | 4.597 | 132 |  | 4.597 | 4 | 0.01 | -99.9 | 20231106_HepCl_ReRun_079 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_079 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 0.232 |  |  |  |  | 20231106_HepCl_ReRun_079 | Blank | Blank | 4.05 | 0.232 | 36 |  | 0.232 | 4 | 0 | -100 | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 10.156 |  |  |  |  | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 10.156 |  |  |  |  | 20231106_HepCl_ReRun_079 | Blank | Blank | 4.36 | 10.156 | 202 |  | 10.156 | 4 | 0 | -100
20231106_HepCl_ReRun_084 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.83 | 20.354 | 553 | [5 empty] | 20231106_HepCl_ReRun_084 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.86 | 108.338 | 2639 | 2.32 | 186.79 |  | 529.44 |  | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.93 | 2.32 | 130 |  | 2.32 | 4 | 0 | -100 | 20231106_HepCl_ReRun_084 | Blank | Blank |  |  |  | 0.156 |  |  |  |  | 20231106_HepCl_ReRun_084 | Blank | Blank | 4.05 | 0.156 | 9 |  | 0.156 | 4 | 0 | -100 | 20231106_HepCl_ReRun_084 | Blank | Blank | 4.06 | 4.429 | 252 | 4.803 | 3.689 |  | 60.97 |  | 20231106_HepCl_ReRun_084 | Blank | Blank |  |  |  | 4.803 |  |  |  |  | 20231106_HepCl_ReRun_084 | Blank | Blank | 4.4 | 4.803 | 181 |  | 4.803 | 4 | 0 | -100
20231106_HepCl_ReRun_085 | Blank | Blank | 3.86 | 1.076 | 86 | [5 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | 3.83 | 1.805 | 126 | [5 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_085 | Blank | Blank | 3.86 | 243.378 | 5428 | [5 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_085 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.06 | 5.79 | 191 | 0.242 | 95.702 |  | 1597.58 |  | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.36 | 4.824 | 525 | 0.242 | 79.736 |  | 276.24 |  | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.42 | 0.242 | 58 |  | 0.242 | 4 | 0 | -100
20231106_HepCl_ReRun_087 | Blank | Blank | 3.89 | 6.497 | 168 | [5 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | 3.82 | 12.764 | 267 | [5 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_087 | Blank | Blank |  |  |  | 3.535 |  |  |  |  | 20231106_HepCl_ReRun_087 | Blank | Blank | 3.91 | 3.535 | 119 |  | 3.535 | 4 | 0 | -100 | 20231106_HepCl_ReRun_087 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_087 | Blank | Blank | 4.09 | 9.932 | 215 | [5 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | 4.36 | 22.948 | 1505 | [5 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 1.329 |  |  |  |  | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.83 | 17.262 | 362 | 1.329 | 51.955 |  | 766.16 |  | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.81 | 1.329 | 48 |  | 1.329 | 4 | 0 | -100 | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.86 | 135.857 | 3418 | 4.894 | 111.04 |  | 314.56 |  | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.97 | 4.894 | 185 |  | 4.894 | 4 | 0 | -100 | 20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 1.338 |  |  |  |  | 20231106_HepCl_ReRun_094 | Blank | Blank | 4.1 | 1.338 | 61 |  | 1.338 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_094 | Blank | Blank | 4.08 | 2.282 | 84 | 0.119 | 76.706 |  | 1280.34 |  | 20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 0.119 |  |  |  |  | 20231106_HepCl_ReRun_094 | Blank | Blank | 4.4 | 0.119 | 22 |  | 0.119 | 4 | 0 | -100
20231106_HepCl_ReRun_101 | Blank | Blank | 3.83 | 17.593 | 249 | [5 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | 3.82 | 6.098 | 182 | [5 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_101 | Blank | Blank | 3.85 | 166.996 | 3172 | [5 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.06 | 1.23 | 72 | [5 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.04 | 13.821 | 319 | 1.22 | 45.315 |  | 756.11 |  | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.42 | 49.545 | 2242 | 1.22 | 162.443 |  | 563.34 |  | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.39 | 1.22 | 126 |  | 1.22 | 4 | 0 | -100
20231106_HepCl_ReRun_108 | Blank | Blank | 3.85 | 6.219 | 242 | [5 empty] | 20231106_HepCl_ReRun_108 | Blank | Blank | 3.82 | 2.795 | 99 | [5 empty] | 20231106_HepCl_ReRun_108 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_108 | Blank | Blank |  |  |  | 3.622 |  |  |  |  | 20231106_HepCl_ReRun_108 | Blank | Blank | 3.97 | 3.622 | 139 |  | 3.622 | 4 | 0 | -100 | 20231106_HepCl_ReRun_108 | Blank | Blank |  |  |  | 0.289 |  |  |  |  | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.1 | 0.289 | 13 |  | 0.289 | 4 | 0 | -100 | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.07 | 2.447 | 167 | 7.197 | 1.36 |  | 22.08 |  | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.44 | 120.509 | 2214 | 7.197 | 66.977 |  | 231.95 |  | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.41 | 7.197 | 270 |  | 7.197 | 4 | 0 | -100
20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 3.311 |  |  |  |  | 20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 3.311 |  |  |  |  | 20231106_HepCl_ReRun_110 | Blank | Blank | 3.81 | 3.311 | 119 |  | 3.311 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_110 | Blank | Blank | 3.85 | 202.452 | 3781 | 1.239 | 653.598 |  | 1853.58 |  | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.01 | 1.239 | 97 |  | 1.239 | 4 | 0 | -100 | 20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 0.254 |  |  |  |  | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.04 | 0.254 | 45 |  | 0.254 | 4 | 0 | -100 | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.06 | 3.336 | 134 | 1.133 | 11.778 |  | 196.05 |  | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.42 | 11.766 | 1260 | 1.133 | 41.539 |  | 143.65 |  | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.38 | 1.133 | 113 |  | 1.133 | 4 | 0 | -100
20231106_HepCl_ReRun_117 | Blank | Blank | 3.91 | 20.044 | 377 | [5 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_117 | Blank | Blank | 3.86 | 219.288 | 4057 | [5 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_117 | Blank | Blank |  |  |  | 0.204 |  |  |  |  | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.06 | 0.204 | 16 |  | 0.204 | 4 | 0 | -100 | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.02 | 3.587 | 216 | 1.785 | 8.038 |  | 133.6 |  | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.4 | 73.177 | 2401 | 1.785 | 163.982 |  | 568.68 |  | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.35 | 1.785 | 169 |  | 1.785 | 4 | 0 | -100
20231106_HepCl_ReRun_124 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank | 3.82 | 10.597 | 266 | [5 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_124 | Blank | Blank | 3.85 | 189.908 | 3607 | [5 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_124 | Blank | Blank |  |  |  | 0.326 |  |  |  |  | 20231106_HepCl_ReRun_124 | Blank | Blank | 4.09 | 0.326 | 19 |  | 0.326 | 4 | 0 | -100 | 20231106_HepCl_ReRun_124 | Blank | Blank | 4.06 | 3.813 | 215 | 8.867 | 1.72 |  | 28.09 |  | 20231106_HepCl_ReRun_124 | Blank | Blank |  |  |  | 8.867 |  |  |  |  | 20231106_HepCl_ReRun_124 | Blank | Blank | 4.39 | 8.867 | 284 |  | 8.867 | 4 | 0 | -100
20231106_HepCl_ReRun_131 | Blank | Blank |  |  |  | 1.14 |  |  |  |  | 20231106_HepCl_ReRun_131 | Blank | Blank |  |  |  | 1.14 |  |  |  |  | 20231106_HepCl_ReRun_131 | Blank | Blank | 3.8 | 1.14 | 41 |  | 1.14 | 4 | 0 | -100 | 20231106_HepCl_ReRun_131 | Blank | Blank | 3.85 | 214.173 | 4079 | 14.521 | 58.997 |  | 166.94 |  | 20231106_HepCl_ReRun_131 | Blank | Blank | 3.92 | 14.521 | 393 |  | 14.521 | 4 | 0 | -100 | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.05 | 0.165 | 30 | 0.336 | 1.964 |  | 10.78 |  | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.09 | 0.336 | 18 |  | 0.336 | 4 | 0 | -100 | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.06 | 0.554 | 114 | 0.21 | 10.552 |  | 175.59 |  | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.43 | 5.209 | 616 | 0.21 | 99.219 |  | 343.87 |  | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.44 | 0.21 | 68 |  | 0.21 | 4 | 0 | -100
20231106_HepCl_ReRun_133 | Blank | Blank | 3.82 | 1.416 | 98 | 0.164 | 34.537 |  | 125.54 |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.81 | 10.285 | 369 | 0.164 | 250.854 |  | 3705.38 |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.81 | 0.164 | 6 |  | 0.164 | 4 | 0 | -100 | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.85 | 208.384 | 3838 | 7.49 | 111.287 |  | 315.26 |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.97 | 7.49 | 223 |  | 7.49 | 4 | 0 | -100 | 20231106_HepCl_ReRun_133 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_133 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.09 | 2.798 | 175 | 1.004 | 11.147 |  | 185.53 |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.39 | 20.099 | 1631 | 1.004 | 80.076 |  | 277.42 |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.38 | 1.004 | 93 |  | 1.004 | 4 | 0 | -100
20231106_HepCl_ReRun_140 | Blank | Blank | 3.88 | 3.301 | 150 | [5 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | 3.81 | 26.678 | 650 | [5 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_140 | Blank | Blank | 3.85 | 363.167 | 8123 | [5 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_140 | Blank | Blank |  |  |  | 0.47 |  |  |  |  | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.06 | 0.47 | 28 |  | 0.47 | 4 | 0 | -100 | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.06 | 0.123 | 16 | 3.553 | 0.138 |  | 1.68 |  | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.43 | 18.612 | 2037 | 3.553 | 20.954 |  | 72.19 |  | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.4 | 3.553 | 153 |  | 3.553 | 4 | 0 | -100
20231106_HepCl_ReRun_142 | Blank | Blank |  |  |  | 2.523 |  |  |  |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.81 | 20.394 | 384 | 2.523 | 32.333 |  | 476.2 |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.72 | 2.523 | 91 |  | 2.523 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.86 | 295.344 | 5078 | [5 empty] | 20231106_HepCl_ReRun_142 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.02 | 0.165 | 18 | 0.21 | 3.143 |  | 17.85 |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.09 | 0.21 | 11 |  | 0.21 | 4 | 0 | -100 | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.1 | 3.614 | 178 | 4.462 | 3.24 |  | 53.47 |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.41 | 27.864 | 1977 | 4.462 | 24.979 |  | 86.16 |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.39 | 4.462 | 185 |  | 4.462 | 4 | 0 | -100
20231106_HepCl_ReRun_143 | Blank | Blank |  |  |  | 5.162 |  |  |  |  | 20231106_HepCl_ReRun_143 | Blank | Blank |  |  |  | 5.162 |  |  |  |  | 20231106_HepCl_ReRun_143 | Blank | Blank | 3.81 | 5.162 | 216 |  | 5.162 | 4 | 0.01 | -99.8 | 20231106_HepCl_ReRun_143 | Blank | Blank | 3.85 | 306.383 | 5571 | 2.895 | 423.327 |  | 1200.4 |  | 20231106_HepCl_ReRun_143 | Blank | Blank | 3.92 | 2.895 | 130 |  | 2.895 | 4 | 0 | -100 | 20231106_HepCl_ReRun_143 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_143 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_152 | Blank | Blank | 3.89 | 12.933 | 398 | 0.579 | 89.347 |  | 325.92 |  | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.81 | 16.621 | 419 | 0.579 | 114.826 |  | 1695.23 |  | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.79 | 0.579 | 34 |  | 0.579 | 4 | 0 | -100 | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.86 | 275.404 | 5049 | [5 empty] | 20231106_HepCl_ReRun_152 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.05 | 0.254 | 15 | [5 empty] | 20231106_HepCl_ReRun_152 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.05 | 4.853 | 276 | 6.476 | 2.998 |  | 49.43 |  | 20231106_HepCl_ReRun_152 | Blank | Blank |  |  |  | 6.476 |  |  |  |  | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.4 | 6.476 | 159 |  | 6.476 | 4 | 0 | -100
20231106_HepCl_ReRun_159 | Blank | Blank | 3.83 | 4.698 | 88 | [5 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.8 | 12.358 | 293 | [5 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.85 | 188.909 | 3291 | 4.77 | 158.414 |  | 448.95 |  | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.95 | 4.77 | 219 |  | 4.77 | 4 | 0 | -100 | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.08 | 0.303 | 18 | [5 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.02 | 4.052 | 238 | 0.75 | 21.611 |  | 360.26 |  | 20231106_HepCl_ReRun_159 | Blank | Blank |  |  |  | 0.75 |  |  |  |  | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.4 | 0.75 | 88 |  | 0.75 | 4 | 0 | -100
20231106_HepCl_ReRun_166 | Blank | Blank |  |  |  | 0.413 |  |  |  |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 3.78 | 0.364 | 40 | 0.413 | 3.525 |  | 50.49 |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 3.82 | 0.413 | 15 |  | 0.413 | 4 | 0 | -100 | 20231106_HepCl_ReRun_166 | Blank | Blank |  |  |  | 6.376 |  |  |  |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 3.96 | 6.376 | 199 |  | 6.376 | 4 | 0 | -100 | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.05 | 0.883 | 52 | 0.343 | 10.297 |  | 60.77 |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.06 | 0.343 | 20 |  | 0.343 | 4 | 0 | -100 | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.04 | 8.785 | 356 | 3.511 | 10.009 |  | 166.51 |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.41 | 1.558 | 302 | 3.511 | 1.775 |  | 5.62 |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.41 | 3.511 | 188 |  | 3.511 | 4 | 0 | -100
20231106_HepCl_ReRun_168 | Blank | Blank |  |  |  | 2.772 |  |  |  |  | 20231106_HepCl_ReRun_168 | Blank | Blank | 3.82 | 10.882 | 381 | 2.772 | 15.703 |  | 230.44 |  | 20231106_HepCl_ReRun_168 | Blank | Blank | 3.81 | 2.772 | 100 |  | 2.772 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_168 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_168 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_168 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_175 | Blank | Blank | 3.89 | 4.209 | 179 | [5 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | 3.82 | 11.069 | 300 | [5 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_175 | Blank | Blank | 3.85 | 215.745 | 4778 | [5 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_175 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.02 | 0.679 | 63 | 6.395 | 0.425 |  | 6.46 |  | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.43 | 88.735 | 1658 | 6.395 | 55.503 |  | 192.12 |  | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.4 | 6.395 | 288 |  | 6.395 | 4 | 0 | -100
20231106_HepCl_ReRun_182 | Blank | Blank | 3.84 | 7.722 | 217 | 1.668 | 18.518 |  | 66.97 |  | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.81 | 0.8 | 62 | 1.668 | 1.918 |  | 26.75 |  | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.81 | 1.668 | 38 |  | 1.668 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.86 | 188.246 | 4310 | 14.625 | 51.486 |  | 145.63 |  | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.96 | 14.625 | 344 |  | 14.625 | 4 | 0 | -100 | 20231106_HepCl_ReRun_182 | Blank | Blank | 4.02 | 0.146 | 9 | 0.546 | 1.07 |  | 5.41 |  | 20231106_HepCl_ReRun_182 | Blank | Blank | 4.09 | 0.546 | 33 |  | 0.546 | 4 | 0 | -100 | 20231106_HepCl_ReRun_182 | Blank | Blank |  |  |  | 0.43 |  |  |  |  | 20231106_HepCl_ReRun_182 | Blank | Blank |  |  |  | 0.43 |  |  |  |  | 20231106_HepCl_ReRun_182 | Blank | Blank | 4.4 | 0.43 | 81 |  | 0.43 | 4 | 0 | -100
20231106_HepCl_ReRun_189 | Blank | Blank | 3.87 | 10.259 | 349 | [5 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | 3.81 | 19.428 | 557 | [5 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_189 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.09 | 0.168 | 9 | 0.191 | 3.518 |  | 20.1 |  | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.06 | 0.191 | 11 |  | 0.191 | 4 | 0 | -100 | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.1 | 4.923 | 175 | 2.904 | 6.781 |  | 112.61 |  | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.43 | 14.778 | 1378 | 2.904 | 20.355 |  | 70.11 |  | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.42 | 2.904 | 87 |  | 2.904 | 4 | 0 | -100
20231106_HepCl_ReRun_191 | Blank | Blank | 3.83 | 5.461 | 215 | [5 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_191 | Blank | Blank | 3.85 | 107.023 | 2703 | 6.924 | 61.827 |  | 174.97 |  | 20231106_HepCl_ReRun_191 | Blank | Blank | 3.92 | 6.924 | 263 |  | 6.924 | 4 | 0 | -100 | 20231106_HepCl_ReRun_191 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.04 | 7.196 | 210 | 2.254 | 12.77 |  | 212.63 |  | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.41 | 3.016 | 787 | 2.254 | 5.352 |  | 18.03 |  | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.39 | 2.254 | 110 |  | 2.254 | 4 | 0 | -100
20231106_HepCl_ReRun_198 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.81 | 18.713 | 586 | [5 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.86 | 207.931 | 4316 | 3.35 | 248.276 |  | 703.85 |  | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.92 | 3.35 | 148 |  | 3.35 | 4 | 0 | -100 | 20231106_HepCl_ReRun_198 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_198 | Blank | Blank |  |  |  | 0.96 |  |  |  |  | 20231106_HepCl_ReRun_198 | Blank | Blank | 4.35 | 76.1 | 1422 | 0.96 | 317.083 |  | 1100.13 |  | 20231106_HepCl_ReRun_198 | Blank | Blank | 4.37 | 0.96 | 105 |  | 0.96 | 4 | 0 | -100
20231106_HepCl_ReRun_200 | Blank | Blank |  |  |  | 0.525 |  |  |  |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.82 | 9.87 | 225 | 0.525 | 75.2 |  | 1109.66 |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.81 | 0.525 | 61 |  | 0.525 | 4 | 0 | -100 | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.85 | 181.452 | 3449 | 2.811 | 258.203 |  | 732.01 |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 4 | 2.811 | 117 |  | 2.811 | 4 | 0 | -100 | 20231106_HepCl_ReRun_200 | Blank | Blank |  |  |  | 0.632 |  |  |  |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.04 | 0.632 | 37 |  | 0.632 | 4 | 0 | -100 | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.04 | 2.128 | 95 | 1.102 | 7.724 |  | 128.36 |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.43 | 185.614 | 2410 | 1.102 | 673.735 |  | 2338.17 |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.4 | 1.102 | 122 |  | 1.102 | 4 | 0 | -100
20231106_HepCl_ReRun_207 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank | 3.81 | 22.558 | 431 | [5 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_207 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.05 | 0.506 | 30 | 1.208 | 1.675 |  | 9.04 |  | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.08 | 1.208 | 32 |  | 1.208 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.1 | 3.972 | 181 | 1.83 | 8.682 |  | 144.36 |  | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.4 | 24.258 | 1634 | 1.83 | 53.023 |  | 183.51 |  | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.4 | 1.83 | 193 |  | 1.83 | 4 | 0 | -100
20231106_HepCl_ReRun_214 | Blank | Blank | 3.87 | 9.578 | 198 | [5 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | 3.82 | 21.328 | 482 | [5 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_214 | Blank | Blank | 3.85 | 12.237 | 909 | [5 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.04 | 0.323 | 19 | [5 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.07 | 12.603 | 469 | 1.23 | 40.985 |  | 683.81 |  | 20231106_HepCl_ReRun_214 | Blank | Blank |  |  |  | 1.23 |  |  |  |  | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.43 | 1.23 | 122 |  | 1.23 | 4 | 0 | -100
20231106_HepCl_ReRun_221 | Blank | Blank | 3.87 | 8.066 | 252 | [5 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_221 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_221 | Blank | Blank |  |  |  | 0.465 |  |  |  |  | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.09 | 0.465 | 26 |  | 0.465 | 4 | 0 | -100 | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.02 | 10.347 | 233 | 0.419 | 98.778 |  | 1648.94 |  | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.36 | 269.191 | 2916 | 0.419 | 2569.842 |  | 8920.05 |  | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.38 | 0.419 | 76 |  | 0.419 | 4 | 0 | -100
20231106_HepCl_ReRun_223 | Blank | Blank | 3.87 | 3.376 | 132 | [5 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | 3.82 | 17.31 | 342 | [5 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_223 | Blank | Blank | 3.85 | 209.716 | 4694 | [5 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_223 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_223 | Blank | Blank | 4.05 | 3.697 | 187 | 5.807 | 2.547 |  | 41.9 |  | 20231106_HepCl_ReRun_223 | Blank | Blank |  |  |  | 5.807 |  |  |  |  | 20231106_HepCl_ReRun_223 | Blank | Blank | 4.39 | 5.807 | 229 |  | 5.807 | 4 | 0 | -100
20231106_HepCl_ReRun_230 | Blank | Blank | 3.83 | 4.137 | 166 | 4.215 | 3.926 |  | 13.62 |  | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.81 | 5.699 | 139 | 4.215 | 5.408 |  | 78.32 |  | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.82 | 4.215 | 142 |  | 4.215 | 4 | 0.01 | -99.9 | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.85 | 165.117 | 3099 | 9.549 | 69.166 |  | 195.78 |  | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.96 | 9.549 | 291 |  | 9.549 | 4 | 0 | -100 | 20231106_HepCl_ReRun_230 | Blank | Blank |  |  |  | 0.182 |  |  |  |  | 20231106_HepCl_ReRun_230 | Blank | Blank | 4.04 | 0.182 | 11 |  | 0.182 | 4 | 0 | -100 | 20231106_HepCl_ReRun_230 | Blank | Blank | 4.08 | 13.86 | 401 | 2.226 | 24.906 |  | 415.29 |  | 20231106_HepCl_ReRun_230 | Blank | Blank |  |  |  | 2.226 |  |  |  |  | 20231106_HepCl_ReRun_230 | Blank | Blank | 4.39 | 2.226 | 148 |  | 2.226 | 4 | 0 | -100
20231106_HepCl_ReRun_237 | Blank | Blank | 3.85 | 11.589 | 261 | 1.939 | 23.907 |  | 86.67 |  | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 1.939 |  |  |  |  | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.74 | 1.939 | 70 |  | 1.939 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.86 | 148.115 | 2568 | 0.334 | 1773.832 |  | 5031.24 |  | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.98 | 0.334 | 38 |  | 0.334 | 4 | 0 | -100 | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 0.132 |  |  |  |  | 20231106_HepCl_ReRun_237 | Blank | Blank | 4.04 | 0.132 | 21 |  | 0.132 | 4 | 0 | -100 | 20231106_HepCl_ReRun_237 | Blank | Blank | 4.06 | 2.31 | 145 | 8.198 | 1.127 |  | 18.19 |  | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 8.198 |  |  |  |  | 20231106_HepCl_ReRun_237 | Blank | Blank | 4.37 | 8.198 | 144 |  | 8.198 | 4 | 0 | -100
20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 0.26 |  |  |  |  | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.82 | 13.663 | 294 | 0.26 | 210.2 |  | 3104.62 |  | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.79 | 0.26 | 9 |  | 0.26 | 4 | 0 | -100 | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.85 | 184.859 | 3229 | [5 empty] | 20231106_HepCl_ReRun_244 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 0.507 |  |  |  |  | 20231106_HepCl_ReRun_244 | Blank | Blank | 4.09 | 0.507 | 30 |  | 0.507 | 4 | 0 | -100 | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 7.8 |  |  |  |  | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 7.8 |  |  |  |  | 20231106_HepCl_ReRun_244 | Blank | Blank | 4.44 | 7.8 | 151 |  | 7.8 | 4 | 0 | -100
20231106_HepCl_ReRun_246 | Blank | Blank | 3.87 | 0.829 | 74 | [5 empty] | 20231106_HepCl_ReRun_246 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_246 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_246 | Blank | Blank | 3.86 | 213.496 | 3757 | 1.2 | 711.653 |  | 2018.26 |  | 20231106_HepCl_ReRun_246 | Blank | Blank | 3.96 | 1.2 | 64 |  | 1.2 | 4 | 0 | -100 | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.03 | 0.616 | 36 | 0.149 | 16.537 |  | 98.2 |  | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.04 | 0.149 | 9 |  | 0.149 | 4 | 0 | -100 | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.08 | 0.727 | 65 | 4.036 | 0.721 |  | 11.4 |  | 20231106_HepCl_ReRun_246 | Blank | Blank |  |  |  | 4.036 |  |  |  |  | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.4 | 4.036 | 240 |  | 4.036 | 4 | 0 | -100
20231106_HepCl_ReRun_253 | Blank | Blank | 3.84 | 5.148 | 173 | [5 empty] | 20231106_HepCl_ReRun_253 | Blank | Blank | 3.81 | 8.129 | 294 | [5 empty] | 20231106_HepCl_ReRun_253 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_253 | Blank | Blank |  |  |  | 12.938 |  |  |  |  | 20231106_HepCl_ReRun_253 | Blank | Blank | 3.96 | 12.938 | 311 |  | 12.938 | 4 | 0 | -100 | 20231106_HepCl_ReRun_253 | Blank | Blank |  |  |  | 0.066 |  |  |  |  | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.04 | 0.066 | 4 |  | 0.066 | 4 | 0 | -100 | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.06 | 0.597 | 60 | 1.238 | 1.929 |  | 31.58 |  | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.43 | 12.221 | 1633 | 1.238 | 39.486 |  | 136.52 |  | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.4 | 1.238 | 124 |  | 1.238 | 4 | 0 | -100
20231106_HepCl_ReRun_254 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | 3.81 | 13.784 | 295 | [5 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_254 | Blank | Blank | 3.86 | 153.336 | 3207 | [5 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_254 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_254 | Blank | Blank | 4.08 | 4.24 | 210 | 7.609 | 2.229 |  | 36.59 |  | 20231106_HepCl_ReRun_254 | Blank | Blank |  |  |  | 7.609 |  |  |  |  | 20231106_HepCl_ReRun_254 | Blank | Blank | 4.39 | 7.609 | 217 |  | 7.609 | 4 | 0 | -100
20231106_HepCl_ReRun_255 | Blank | Blank | 3.86 | 7.01 | 252 | [5 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | 3.82 | 17.656 | 487 | [5 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_255 | Blank | Blank | 3.85 | 181.382 | 3786 | [5 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_255 | Blank | Blank |  |  |  | 0.647 |  |  |  |  | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.06 | 0.647 | 38 |  | 0.647 | 4 | 0 | -100 | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.01 | 0.367 | 74 | 3 | 0.489 |  | 7.54 |  | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.43 | 36.339 | 1486 | 3 | 48.452 |  | 167.65 |  | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.41 | 3 | 179 |  | 3 | 4 | 0 | -100
20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.9 | 7.372 | 255 | [5 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.82 | 33.317 | 899 | [5 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.86 | 475.916 | 8865 | [5 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.06 | 0.326 | 19 | 0.572 | 2.28 |  | 12.67 |  | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.06 | 0.572 | 55 |  | 0.572 | 4 | 0 | -100 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.01 | 0.655 | 65 | 3.127 | 0.838 |  | 13.36 |  | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.37 | 140.794 | 2229 | 3.127 | 180.101 |  | 624.63 |  | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.38 | 3.127 | 303 |  | 3.127 | 4 | 0 | -100
20231106_HepCl_ReRun_028 | Crash Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.82 | 17.326 | 404 | [5 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.86 | 244.642 | 4634 | 9.577 | 102.179 |  | 289.43 |  | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.98 | 9.577 | 349 |  | 9.577 | 4 | 0 | -100 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.08 | 0.433 | 25 | [5 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.03 | 3.303 | 150 | 4.24 | 3.116 |  | 51.41 |  | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.36 | 23.582 | 2277 | 4.24 | 22.247 |  | 76.68 |  | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.39 | 4.24 | 142 |  | 4.24 | 4 | 0 | -100
20231106_HepCl_ReRun_046 | Crash Blank | Blank | 3.82 | 2.362 | 85 | [5 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 3.81 | 26.536 | 537 | [5 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.02 | 0.155 | 15 | 0.087 | 7.126 |  | 41.75 |  | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.04 | 0.087 | 5 |  | 0.087 | 4 | 0 | -100 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.04 | 3.32 | 197 | 1.228 | 10.814 |  | 179.97 |  | 20231106_HepCl_ReRun_046 | Crash Blank | Blank |  |  |  | 1.228 |  |  |  |  | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.42 | 1.228 | 138 |  | 1.228 | 4 | 0 | -100
20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.328 |  |  |  |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.328 |  |  |  |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 3.81 | 1.328 | 48 |  | 1.328 | 4 | 0 | -100 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 3.86 | 250.409 | 4456 | 12.974 | 77.203 |  | 218.58 |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 3.97 | 12.974 | 336 |  | 12.974 | 4 | 0 | -100 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.05 | 0.073 | 18 | 0.348 | 0.839 |  | 4.03 |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.1 | 0.348 | 15 |  | 0.348 | 4 | 0 | -100 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.03 | 3.021 | 145 | 1.469 | 8.226 |  | 136.74 |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.469 |  |  |  |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.39 | 1.469 | 120 |  | 1.469 | 4 | 0 | -100
20231106_HepCl_ReRun_004 | Matrix Blank | Blank |  |  |  | 8.997 |  |  |  |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.81 | 29.492 | 926 | 8.997 | 13.112 |  | 192.16 |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.81 | 8.997 | 224 |  | 8.997 | 4 | 0.01 | -99.7 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.85 | 607.958 | 11669 | 0.638 | 3811.649 |  | 10811.72 |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.96 | 0.638 | 27 |  | 0.638 | 4 | 0 | -100 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.08 | 1.594 | 79 | 0.86 | 7.414 |  | 43.47 |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.05 | 0.86 | 50 |  | 0.86 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.07 | 1.874 | 69 | 5.073 | 1.478 |  | 24.05 |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.43 | 13.412 | 1130 | 5.073 | 10.575 |  | 36.17 |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.41 | 5.073 | 225 |  | 5.073 | 4 | 0 | -100
20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.87 | 35.183 | 628 | 1.072 | 131.28 |  | 479.23 |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.83 | 19.347 | 590 | 1.072 | 72.19 |  | 1065.19 |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.81 | 1.072 | 39 |  | 1.072 | 4 | 0 | -100 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank |  |  |  | 0.598 |  |  |  |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.07 | 0.598 | 35 |  | 0.598 | 4 | 0 | -100 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.07 | 4.099 | 215 | 16.466 | 0.996 |  | 16 |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.38 | 59.634 | 2149 | 16.466 | 14.487 |  | 49.74 |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.41 | 16.466 | 203 |  | 16.466 | 4 | 0 | -100
20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.9 | 2.746 | 126 | [5 empty] | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.81 | 16.706 | 473 | [5 empty] | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.86 | 183.34 | 3807 | 8.655 | 84.733 |  | 239.94 |  | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 4 | 8.655 | 211 |  | 8.655 | 4 | 0 | -100 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 0.278 |  |  |  |  | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 4.1 | 0.278 | 12 |  | 0.278 | 4 | 0 | -100 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 9.694 |  |  |  |  | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 9.694 |  |  |  |  | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 4.43 | 9.694 | 259 |  | 9.694 | 4 | 0 | -100
20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 0.37 |  |  |  |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.81 | 13.005 | 297 | 0.37 | 140.595 |  | 2076.03 |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.81 | 0.37 | 26 |  | 0.37 | 4 | 0 | -100 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.86 | 175.411 | 3826 | 0.454 | 1545.471 |  | 4383.47 |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4 | 0.454 | 37 |  | 0.454 | 4 | 0 | -100 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 3.149 |  |  |  |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4.02 | 3.149 | 105 |  | 3.149 | 4 | 0.01 | -99.8 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 6.959 |  |  |  |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4.43 | 0.86 | 403 | 6.959 | 0.494 |  | 1.17 |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4.37 | 6.959 | 192 |  | 6.959 | 4 | 0 | -100

## Working Sheet 2
|  |  |  | 6:3 FTCA | [9 empty] | 6:2 FTS | [11 empty] | M2-6:2 FTS | [10 empty] | 9H-PFNA | [10 empty] | M9PFNA | [9 empty] | 8:2 FTS | [11 empty] | M2-8:2 FTS | [11 empty] | 11H-PFUnDA | [10 empty] | PFUnDA | [9 empty] | M7PFUdA
Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
20231106_HepCl_ReRun_008 | CC01 | Analyte | 3.89 | 67.971 | 1011 | 3119.796 | 0.087 | 0.175 |  |  | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 3.78 | 67.876 | 1854 | 3119.796 | 0.087 | 0.175 |  |  | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 3.78 | 3119.796 | 107592 |  | 3119.796 | 4 | 3.86 | -3.6 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 3.81 | 1564.92 | 41924 | 75194.047 | 0.083 | 0.175 |  |  | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 3.96 | 75194.047 | 2065091 |  | 75194.047 | 4 | 4.54 | 13.4 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 4.05 | 72.711 | 2860 | 1404.893 | 0.207 | 0.175 | 0.23 | 33.9 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 4.05 | 1404.893 | 51799 |  | 1404.893 | 4 | 4.18 | 4.5 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 4.07 | 141.617 | 5583 | 49528.125 | 0.011 | 0.175 |  |  | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 4.4 | 2611.916 | 113399 | 49528.125 | 0.211 | 0.175 | 0.19 | 7.5 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 4.4 | 49528.125 | 2117331 |  | 49528.125 | 4 | 5.54 | 38.5
20231106_HepCl_ReRun_047 | CC01 | Analyte | 3.88 | 43.373 | 826 | 2990.745 | 0.058 | 0.175 |  |  | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 3.78 | 44.494 | 1585 | 2990.745 | 0.06 | 0.175 |  |  | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 3.78 | 2990.745 | 102310 |  | 2990.745 | 4 | 3.7 | -7.6 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 3.81 | 1474.341 | 41632 | 66380.367 | 0.089 | 0.175 |  |  | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 3.96 | 66380.367 | 1787191 |  | 66380.367 | 4 | 4 | 0.1 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 4.05 | 62.805 | 2424 | 1396.361 | 0.18 | 0.175 | 0.07 | -59.1 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 4.05 | 1396.361 | 52081 |  | 1396.361 | 4 | 4.15 | 3.8 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 4.06 | 131.287 | 5100 | 31262.33 | 0.017 | 0.175 |  |  | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 4.4 | 1618.11 | 66118 | 31262.33 | 0.207 | 0.175 | 0.17 | -0.3 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 4.4 | 31262.33 | 1253093 |  | 31262.33 | 4 | 3.5 | -12.6
20231106_HepCl_ReRun_009 | CC02 | Analyte | 3.87 | 21.358 | 767 | 3010.686 | 0.028 | 0.281 |  |  | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 3.78 | 99.02 | 2305 | 3010.686 | 0.132 | 0.281 | 0.34 | 21.5 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 3.78 | 3010.686 | 103804 |  | 3010.686 | 4 | 3.72 | -6.9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 3.81 | 2451.184 | 72295 | 75489.938 | 0.13 | 0.281 |  |  | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 3.97 | 75489.938 | 2134003 |  | 75489.938 | 4 | 4.55 | 13.9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 4.05 | 68.343 | 2531 | 1292.746 | 0.211 | 0.281 | 0.26 | -7.2 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 4.05 | 1292.746 | 46823 |  | 1292.746 | 4 | 3.85 | -3.9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 4.07 | 255.263 | 10129 | 47045.578 | 0.022 | 0.281 |  |  | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 4.4 | 2187.755 | 92047 | 47045.578 | 0.186 | 0.281 | 0.1 | -63.9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 4.4 | 47045.578 | 1943068 |  | 47045.578 | 4 | 5.26 | 31.6
20231106_HepCl_ReRun_048 | CC02 | Analyte | 3.88 | 30.829 | 1001 | 3253.131 | 0.038 | 0.281 |  |  | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 3.78 | 89.744 | 2579 | 3253.131 | 0.11 | 0.281 | 0.03 | -90 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 3.78 | 3253.131 | 112734 |  | 3253.131 | 4 | 4.02 | 0.6 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 3.81 | 2342.242 | 73223 | 68137.133 | 0.138 | 0.281 |  |  | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 3.96 | 68137.133 | 1884396 |  | 68137.133 | 4 | 4.11 | 2.8 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 4.05 | 45.842 | 1834 | 1232.084 | 0.149 | 0.281 |  |  | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 4.05 | 1232.084 | 47762 |  | 1232.084 | 4 | 3.66 | -8.4 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 4.07 | 206.615 | 7775 | 31310.842 | 0.026 | 0.281 |  |  | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 4.4 | 1399.847 | 59255 | 31310.842 | 0.179 | 0.281 | 0.08 | -72.8 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 4.4 | 31310.842 | 1311558 |  | 31310.842 | 4 | 3.5 | -12.4
20231106_HepCl_ReRun_010 | CC03 | Analyte | 3.86 | 120.148 | 2098 | 2911.38 | 0.165 | 0.444 |  |  | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 3.78 | 95.6 | 2399 | 2911.38 | 0.131 | 0.444 | 0.34 | -23.7 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 3.77 | 2911.38 | 99233 |  | 2911.38 | 4 | 3.6 | -10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 3.81 | 3388.928 | 104660 | 66257.367 | 0.205 | 0.444 | 0.17 | -62.2 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 3.96 | 66257.367 | 1850615 |  | 66257.367 | 4 | 4 | -0.1 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 4.05 | 57.947 | 2438 | 1261.171 | 0.184 | 0.444 | 0.09 | -78.6 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 4.05 | 1261.171 | 45058 |  | 1261.171 | 4 | 3.75 | -6.2 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 4.06 | 284.665 | 9000 | 40876.117 | 0.028 | 0.444 |  |  | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 4.4 | 2251.235 | 90945 | 40876.117 | 0.22 | 0.444 | 0.22 | -50.3 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 4.4 | 40876.117 | 1656387 |  | 40876.117 | 4 | 4.57 | 14.3
20231106_HepCl_ReRun_049 | CC03 | Analyte | 3.88 | 110.652 | 1988 | 2768.087 | 0.16 | 0.444 |  |  | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 3.78 | 79.804 | 2635 | 2768.087 | 0.115 | 0.444 | 0.1 | -77.1 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 3.78 | 2768.087 | 94918 |  | 2768.087 | 4 | 3.42 | -14.4 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 3.81 | 3022.632 | 95828 | 60965.863 | 0.198 | 0.444 | 0.15 | -66.2 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 3.96 | 60965.863 | 1687426 |  | 60965.863 | 4 | 3.68 | -8.1 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 4.05 | 50.07 | 2160 | 1101.421 | 0.182 | 0.444 | 0.08 | -81.3 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 4.05 | 1101.421 | 39699 |  | 1101.421 | 4 | 3.28 | -18.1 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 4.06 | 234.735 | 8332 | 28202.996 | 0.033 | 0.444 |  |  | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 4.4 | 1610.389 | 65901 | 28202.996 | 0.228 | 0.444 | 0.25 | -44 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 4.4 | 28202.996 | 1139326 |  | 28202.996 | 4 | 3.15 | -21.1
20231106_HepCl_ReRun_011 | CC04 | Analyte | 3.88 | 179.938 | 3117 | 3209.637 | 0.224 | 0.713 | 0.09 | -87.4 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 3.78 | 141.693 | 3935 | 3209.637 | 0.177 | 0.713 | 1.01 | 41.3 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 3.78 | 3209.637 | 111137 |  | 3209.637 | 4 | 3.97 | -0.8 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 3.8 | 4601.388 | 140744 | 73904.594 | 0.249 | 0.713 | 0.29 | -58.8 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 3.97 | 73904.594 | 2064188 |  | 73904.594 | 4 | 4.46 | 11.5 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 4.05 | 158.026 | 5515 | 1380.42 | 0.458 | 0.713 | 1.74 | 144.1 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 4.05 | 1380.42 | 51580 |  | 1380.42 | 4 | 4.11 | 2.7 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 4.06 | 528.606 | 20359 | 45863.699 | 0.046 | 0.713 | 0.14 | -80.5 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 4.4 | 5824.572 | 243158 | 45863.699 | 0.508 | 0.713 | 1.22 | 71.1 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 4.4 | 45863.699 | 1923637 |  | 45863.699 | 4 | 5.13 | 28.3
20231106_HepCl_ReRun_029 | CC04 | Analyte | 3.88 | 221.882 | 3836 | 2902.343 | 0.306 | 0.713 | 0.39 | -45.6 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 3.78 | 92.881 | 2263 | 2902.343 | 0.128 | 0.713 | 0.29 | -59.4 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 3.78 | 2902.343 | 99894 |  | 2902.343 | 4 | 3.59 | -10.3 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 3.81 | 4751.315 | 151755 | 72754.594 | 0.261 | 0.713 | 0.33 | -53.9 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 3.96 | 72754.594 | 1997669 |  | 72754.594 | 4 | 4.39 | 9.7 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 4.05 | 143.369 | 5623 | 1378.086 | 0.416 | 0.713 | 1.49 | 109 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 4.05 | 1378.086 | 49348 |  | 1378.086 | 4 | 4.1 | 2.5 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 4.07 | 442.919 | 18043 | 38003.336 | 0.047 | 0.713 | 0.15 | -79.3 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 4.4 | 5001.823 | 214961 | 38003.336 | 0.526 | 0.713 | 1.28 | 80.1 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 4.4 | 38003.336 | 1571169 |  | 38003.336 | 4 | 4.25 | 6.3
20231106_HepCl_ReRun_033 | CC04 | Analyte | 3.88 | 215.362 | 4509 | 2884.729 | 0.299 | 0.713 | 0.36 | -49.3 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 3.78 | 88.341 | 2578 | 2884.729 | 0.122 | 0.713 | 0.21 | -70.8 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 3.78 | 2884.729 | 98797 |  | 2884.729 | 4 | 3.57 | -10.8 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 3.81 | 5260.333 | 171299 | 73487.75 | 0.286 | 0.713 | 0.4 | -43.9 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 3.96 | 73487.75 | 2025898 |  | 73487.75 | 4 | 4.43 | 10.8 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 4.05 | 150.419 | 6273 | 1402.833 | 0.429 | 0.713 | 1.57 | 119.7 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 4.05 | 1402.833 | 55016 |  | 1402.833 | 4 | 4.17 | 4.3 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 4.07 | 430.506 | 15772 | 37383.543 | 0.046 | 0.713 | 0.14 | -80.6 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 4.4 | 4821.958 | 198195 | 37383.543 | 0.516 | 0.713 | 1.25 | 75 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 4.4 | 37383.543 | 1509787 |  | 37383.543 | 4 | 4.18 | 4.5
20231106_HepCl_ReRun_037 | CC04 | Analyte | 3.88 | 192.081 | 3259 | 2898.57 | 0.265 | 0.713 | 0.24 | -66.5 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 3.77 | 115.598 | 3475 | 2898.57 | 0.16 | 0.713 | 0.75 | 6 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 3.77 | 2898.57 | 98875 |  | 2898.57 | 4 | 3.58 | -10.4 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 3.8 | 4919.652 | 157957 | 70830.211 | 0.278 | 0.713 | 0.38 | -47.3 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 3.96 | 70830.211 | 1970689 |  | 70830.211 | 4 | 4.27 | 6.8 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 4.05 | 144.658 | 5667 | 1405.047 | 0.412 | 0.713 | 1.46 | 105.3 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 4.05 | 1405.047 | 55126 |  | 1405.047 | 4 | 4.18 | 4.5 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 4.06 | 532.507 | 21317 | 34334.141 | 0.062 | 0.713 | 0.41 | -43.1 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 4.4 | 4405.595 | 180554 | 34334.141 | 0.513 | 0.713 | 1.24 | 73.7 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 4.39 | 34334.141 | 1351660 |  | 34334.141 | 4 | 3.84 | -4
20231106_HepCl_ReRun_050 | CC04 | Analyte | 3.88 | 205.072 | 3755 | 3070.196 | 0.267 | 0.713 | 0.25 | -65.4 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 3.78 | 111.645 | 3306 | 3070.196 | 0.145 | 0.713 | 0.55 | -23.2 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 3.78 | 3070.196 | 107053 |  | 3070.196 | 4 | 3.8 | -5.1 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 3.81 | 5242.089 | 172577 | 67549.781 | 0.31 | 0.713 | 0.47 | -34.3 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 3.96 | 67549.781 | 1870626 |  | 67549.781 | 4 | 4.08 | 1.9 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 4.05 | 147.42 | 5938 | 1415.325 | 0.417 | 0.713 | 1.49 | 109.4 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 4.05 | 1415.325 | 53894 |  | 1415.325 | 4 | 4.21 | 5.2 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 4.06 | 428.645 | 14499 | 29024.73 | 0.059 | 0.713 | 0.36 | -50.1 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 4.4 | 3810.877 | 157128 | 29024.73 | 0.525 | 0.713 | 1.28 | 79.5 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 4.4 | 29024.73 | 1111171 |  | 29024.73 | 4 | 3.25 | -18.8
20231106_HepCl_ReRun_068 | CC04 | Analyte | 3.88 | 218.265 | 4061 | 3113.354 | 0.28 | 0.713 | 0.29 | -58.6 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 3.78 | 117.315 | 3438 | 3113.354 | 0.151 | 0.713 | 0.62 | -12.3 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 3.78 | 3113.354 | 106715 |  | 3113.354 | 4 | 3.85 | -3.8 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 3.81 | 5032.04 | 161812 | 70878.07 | 0.284 | 0.713 | 0.39 | -44.9 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 3.96 | 70878.07 | 1984627 |  | 70878.07 | 4 | 4.28 | 6.9 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 4.05 | 124.174 | 4828 | 1319.151 | 0.377 | 0.713 | 1.25 | 75.6 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 4.05 | 1319.151 | 49822 |  | 1319.151 | 4 | 3.92 | -1.9 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 4.06 | 422.18 | 15499 | 26842.035 | 0.063 | 0.713 | 0.42 | -41.1 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 4.4 | 3619.746 | 142326 | 26842.035 | 0.539 | 0.713 | 1.33 | 86.4 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 4.4 | 26842.035 | 1085838 |  | 26842.035 | 4 | 3 | -24.9
20231106_HepCl_ReRun_072 | CC04 | Analyte | 3.88 | 154.409 | 2781 | 2424.344 | 0.255 | 0.713 | 0.2 | -71.8 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 3.78 | 94.069 | 2660 | 2424.344 | 0.155 | 0.713 | 0.69 | -3 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 3.78 | 2424.344 | 82779 |  | 2424.344 | 4 | 3 | -25.1 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 3.81 | 4577.015 | 146341 | 68083.242 | 0.269 | 0.713 | 0.35 | -50.9 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 3.96 | 68083.242 | 1884270 |  | 68083.242 | 4 | 4.11 | 2.7 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 4.05 | 135.904 | 5456 | 1228.853 | 0.442 | 0.713 | 1.65 | 131.1 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 4.05 | 1228.853 | 46834 |  | 1228.853 | 4 | 3.66 | -8.6 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 4.07 | 372.902 | 13666 | 25236.123 | 0.059 | 0.713 | 0.36 | -50 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 4.4 | 3276.574 | 133935 | 25236.123 | 0.519 | 0.713 | 1.26 | 76.7 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 4.4 | 25236.123 | 989841 |  | 25236.123 | 4 | 2.82 | -29.4
20231106_HepCl_ReRun_076 | CC04 | Analyte | 3.88 | 181.043 | 2968 | 3100.597 | 0.234 | 0.713 | 0.12 | -82.7 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 3.78 | 90.3 | 2912 | 3100.597 | 0.116 | 0.713 | 0.12 | -83.3 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 3.78 | 3100.597 | 106576 |  | 3100.597 | 4 | 3.83 | -4.2 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 3.81 | 5230.587 | 169974 | 66670.469 | 0.314 | 0.713 | 0.48 | -33 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 3.96 | 66670.469 | 1849073 |  | 66670.469 | 4 | 4.02 | 0.6 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 4.05 | 113.447 | 4183 | 1246.689 | 0.364 | 0.713 | 1.18 | 65.1 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 4.05 | 1246.689 | 48897 |  | 1246.689 | 4 | 3.71 | -7.3 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 4.06 | 401.949 | 14818 | 24837.654 | 0.065 | 0.713 | 0.45 | -36.8 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 4.4 | 3055.906 | 131424 | 24837.654 | 0.492 | 0.713 | 1.16 | 63.4 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 4.4 | 24837.654 | 1011861 |  | 24837.654 | 4 | 2.78 | -30.5
20231106_HepCl_ReRun_012 | CC05 | Analyte | 3.85 | 264.826 | 5128 | 3047.267 | 0.348 | 1.138 | 0.54 | -52.5 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 3.77 | 160.908 | 4889 | 3047.267 | 0.211 | 1.138 | 1.52 | 33.5 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 3.77 | 3047.267 | 103322 |  | 3047.267 | 4 | 3.77 | -5.8 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 3.8 | 8336.581 | 273803 | 72824.063 | 0.458 | 1.138 | 0.89 | -22.1 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 3.96 | 72824.063 | 2060878 |  | 72824.063 | 4 | 4.39 | 9.8 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 4.05 | 145.233 | 5514 | 1318.556 | 0.441 | 1.138 | 1.64 | 43.8 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 4.05 | 1318.556 | 49054 |  | 1318.556 | 4 | 3.92 | -1.9 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 4.06 | 792.861 | 30425 | 43603.629 | 0.073 | 1.138 | 0.58 | -48.7 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 4.4 | 4827.899 | 196931 | 43603.629 | 0.443 | 1.138 | 0.99 | -12.7 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 4.4 | 43603.629 | 1755856 |  | 43603.629 | 4 | 4.88 | 21.9
20231106_HepCl_ReRun_030 | CC05 | Analyte | 3.88 | 379.244 | 6609 | 2148.496 | 0.706 | 1.138 | 1.85 | 62.7 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 3.78 | 92.564 | 2777 | 2148.496 | 0.172 | 1.138 | 0.94 | -17 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 3.78 | 2148.496 | 73093 |  | 2148.496 | 4 | 2.66 | -33.6 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 3.81 | 6234.799 | 201623 | 67752.633 | 0.368 | 1.138 | 0.63 | -44.5 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 3.97 | 67752.633 | 1891724 |  | 67752.633 | 4 | 4.09 | 2.2 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 4.05 | 121.842 | 4186 | 1331.639 | 0.366 | 1.138 | 1.19 | 4.4 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 4.05 | 1331.639 | 48209 |  | 1331.639 | 4 | 3.96 | -1 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 4.07 | 702.902 | 25764 | 35276.152 | 0.08 | 1.138 | 0.7 | -38.4 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 4.4 | 4156.563 | 170807 | 35276.152 | 0.471 | 1.138 | 1.09 | -4 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 4.4 | 35276.152 | 1423350 |  | 35276.152 | 4 | 3.95 | -1.3
20231106_HepCl_ReRun_034 | CC05 | Analyte | 3.88 | 245.912 | 4041 | 3311.456 | 0.297 | 1.138 | 0.36 | -68.7 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 3.78 | 114.997 | 4024 | 3311.456 | 0.139 | 1.138 | 0.45 | -60.4 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 3.78 | 3311.456 | 113680 |  | 3311.456 | 4 | 4.09 | 2.4 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 3.81 | 6547.865 | 214088 | 66896.563 | 0.392 | 1.138 | 0.7 | -38.6 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 3.96 | 66896.563 | 1878349 |  | 66896.563 | 4 | 4.04 | 0.9 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 4.05 | 113.359 | 4706 | 1393.114 | 0.325 | 1.138 | 0.94 | -16.9 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 4.05 | 1393.114 | 51826 |  | 1393.114 | 4 | 4.14 | 3.6 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 4.07 | 699.594 | 26547 | 33432.625 | 0.084 | 1.138 | 0.77 | -32.6 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 4.4 | 4101.203 | 162869 | 33432.625 | 0.491 | 1.138 | 1.16 | 1.9 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 4.4 | 33432.625 | 1300150 |  | 33432.625 | 4 | 3.74 | -6.5
20231106_HepCl_ReRun_038 | CC05 | Analyte | 3.86 | 288.183 | 4636 | 2369.425 | 0.487 | 1.138 | 1.05 | -7.8 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 3.78 | 120.169 | 3550 | 2369.425 | 0.203 | 1.138 | 1.4 | 22.7 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 3.78 | 2369.425 | 80566 |  | 2369.425 | 4 | 2.93 | -26.8 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 3.8 | 7019.847 | 228077 | 64723.23 | 0.434 | 1.138 | 0.82 | -28.1 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 3.96 | 64723.23 | 1829755 |  | 64723.23 | 4 | 3.9 | -2.4 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 4.05 | 129.741 | 5334 | 1291.964 | 0.402 | 1.138 | 1.4 | 23.3 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 4.05 | 1291.964 | 49703 |  | 1291.964 | 4 | 3.84 | -3.9 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 4.06 | 689.775 | 23705 | 32710.328 | 0.084 | 1.138 | 0.78 | -31.6 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 4.39 | 3882.14 | 161436 | 32710.328 | 0.475 | 1.138 | 1.1 | -3 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 4.39 | 32710.328 | 1335617 |  | 32710.328 | 4 | 3.66 | -8.5
20231106_HepCl_ReRun_051 | CC05 | Analyte | 3.88 | 305.641 | 5035 | 2763.465 | 0.442 | 1.138 | 0.89 | -22 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 3.78 | 159.707 | 4651 | 2763.465 | 0.231 | 1.138 | 1.81 | 59.4 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 3.78 | 2763.465 | 93414 |  | 2763.465 | 4 | 3.42 | -14.6 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 3.81 | 6910.878 | 225487 | 68209.023 | 0.405 | 1.138 | 0.74 | -35.2 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 3.96 | 68209.023 | 1921450 |  | 68209.023 | 4 | 4.11 | 2.9 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 4.05 | 123.641 | 5052 | 1218 | 0.406 | 1.138 | 1.43 | 25.6 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 4.05 | 1218 | 45808 |  | 1218 | 4 | 3.62 | -9.4 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 4.07 | 632.781 | 22856 | 28126.109 | 0.09 | 1.138 | 0.87 | -23.3 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 4.4 | 3269.672 | 137616 | 28126.109 | 0.465 | 1.138 | 1.07 | -5.9 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 4.4 | 28126.109 | 1122780 |  | 28126.109 | 4 | 3.15 | -21.3
20231106_HepCl_ReRun_069 | CC05 | Analyte | 3.88 | 284.915 | 4827 | 2642.521 | 0.431 | 1.138 | 0.85 | -25.6 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 3.78 | 117.511 | 3059 | 2642.521 | 0.178 | 1.138 | 1.03 | -9.8 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 3.78 | 2642.521 | 90862 |  | 2642.521 | 4 | 3.27 | -18.3 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 3.81 | 6854.112 | 229917 | 65752.898 | 0.417 | 1.138 | 0.77 | -32.3 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 3.97 | 65752.898 | 1872039 |  | 65752.898 | 4 | 3.97 | -0.8 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 4.05 | 102.804 | 3685 | 1336.031 | 0.308 | 1.138 | 0.84 | -26.3 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 4.05 | 1336.031 | 51239 |  | 1336.031 | 4 | 3.97 | -0.6 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 4.07 | 686.449 | 26862 | 25787.34 | 0.106 | 1.138 | 1.15 | 0.9 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 4.4 | 3183.645 | 129738 | 25787.34 | 0.494 | 1.138 | 1.17 | 2.9 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 4.4 | 25787.34 | 1064513 |  | 25787.34 | 4 | 2.88 | -27.9
20231106_HepCl_ReRun_073 | CC05 | Analyte | 3.88 | 256.321 | 5150 | 3320.192 | 0.309 | 1.138 | 0.4 | -65 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 3.78 | 162.413 | 5489 | 3320.192 | 0.196 | 1.138 | 1.29 | 13.3 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 3.78 | 3320.192 | 115358 |  | 3320.192 | 4 | 4.11 | 2.6 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 3.81 | 5772.026 | 189798 | 64498.293 | 0.358 | 1.138 | 0.6 | -47 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 3.96 | 64498.293 | 1769525 |  | 64498.293 | 4 | 3.89 | -2.7 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 4.05 | 102.66 | 4075 | 1286.003 | 0.319 | 1.138 | 0.91 | -20.2 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 4.05 | 1286.003 | 49431 |  | 1286.003 | 4 | 3.83 | -4.4 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 4.06 | 648.499 | 25291 | 25884.043 | 0.1 | 1.138 | 1.04 | -8.3 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 4.4 | 3061.236 | 126429 | 25884.043 | 0.473 | 1.138 | 1.1 | -3.5 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 4.4 | 25884.043 | 1063068 |  | 25884.043 | 4 | 2.9 | -27.6
20231106_HepCl_ReRun_077 | CC05 | Analyte | 3.87 | 303.546 | 4891 | 3376.185 | 0.36 | 1.138 | 0.58 | -48.6 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 3.78 | 136.129 | 4459 | 3376.185 | 0.161 | 1.138 | 0.78 | -31.4 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 3.78 | 3376.185 | 117099 |  | 3376.185 | 4 | 4.17 | 4.4 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 3.81 | 5196.095 | 167226 | 63275.16 | 0.328 | 1.138 | 0.52 | -54.4 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 3.97 | 63275.16 | 1767799 |  | 63275.16 | 4 | 3.82 | -4.6 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 4.05 | 100.041 | 3971 | 1098.067 | 0.364 | 1.138 | 1.18 | 3.6 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 4.05 | 1098.067 | 41974 |  | 1098.067 | 4 | 3.27 | -18.3 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 4.07 | 627.016 | 24357 | 23961.09 | 0.105 | 1.138 | 1.12 | -1.8 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 4.4 | 2840.368 | 113787 | 23961.09 | 0.474 | 1.138 | 1.1 | -3.1 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 4.4 | 23961.09 | 987606 |  | 23961.09 | 4 | 2.68 | -33
20231106_HepCl_ReRun_013 | CC06 | Standard | 3.88 | 447.461 | 8778 | 3034.11 | 0.59 | 1.825 | 1.43 | -21.8 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.78 | 171.963 | 5263 | 3034.11 | 0.227 | 1.825 | 1.75 | -4.2 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.77 | 3034.11 | 104487 |  | 3034.11 | 4 | 3.75 | -6.2 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.8 | 11641.195 | 382511 | 68219.305 | 0.683 | 1.825 | 1.52 | -16.5 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.96 | 68219.305 | 1940316 |  | 68219.305 | 4 | 4.12 | 2.9 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.05 | 183.985 | 6873 | 1285.358 | 0.573 | 1.825 | 2.43 | 33 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.05 | 1285.358 | 48475 |  | 1285.358 | 4 | 3.82 | -4.4 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.06 | 1062.268 | 36328 | 40008.617 | 0.106 | 1.825 | 1.14 | -37.4 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.4 | 7043 | 277350 | 40008.617 | 0.704 | 1.825 | 1.9 | 4.1 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.4 | 40008.617 | 1617090 |  | 40008.617 | 4 | 4.48 | 11.9
20231106_HepCl_ReRun_031 | CC06 | Standard | 3.88 | 478.13 | 8539 | 2368.634 | 0.807 | 1.825 | 2.22 | 21.7 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.78 | 168.791 | 5259 | 2368.634 | 0.285 | 1.825 | 2.61 | 43 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.78 | 2368.634 | 80581 |  | 2368.634 | 4 | 2.93 | -26.8 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.81 | 10848.09 | 359772 | 65001.293 | 0.668 | 1.825 | 1.48 | -18.9 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.96 | 65001.293 | 1804041 |  | 65001.293 | 4 | 3.92 | -2 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.05 | 134.609 | 4719 | 1266.854 | 0.425 | 1.825 | 1.54 | -15.5 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.05 | 1266.854 | 47688 |  | 1266.854 | 4 | 3.77 | -5.8 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.07 | 1106.86 | 42580 | 33664.371 | 0.132 | 1.825 | 1.57 | -14.2 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.4 | 5588.62 | 236229 | 33664.371 | 0.664 | 1.825 | 1.76 | -3.5 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.4 | 33664.371 | 1373684 |  | 33664.371 | 4 | 3.77 | -5.9
20231106_HepCl_ReRun_035 | CC06 | Standard | 3.87 | 418.264 | 7300 | 2974.036 | 0.563 | 1.825 | 1.33 | -27.3 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.78 | 141.014 | 4907 | 2974.036 | 0.19 | 1.825 | 1.2 | -34.2 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.77 | 2974.036 | 102368 |  | 2974.036 | 4 | 3.68 | -8.1 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.8 | 11571.36 | 384229 | 66638.734 | 0.695 | 1.825 | 1.56 | -14.7 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.96 | 66638.734 | 1854201 |  | 66638.734 | 4 | 4.02 | 0.5 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.05 | 139.414 | 5172 | 1260.427 | 0.442 | 1.825 | 1.65 | -9.8 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.05 | 1260.427 | 46864 |  | 1260.427 | 4 | 3.75 | -6.3 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.06 | 1030.858 | 39085 | 31861.572 | 0.129 | 1.825 | 1.53 | -16.1 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.4 | 5433.467 | 207631 | 31861.572 | 0.682 | 1.825 | 1.82 | -0.1 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.4 | 31861.572 | 1318096 |  | 31861.572 | 4 | 3.56 | -10.9
20231106_HepCl_ReRun_039 | CC06 | Standard | 3.88 | 523.429 | 9106 | 2930.851 | 0.714 | 1.825 | 1.88 | 3.1 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.78 | 177.583 | 5236 | 2930.851 | 0.242 | 1.825 | 1.98 | 8.4 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.78 | 2930.851 | 99746 |  | 2930.851 | 4 | 3.62 | -9.4 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.81 | 11606.056 | 384908 | 61013.957 | 0.761 | 1.825 | 1.75 | -4.3 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.96 | 61013.957 | 1674807 |  | 61013.957 | 4 | 3.68 | -8 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.05 | 140.109 | 4945 | 1310.875 | 0.428 | 1.825 | 1.56 | -14.7 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.05 | 1310.875 | 49191 |  | 1310.875 | 4 | 3.9 | -2.5 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.07 | 1048.295 | 38562 | 30194.342 | 0.139 | 1.825 | 1.69 | -7.5 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.4 | 5458.212 | 220757 | 30194.342 | 0.723 | 1.825 | 1.97 | 7.7 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.4 | 30194.342 | 1255627 |  | 30194.342 | 4 | 3.38 | -15.6
20231106_HepCl_ReRun_052 | CC06 | Standard | 3.88 | 452.235 | 7649 | 2982.801 | 0.606 | 1.825 | 1.49 | -18.5 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.78 | 157.768 | 4759 | 2982.801 | 0.212 | 1.825 | 1.52 | -16.5 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.78 | 2982.801 | 103645 |  | 2982.801 | 4 | 3.69 | -7.8 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.81 | 11166.135 | 371375 | 64340.457 | 0.694 | 1.825 | 1.56 | -14.7 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.96 | 64340.457 | 1828969 |  | 64340.457 | 4 | 3.88 | -3 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.05 | 146.812 | 5980 | 1179.859 | 0.498 | 1.825 | 1.98 | 8.4 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.05 | 1179.859 | 41926 |  | 1179.859 | 4 | 3.51 | -12.3 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.06 | 1061.033 | 39434 | 27641.32 | 0.154 | 1.825 | 1.93 | 5.9 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.4 | 4747.757 | 201851 | 27641.32 | 0.687 | 1.825 | 1.84 | 0.9 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.4 | 27641.32 | 1122443 |  | 27641.32 | 4 | 3.09 | -22.7
20231106_HepCl_ReRun_070 | CC06 | Standard | 3.88 | 472.381 | 7598 | 3064.16 | 0.617 | 1.825 | 1.52 | -16.5 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.78 | 156.739 | 4706 | 3064.16 | 0.205 | 1.825 | 1.42 | -22.1 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.78 | 3064.16 | 104987 |  | 3064.16 | 4 | 3.79 | -5.3 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.81 | 11005.288 | 365567 | 63387.512 | 0.694 | 1.825 | 1.56 | -14.7 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.96 | 63387.512 | 1742378 |  | 63387.512 | 4 | 3.82 | -4.4 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.05 | 120.27 | 4640 | 1220.534 | 0.394 | 1.825 | 1.36 | -25.6 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.05 | 1220.534 | 47238 |  | 1220.534 | 4 | 3.63 | -9.2 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.06 | 1028.637 | 39775 | 26664.004 | 0.154 | 1.825 | 1.95 | 6.6 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.4 | 4767.315 | 194365 | 26664.004 | 0.715 | 1.825 | 1.94 | 6.2 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.4 | 26664.004 | 1080929 |  | 26664.004 | 4 | 2.98 | -25.4
20231106_HepCl_ReRun_074 | CC06 | Standard | 3.88 | 432.128 | 7512 | 3262.376 | 0.53 | 1.825 | 1.21 | -33.9 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.78 | 145.769 | 5063 | 3262.376 | 0.179 | 1.825 | 1.04 | -43.1 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.77 | 3262.376 | 111380 |  | 3262.376 | 4 | 4.03 | 0.8 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.81 | 10255.913 | 336728 | 62066.836 | 0.661 | 1.825 | 1.46 | -19.9 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.96 | 62066.836 | 1728402 |  | 62066.836 | 4 | 3.74 | -6.4 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.05 | 143.867 | 6016 | 1181.049 | 0.487 | 1.825 | 1.92 | 5 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.05 | 1181.049 | 43690 |  | 1181.049 | 4 | 3.51 | -12.2 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.06 | 967.527 | 35145 | 27006.176 | 0.143 | 1.825 | 1.76 | -3.4 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.4 | 4664.719 | 193278 | 27006.176 | 0.691 | 1.825 | 1.85 | 1.6 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.4 | 27006.176 | 1109378 |  | 27006.176 | 4 | 3.02 | -24.5
20231106_HepCl_ReRun_078 | CC06 | Standard | 3.88 | 465.458 | 7889 | 3313.979 | 0.562 | 1.825 | 1.32 | -27.5 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.78 | 168.728 | 5698 | 3313.979 | 0.204 | 1.825 | 1.41 | -22.9 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.77 | 3313.979 | 112946 |  | 3313.979 | 4 | 4.1 | 2.4 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.8 | 11621.744 | 390664 | 61289.246 | 0.758 | 1.825 | 1.74 | -4.7 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.96 | 61289.246 | 1692146 |  | 61289.246 | 4 | 3.7 | -7.6 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.05 | 129.094 | 5605 | 1218.418 | 0.424 | 1.825 | 1.53 | -15.9 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.05 | 1218.418 | 46861 |  | 1218.418 | 4 | 3.62 | -9.4 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.06 | 993.691 | 37872 | 24890.996 | 0.16 | 1.825 | 2.04 | 11.6 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.39 | 4229.273 | 162426 | 24890.996 | 0.68 | 1.825 | 1.82 | -0.5 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.4 | 24890.996 | 992358 |  | 24890.996 | 4 | 2.78 | -30.4
20231106_HepCl_ReRun_015 | CC07 | Standard | 3.88 | 779.956 | 11885 | 2993.711 | 1.042 | 2.919 | 3.08 | 5.5 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.78 | 238.272 | 7112 | 2993.711 | 0.318 | 2.919 | 3.1 | 6.3 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.78 | 2993.711 | 101914 |  | 2993.711 | 4 | 3.7 | -7.5 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.81 | 21097.494 | 703071 | 64042.156 | 1.318 | 2.919 | 3.33 | 13.9 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.96 | 64042.156 | 1714197 |  | 64042.156 | 4 | 3.86 | -3.4 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.05 | 221.348 | 8806 | 1359.29 | 0.651 | 2.919 | 2.9 | -0.6 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.05 | 1359.29 | 50339 |  | 1359.29 | 4 | 4.04 | 1.1 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.06 | 1882.456 | 72625 | 42977.938 | 0.175 | 2.919 | 2.3 | -21.4 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.4 | 10383.962 | 423591 | 42977.938 | 0.966 | 2.919 | 2.81 | -3.7 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.4 | 42977.938 | 1740176 |  | 42977.938 | 4 | 4.81 | 20.2
20231106_HepCl_ReRun_054 | CC07 | Standard | 3.88 | 712.941 | 14349 | 3350.341 | 0.851 | 2.919 | 2.38 | -18.4 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.78 | 193.975 | 6727 | 3350.341 | 0.232 | 2.919 | 1.82 | -37.7 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.77 | 3350.341 | 114845 |  | 3350.341 | 4 | 4.14 | 3.6 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.8 | 20381.063 | 685582 | 67724.766 | 1.204 | 2.919 | 3 | 2.8 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.96 | 67724.766 | 1893346 |  | 67724.766 | 4 | 4.09 | 2.1 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.05 | 193.381 | 7888 | 1307.154 | 0.592 | 2.919 | 2.54 | -12.9 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.05 | 1307.154 | 48444 |  | 1307.154 | 4 | 3.89 | -2.8 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.06 | 1728.86 | 67243 | 30703.615 | 0.225 | 2.919 | 3.13 | 7.3 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.39 | 6970.027 | 276378 | 30703.615 | 0.908 | 2.919 | 2.61 | -10.7 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.39 | 30703.615 | 1272896 |  | 30703.615 | 4 | 3.43 | -14.1
20231106_HepCl_ReRun_016 | CC08 | Standard | 3.88 | 1201.809 | 19563 | 3101.963 | 1.55 | 4.669 | 4.94 | 5.7 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.78 | 347.814 | 11862 | 3101.963 | 0.449 | 4.669 | 5.03 | 7.6 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.78 | 3101.963 | 107129 |  | 3101.963 | 4 | 3.84 | -4.1 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.81 | 29158.406 | 967248 | 65068.902 | 1.792 | 4.669 | 4.67 | 0.1 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.97 | 65068.902 | 1781481 |  | 65068.902 | 4 | 3.93 | -1.9 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.05 | 362.115 | 14411 | 1397.361 | 1.037 | 4.669 | 5.21 | 11.6 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.05 | 1397.361 | 52204 |  | 1397.361 | 4 | 4.16 | 3.9 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.06 | 3141.384 | 120411 | 44416.688 | 0.283 | 4.669 | 4.09 | -12.3 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.4 | 16841.48 | 684362 | 44416.688 | 1.517 | 4.669 | 4.72 | 1.1 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.4 | 44416.688 | 1823227 |  | 44416.688 | 4 | 4.97 | 24.2
20231106_HepCl_ReRun_055 | CC08 | Standard | 3.87 | 1077.497 | 18622 | 2541.916 | 1.696 | 4.669 | 5.47 | 17.1 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.78 | 335.471 | 11180 | 2541.916 | 0.528 | 4.669 | 6.2 | 32.8 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.78 | 2541.916 | 86105 |  | 2541.916 | 4 | 3.14 | -21.4 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.81 | 26398.207 | 865225 | 64294.883 | 1.642 | 4.669 | 4.25 | -9.1 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.96 | 64294.883 | 1769355 |  | 64294.883 | 4 | 3.88 | -3 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.05 | 279.907 | 10543 | 1383.121 | 0.809 | 4.669 | 3.85 | -17.6 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.05 | 1383.121 | 50458 |  | 1383.121 | 4 | 4.11 | 2.9 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.06 | 2842.049 | 102842 | 30452.596 | 0.373 | 4.669 | 5.6 | 20 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.4 | 11831.235 | 485726 | 30452.596 | 1.554 | 4.669 | 4.85 | 3.9 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.4 | 30452.596 | 1229852 |  | 30452.596 | 4 | 3.41 | -14.8
20231106_HepCl_ReRun_017 | CC09 | Standard | 3.88 | 1982.509 | 35805 | 3138.186 | 2.527 | 7.45 | 8.51 | 14.2 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.78 | 584.512 | 18989 | 3138.186 | 0.745 | 7.45 | 9.41 | 26.3 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.78 | 3138.186 | 108340 |  | 3138.186 | 4 | 3.88 | -3 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.81 | 48933.699 | 1647090 | 70109.063 | 2.792 | 7.45 | 7.51 | 0.8 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.96 | 70109.063 | 1960652 |  | 70109.063 | 4 | 4.23 | 5.7 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.05 | 517.63 | 19667 | 1364.509 | 1.517 | 7.45 | 8.1 | 8.7 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.05 | 1364.509 | 50516 |  | 1364.509 | 4 | 4.06 | 1.5 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.07 | 4509.686 | 158142 | 41387.938 | 0.436 | 7.45 | 6.65 | -10.8 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.4 | 23083.746 | 919942 | 41387.938 | 2.231 | 7.45 | 7.2 | -3.4 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.4 | 41387.938 | 1671639 |  | 41387.938 | 4 | 4.63 | 15.7
20231106_HepCl_ReRun_056 | CC09 | Standard | 3.87 | 2015.593 | 34908 | 3151.321 | 2.558 | 7.45 | 8.62 | 15.7 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.78 | 445.528 | 14562 | 3151.321 | 0.566 | 7.45 | 6.75 | -9.3 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.78 | 3151.321 | 108098 |  | 3151.321 | 4 | 3.9 | -2.6 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.81 | 46315.258 | 1565880 | 63195.191 | 2.932 | 7.45 | 7.9 | 6.1 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.96 | 63195.191 | 1738904 |  | 63195.191 | 4 | 3.81 | -4.7 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.05 | 502.933 | 19943 | 1332.81 | 1.509 | 7.45 | 8.05 | 8 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.05 | 1332.81 | 48860 |  | 1332.81 | 4 | 3.96 | -0.9 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.06 | 4361.453 | 168909 | 29372.057 | 0.594 | 7.45 | 9.29 | 24.7 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.4 | 16286.906 | 675195 | 29372.057 | 2.218 | 7.45 | 7.16 | -4 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.4 | 29372.057 | 1219939 |  | 29372.057 | 4 | 3.29 | -17.9
20231106_HepCl_ReRun_018 | CC10 | Standard | 3.87 | 2973.767 | 47181 | 3163.276 | 3.76 | 11.875 | 13.02 | 9.6 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.78 | 761.471 | 26224 | 3163.276 | 0.963 | 11.875 | 12.63 | 6.3 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.77 | 3163.276 | 108367 |  | 3163.276 | 4 | 3.91 | -2.2 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.81 | 79207.258 | 2656081 | 69572.93 | 4.554 | 11.875 | 12.5 | 5.3 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.96 | 69572.93 | 1929098 |  | 69572.93 | 4 | 4.2 | 4.9 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.05 | 764.35 | 28612 | 1367.205 | 2.236 | 11.875 | 12.41 | 4.5 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.05 | 1367.205 | 51313 |  | 1367.205 | 4 | 4.07 | 1.7 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.06 | 7758.193 | 294767 | 43287.93 | 0.717 | 11.875 | 11.34 | -4.5 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.4 | 36668.535 | 1499343 | 43287.93 | 3.388 | 11.875 | 11.22 | -5.5 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.4 | 43287.93 | 1755786 |  | 43287.93 | 4 | 4.84 | 21.1
20231106_HepCl_ReRun_057 | CC10 | Standard | 3.88 | 3366.109 | 59308 | 3307.122 | 4.071 | 11.875 | 14.15 | 19.2 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.78 | 725.711 | 25169 | 3307.122 | 0.878 | 11.875 | 11.37 | -4.3 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.78 | 3307.122 | 114196 |  | 3307.122 | 4 | 4.09 | 2.2 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.81 | 75334.328 | 2526905 | 63700.488 | 4.731 | 11.875 | 13.01 | 9.5 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.96 | 63700.488 | 1730554 |  | 63700.488 | 4 | 3.84 | -3.9 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.05 | 682.924 | 25853 | 1340.69 | 2.038 | 11.875 | 11.22 | -5.5 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.05 | 1340.69 | 49527 |  | 1340.69 | 4 | 3.99 | -0.3 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.07 | 6099.027 | 213476 | 28786.756 | 0.847 | 11.875 | 13.52 | 13.9 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.4 | 23608.922 | 980457 | 28786.756 | 3.281 | 11.875 | 10.84 | -8.7 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.4 | 28786.756 | 1209493 |  | 28786.756 | 4 | 3.22 | -19.5
20231106_HepCl_ReRun_019 | CC11 | Standard | 3.88 | 5696.95 | 94449 | 3229.261 | 7.057 | 19.125 | 25.07 | 31.1 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.78 | 1328.398 | 45730 | 3229.261 | 1.645 | 19.125 | 22.71 | 18.8 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.78 | 3229.261 | 110694 |  | 3229.261 | 4 | 3.99 | -0.2 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.81 | 161645.313 | 5450768 | 72849.406 | 8.876 | 19.125 | 24.76 | 29.5 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.96 | 72849.406 | 2016856 |  | 72849.406 | 4 | 4.39 | 9.9 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.05 | 1354.45 | 51659 | 1491.731 | 3.632 | 19.125 | 20.78 | 8.7 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.05 | 1491.731 | 55680 |  | 1491.731 | 4 | 4.44 | 10.9 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.07 | 13921.18 | 525530 | 43650.934 | 1.276 | 19.125 | 20.67 | 8.1 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.4 | 63299.324 | 2670024 | 43650.934 | 5.801 | 19.125 | 19.59 | 2.4 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.4 | 43650.934 | 1729358 |  | 43650.934 | 4 | 4.88 | 22.1
20231106_HepCl_ReRun_058 | CC11 | Standard | 3.88 | 5013.81 | 76829 | 3103.845 | 6.461 | 19.125 | 22.89 | 19.7 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.78 | 1252.312 | 42857 | 3103.845 | 1.614 | 19.125 | 22.25 | 16.3 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.78 | 3103.845 | 107523 |  | 3103.845 | 4 | 3.84 | -4.1 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.81 | 152744.953 | 5169925 | 68771.055 | 8.884 | 19.125 | 24.79 | 29.6 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.96 | 68771.055 | 1866345 |  | 68771.055 | 4 | 4.15 | 3.7 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.05 | 1244.492 | 46411 | 1394.592 | 3.569 | 19.125 | 20.41 | 6.7 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.05 | 1394.592 | 50530 |  | 1394.592 | 4 | 4.15 | 3.7 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.06 | 12133.325 | 451289 | 31747.754 | 1.529 | 19.125 | 24.9 | 30.2 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.4 | 45632.473 | 1810515 | 31747.754 | 5.749 | 19.125 | 19.41 | 1.5 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.4 | 31747.754 | 1296720 |  | 31747.754 | 4 | 3.55 | -11.2
20231106_HepCl_ReRun_020 | CC12 | Standard | 3.87 | 6181.179 | 122226 | 2767.938 | 8.933 | 30.5 | 31.93 | 4.7 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.77 | 1914.911 | 65185 | 2767.938 | 2.767 | 30.5 | 39.29 | 28.8 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.77 | 2767.938 | 94438 |  | 2767.938 | 4 | 3.42 | -14.4 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.8 | 217131.609 | 7364952 | 60371.941 | 14.386 | 30.5 | 40.4 | 32.4 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.96 | 60371.941 | 1653881 |  | 60371.941 | 4 | 3.64 | -8.9 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.05 | 2082.749 | 78265 | 1313.037 | 6.345 | 30.5 | 37.06 | 21.5 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.05 | 1313.037 | 49232 |  | 1313.037 | 4 | 3.91 | -2.4 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.06 | 18356.199 | 719052 | 43166.441 | 1.701 | 30.5 | 27.78 | -8.9 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.39 | 110621.594 | 4408315 | 43166.441 | 10.251 | 30.5 | 35.04 | 14.9 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.4 | 43166.441 | 1736996 |  | 43166.441 | 4 | 4.83 | 20.7
20231106_HepCl_ReRun_059 | CC12 | Standard | 3.88 | 8168.249 | 142368 | 3365.248 | 9.709 | 30.5 | 34.77 | 14 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.78 | 2141.835 | 73548 | 3365.248 | 2.546 | 30.5 | 36.02 | 18.1 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.77 | 3365.248 | 115566 |  | 3365.248 | 4 | 4.16 | 4 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.81 | 215742.5 | 7243975 | 69718.367 | 12.378 | 30.5 | 34.7 | 13.8 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.96 | 69718.367 | 1978759 |  | 69718.367 | 4 | 4.21 | 5.1 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.05 | 2029.725 | 77604 | 1381.883 | 5.875 | 30.5 | 34.24 | 12.3 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.05 | 1381.883 | 52583 |  | 1381.883 | 4 | 4.11 | 2.8 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.06 | 19046.547 | 718133 | 31395.316 | 2.427 | 30.5 | 39.89 | 30.8 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.4 | 81080.563 | 3271567 | 31395.316 | 10.33 | 30.5 | 35.31 | 15.8 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.4 | 31395.316 | 1232338 |  | 31395.316 | 4 | 3.51 | -12.2
20231106_HepCl_ReRun_086 | CC12 | Analyte | 3.88 | 8971.841 | 138720 | 3342.74 | 10.736 | 30.5 | 38.52 | 26.3 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 3.78 | 2152.861 | 74098 | 3342.74 | 2.576 | 30.5 | 36.47 | 19.6 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 3.77 | 3342.74 | 115401 |  | 3342.74 | 4 | 4.13 | 3.3 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 3.81 | 163088.453 | 5391557 | 70980.813 | 9.191 | 30.5 | 25.66 | -15.9 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 3.96 | 70980.813 | 1997170 |  | 70980.813 | 4 | 4.28 | 7.1 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 4.05 | 1957.167 | 73926 | 1311.876 | 5.968 | 30.5 | 34.79 | 14.1 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 4.05 | 1311.876 | 49396 |  | 1311.876 | 4 | 3.9 | -2.4 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 4.06 | 18915.379 | 692422 | 25853.348 | 2.927 | 30.5 | 48.24 | 58.2 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 4.4 | 64428.176 | 2543358 | 25853.348 | 9.968 | 30.5 | 34.06 | 11.7 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 4.4 | 25853.348 | 1085943 |  | 25853.348 | 4 | 2.89 | -27.7
20231106_HepCl_ReRun_109 | CC12 | Analyte | 3.88 | 7950.584 | 118420 | 3304.299 | 9.625 | 30.5 | 34.46 | 13 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 3.78 | 2132.004 | 73766 | 3304.299 | 2.581 | 30.5 | 36.54 | 19.8 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 3.78 | 3304.299 | 114603 |  | 3304.299 | 4 | 4.09 | 2.1 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 3.81 | 203688.5 | 6846477 | 68778 | 11.846 | 30.5 | 33.19 | 8.8 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 3.96 | 68778 | 1876567 |  | 68778 | 4 | 4.15 | 3.7 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 4.05 | 1863.931 | 71024 | 1289.396 | 5.782 | 30.5 | 33.68 | 10.4 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 4.05 | 1289.396 | 48638 |  | 1289.396 | 4 | 3.84 | -4.1 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 4.06 | 18737.344 | 706695 | 23824.189 | 3.146 | 30.5 | 51.91 | 70.2 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 4.4 | 57135.223 | 2375849 | 23824.189 | 9.593 | 30.5 | 32.75 | 7.4 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 4.4 | 23824.189 | 1007547 |  | 23824.189 | 4 | 2.67 | -33.4
20231106_HepCl_ReRun_132 | CC12 | Analyte | 3.87 | 7783.227 | 105462 | 3604.25 | 8.638 | 30.5 | 30.85 | 1.1 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 3.78 | 2330.145 | 80404 | 3604.25 | 2.586 | 30.5 | 36.61 | 20 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 3.78 | 3604.25 | 123885 |  | 3604.25 | 4 | 4.46 | 11.4 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 3.81 | 202991.078 | 6881815 | 65707.844 | 12.357 | 30.5 | 34.64 | 13.6 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 3.96 | 65707.844 | 1818297 |  | 65707.844 | 4 | 3.96 | -0.9 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 4.05 | 1845.258 | 70940 | 1187.432 | 6.216 | 30.5 | 36.28 | 19 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 4.05 | 1187.432 | 45510 |  | 1187.432 | 4 | 3.53 | -11.7 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 4.06 | 17263.254 | 649328 | 21559.432 | 3.203 | 30.5 | 52.86 | 73.3 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 4.4 | 52353.785 | 2100787 | 21559.432 | 9.713 | 30.5 | 33.17 | 8.8 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 4.4 | 21559.432 | 889577 |  | 21559.432 | 4 | 2.41 | -39.7
20231106_HepCl_ReRun_141 | CC12 | Analyte | 3.87 | 7408.023 | 99582 | 3653.36 | 8.111 | 30.5 | 28.92 | -5.2 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 3.78 | 2373.712 | 81651 | 3653.36 | 2.599 | 30.5 | 36.8 | 20.7 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 3.77 | 3653.36 | 124641 |  | 3653.36 | 4 | 4.52 | 12.9 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 3.8 | 205355.906 | 6923133 | 64887.246 | 12.659 | 30.5 | 35.5 | 16.4 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 3.96 | 64887.246 | 1825407 |  | 64887.246 | 4 | 3.91 | -2.1 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 4.05 | 1715.24 | 65787 | 1154.56 | 5.942 | 30.5 | 34.64 | 13.6 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 4.05 | 1154.56 | 44383 |  | 1154.56 | 4 | 3.43 | -14.1 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 4.06 | 18161.553 | 695896 | 19427.527 | 3.739 | 30.5 | 61.82 | 102.7 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 4.4 | 46493.293 | 1956441 | 19427.527 | 9.573 | 30.5 | 32.69 | 7.2 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 4.4 | 19427.527 | 791137 |  | 19427.527 | 4 | 2.17 | -45.7
20231106_HepCl_ReRun_167 | CC12 | Analyte | 3.87 | 6318.074 | 98306 | 3208.256 | 7.877 | 30.5 | 28.07 | -8 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 3.78 | 1761.065 | 59448 | 3208.256 | 2.196 | 30.5 | 30.84 | 1.1 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 3.78 | 3208.256 | 110436 |  | 3208.256 | 4 | 3.97 | -0.8 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 3.8 | 164249.453 | 5518383 | 62184.316 | 10.565 | 30.5 | 29.56 | -3.1 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 3.96 | 62184.316 | 1723499 |  | 62184.316 | 4 | 3.75 | -6.2 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 4.05 | 1524.464 | 59642 | 1024.203 | 5.954 | 30.5 | 34.71 | 13.8 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 4.05 | 1024.203 | 38095 |  | 1024.203 | 4 | 3.05 | -23.8 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 4.06 | 15258.011 | 543415 | 19198.016 | 3.179 | 30.5 | 52.46 | 72 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 4.4 | 46230.961 | 1938281 | 19198.016 | 9.632 | 30.5 | 32.89 | 7.8 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 4.4 | 19198.016 | 739716 |  | 19198.016 | 4 | 2.15 | -46.3
20231106_HepCl_ReRun_190 | CC12 | Analyte | 3.87 | 3027.379 | 74803 | 3623.284 | 3.342 | 30.5 | 11.49 | -62.3 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 3.78 | 2071.804 | 70999 | 3623.284 | 2.287 | 30.5 | 32.2 | 5.6 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 3.78 | 3623.284 | 125981 |  | 3623.284 | 4 | 4.48 | 12 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 3.81 | 187763.328 | 6425514 | 65820.523 | 11.411 | 30.5 | 31.95 | 4.8 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 3.96 | 65820.523 | 1851785 |  | 65820.523 | 4 | 3.97 | -0.7 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 4.05 | 1400.518 | 54563 | 996.986 | 5.619 | 30.5 | 32.7 | 7.2 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 4.05 | 996.986 | 38354 |  | 996.986 | 4 | 2.97 | -25.9 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 4.07 | 16958.822 | 643563 | 21716.965 | 3.124 | 30.5 | 51.53 | 69 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 4.4 | 50770.645 | 1929792 | 21716.965 | 9.351 | 30.5 | 31.92 | 4.6 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 4.4 | 21716.965 | 892153 |  | 21716.965 | 4 | 2.43 | -39.3
20231106_HepCl_ReRun_199 | CC12 | Analyte | 3.87 | 2826.421 | 68207 | 2966.934 | 3.811 | 30.5 | 13.2 | -56.7 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 3.78 | 2376.775 | 81275 | 2966.934 | 3.204 | 30.5 | 45.75 | 50 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 3.78 | 2966.934 | 101754 |  | 2966.934 | 4 | 3.67 | -8.3 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 3.81 | 180207.031 | 6069535 | 51659.656 | 13.953 | 30.5 | 39.17 | 28.4 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 3.97 | 51659.656 | 1392823 |  | 51659.656 | 4 | 3.12 | -22.1 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 4.05 | 1391.144 | 54588 | 1022.509 | 5.442 | 30.5 | 31.64 | 3.7 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 4.05 | 1022.509 | 39451 |  | 1022.509 | 4 | 3.04 | -24 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 4.07 | 17834.35 | 672807 | 19258.826 | 3.704 | 30.5 | 61.23 | 100.7 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 4.4 | 46421.027 | 1889998 | 19258.826 | 9.642 | 30.5 | 32.92 | 7.9 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 4.4 | 19258.826 | 790180 |  | 19258.826 | 4 | 2.15 | -46.1
20231106_HepCl_ReRun_222 | CC12 | Analyte | 3.87 | 2645.947 | 69209 | 3886.825 | 2.723 | 30.5 | 9.22 | -69.8 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 3.78 | 2442.702 | 84129 | 3886.825 | 2.514 | 30.5 | 35.55 | 16.5 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 3.78 | 3886.825 | 134191 |  | 3886.825 | 4 | 4.81 | 20.1 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 3.81 | 190585.047 | 6503453 | 59542.715 | 12.803 | 30.5 | 35.91 | 17.7 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 3.96 | 59542.715 | 1668513 |  | 59542.715 | 4 | 3.59 | -10.2 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 4.05 | 1477.327 | 58320 | 1040.056 | 5.682 | 30.5 | 33.08 | 8.5 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 4.05 | 1040.056 | 41836 |  | 1040.056 | 4 | 3.09 | -22.7 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 4.07 | 17293.111 | 658177 | 19805.637 | 3.493 | 30.5 | 57.69 | 89.2 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 4.4 | 48423.484 | 2028444 | 19805.637 | 9.78 | 30.5 | 33.4 | 9.5 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 4.4 | 19805.637 | 827016 |  | 19805.637 | 4 | 2.22 | -44.6
20231106_HepCl_ReRun_245 | CC12 | Analyte | 3.86 | 2143.984 | 53680 | 3753.745 | 2.285 | 30.5 | 7.62 | -75 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 3.78 | 2361.131 | 81590 | 3753.745 | 2.516 | 30.5 | 35.58 | 16.6 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 3.78 | 3753.745 | 129845 |  | 3753.745 | 4 | 4.64 | 16 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 3.81 | 152733.531 | 5120461 | 53564.039 | 11.406 | 30.5 | 31.94 | 4.7 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 3.97 | 53564.039 | 1450588 |  | 53564.039 | 4 | 3.23 | -19.2 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 4.05 | 1451.808 | 56361 | 952.141 | 6.099 | 30.5 | 35.58 | 16.7 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 4.05 | 952.141 | 36160 |  | 952.141 | 4 | 2.83 | -29.2 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 4.07 | 18489.309 | 735656 | 20601.318 | 3.59 | 30.5 | 59.32 | 94.5 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 4.4 | 50404.621 | 1991044 | 20601.318 | 9.787 | 30.5 | 33.43 | 9.6 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 4.4 | 20601.318 | 819968 |  | 20601.318 | 4 | 2.3 | -42.4
20231106_HepCl_ReRun_022 | CC13 | Standard | 3.88 | 11224.8 | 178932 | 3439.733 | 13.053 | 48.75 | 46.99 | -3.6 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.78 | 3139.108 | 108007 | 3439.733 | 3.65 | 48.75 | 52.34 | 7.4 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.78 | 3439.733 | 118429 |  | 3439.733 | 4 | 4.25 | 6.3 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.81 | 279932.938 | 9373593 | 64329.785 | 17.406 | 48.75 | 48.96 | 0.4 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.96 | 64329.785 | 1726832 |  | 64329.785 | 4 | 3.88 | -3 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.05 | 2907.99 | 115954 | 1391.243 | 8.361 | 48.75 | 49.15 | 0.8 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.05 | 1391.243 | 51807 |  | 1391.243 | 4 | 4.14 | 3.5 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.06 | 26155.953 | 903097 | 44292.91 | 2.362 | 48.75 | 38.82 | -20.4 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.4 | 142737.063 | 5848686 | 44292.91 | 12.89 | 48.75 | 44.2 | -9.3 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.4 | 44292.91 | 1771998 |  | 44292.91 | 4 | 4.95 | 23.9
20231106_HepCl_ReRun_061 | CC13 | Standard | 3.88 | 12265.34 | 222118 | 3673.549 | 13.355 | 48.75 | 48.1 | -1.3 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.78 | 3087.364 | 107125 | 3673.549 | 3.362 | 48.75 | 48.08 | -1.4 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.78 | 3673.549 | 126337 |  | 3673.549 | 4 | 4.54 | 13.6 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.8 | 275087.375 | 9279365 | 69190.055 | 15.903 | 48.75 | 44.7 | -8.3 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.96 | 69190.055 | 1947788 |  | 69190.055 | 4 | 4.17 | 4.4 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.05 | 2793.315 | 105829 | 1344.772 | 8.309 | 48.75 | 48.84 | 0.2 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.05 | 1344.772 | 48392 |  | 1344.772 | 4 | 4 | 0 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.06 | 28121.049 | 1057346 | 34492.879 | 3.261 | 48.75 | 53.83 | 10.4 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.4 | 110454.477 | 4597438 | 34492.879 | 12.809 | 48.75 | 43.92 | -9.9 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.4 | 34492.879 | 1450823 |  | 34492.879 | 4 | 3.86 | -3.5
20231106_HepCl_ReRun_023 | CC14 | Standard | 3.88 | 22367.092 | 397198 | 3919.615 | 22.826 | 78.125 | 82.72 | 5.9 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.78 | 4940.688 | 170489 | 3919.615 | 5.042 | 78.125 | 72.91 | -6.7 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.78 | 3919.615 | 134613 |  | 3919.615 | 4 | 4.85 | 21.2 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.81 | 510067.344 | 17057968 | 72004.906 | 28.335 | 78.125 | 79.96 | 2.4 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.96 | 72004.906 | 1945760 |  | 72004.906 | 4 | 4.34 | 8.6 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.05 | 4992.377 | 192789 | 1497.735 | 13.333 | 78.125 | 78.98 | 1.1 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.05 | 1497.735 | 54749 |  | 1497.735 | 4 | 4.46 | 11.4 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.06 | 50974.664 | 1864530 | 50112.5 | 4.069 | 78.125 | 67.32 | -13.8 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.4 | 284172 | 11155492 | 50112.5 | 22.683 | 78.125 | 78.19 | 0.1 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.4 | 50112.5 | 1895970 |  | 50112.5 | 4 | 5.61 | 40.1
20231106_HepCl_ReRun_062 | CC14 | Standard | 3.88 | 21361.063 | 375481 | 4066.942 | 21.009 | 78.125 | 76.08 | -2.6 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.78 | 5185.91 | 177930 | 4066.942 | 5.101 | 78.125 | 73.77 | -5.6 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.78 | 4066.942 | 139930 |  | 4066.942 | 4 | 5.03 | 25.7 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.81 | 494636.938 | 16853748 | 75166.273 | 26.322 | 78.125 | 74.25 | -5 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.96 | 75166.273 | 2079953 |  | 75166.273 | 4 | 4.53 | 13.4 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.05 | 4922.576 | 187116 | 1523.123 | 12.928 | 78.125 | 76.55 | -2 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.05 | 1523.123 | 57253 |  | 1523.123 | 4 | 4.53 | 13.3 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.06 | 49863.824 | 1864208 | 41397.715 | 4.818 | 78.125 | 79.83 | 2.2 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.4 | 222048.391 | 9055133 | 41397.715 | 21.455 | 78.125 | 73.93 | -5.4 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.4 | 41397.715 | 1662759 |  | 41397.715 | 4 | 4.63 | 15.8
20231106_HepCl_ReRun_024 | CC15 | Standard | 3.88 | 30114.336 | 533010 | 3888.294 | 30.979 | 125 | 112.53 | -10 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.78 | 7579.503 | 259401 | 3888.294 | 7.797 | 125 | 113.62 | -9.1 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.78 | 3888.294 | 133357 |  | 3888.294 | 4 | 4.81 | 20.2 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.81 | 653154.375 | 22047396 | 67768.188 | 38.552 | 125 | 108.94 | -12.8 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.97 | 67768.188 | 1918844 |  | 67768.188 | 4 | 4.09 | 2.2 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.05 | 7282.961 | 267543 | 1422.099 | 20.485 | 125 | 121.89 | -2.5 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.05 | 1422.099 | 52406 |  | 1422.099 | 4 | 4.23 | 5.8 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.07 | 71695.906 | 2683730 | 38008.148 | 7.545 | 125 | 125.37 | 0.3 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.4 | 376289.313 | 15172176 | 38008.148 | 39.601 | 125 | 136.92 | 9.5 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.4 | 38008.148 | 1745132 |  | 38008.148 | 4 | 4.25 | 6.3
20231106_HepCl_ReRun_063 | CC15 | Standard | 3.88 | 32258.174 | 560369 | 3980.426 | 32.417 | 125 | 117.79 | -5.8 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.78 | 8270.189 | 287977 | 3980.426 | 8.311 | 125 | 121.21 | -3 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.78 | 3980.426 | 148683 |  | 3980.426 | 4 | 4.92 | 23 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.81 | 687010.063 | 23100348 | 64113.281 | 42.862 | 125 | 121.17 | -3.1 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.96 | 64113.281 | 1795978 |  | 64113.281 | 4 | 3.87 | -3.3 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.05 | 6916.329 | 258533 | 1427.197 | 19.384 | 125 | 115.29 | -7.8 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.05 | 1427.197 | 52880 |  | 1427.197 | 4 | 4.25 | 6.1 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.06 | 70819.211 | 2617717 | 38112.547 | 7.433 | 125 | 123.49 | -1.2 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.4 | 327021.594 | 13200246 | 38112.547 | 34.322 | 125 | 118.6 | -5.1 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.4 | 38112.547 | 1543227 |  | 38112.547 | 4 | 4.26 | 6.6
20231106_HepCl_ReRun_025 | CC16 | Analyte | 3.88 | 51578.672 | 891223 | 4835.771 | 42.664 | 200 | 155.25 | -22.4 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 3.78 | 10693.694 | 369588 | 4835.771 | 8.845 | 200 | 129.11 | -35.4 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 3.78 | 4835.771 | 166839 |  | 4835.771 | 4 | 5.98 | 49.5 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 3.81 | 1079783.75 | 36050296 | 89323.164 | 48.354 | 200 | 136.75 | -31.6 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 3.97 | 89323.164 | 2487380 |  | 89323.164 | 4 | 5.39 | 34.7 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 4.05 | 11365.859 | 404654 | 1729.245 | 26.291 | 200 | 156.72 | -21.6 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 4.05 | 1729.245 | 60671 |  | 1729.245 | 4 | 5.14 | 28.6 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 4.06 | 131145.125 | 4748852 | 52133.781 | 10.062 | 200 | 167.41 | -16.3 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 4.4 | 619418.688 | 22182238 | 52133.781 | 47.525 | 200 | 164.43 | -17.8 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 4.4 | 52133.781 | 1791967 |  | 52133.781 | 4 | 5.83 | 45.8
20231106_HepCl_ReRun_064 | CC16 | Analyte | 3.88 | 47232.074 | 763949 | 4885.239 | 38.673 | 200 | 140.66 | -29.7 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 3.78 | 8570.126 | 291253 | 4885.239 | 7.017 | 200 | 102.09 | -49 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 3.78 | 4885.239 | 167128 |  | 4885.239 | 4 | 6.04 | 51 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 3.81 | 1203123.5 | 40178528 | 89929.188 | 53.514 | 200 | 151.39 | -24.3 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 3.96 | 89929.188 | 2487582 |  | 89929.188 | 4 | 5.43 | 35.6 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 4.05 | 11634.63 | 427689 | 1844.922 | 25.225 | 200 | 150.33 | -24.8 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 4.05 | 1844.922 | 67300 |  | 1844.922 | 4 | 5.49 | 37.2 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 4.06 | 139249.141 | 5104794 | 57215.652 | 9.735 | 200 | 161.94 | -19 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 4.4 | 652959.25 | 26559178 | 57215.652 | 45.649 | 200 | 157.92 | -21 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 4.4 | 57215.652 | 2313839 |  | 57215.652 | 4 | 6.4 | 60
20231106_HepCl_ReRun_026 | CC17 | Analyte | 3.88 | 91458.242 | 1662166 | 5718.93 | 63.969 | 325.625 | 233.14 | -28.4 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 3.78 | 18052.461 | 621276 | 5718.93 | 12.626 | 325.625 | 184.98 | -43.2 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 3.78 | 5718.93 | 196763 |  | 5718.93 | 4 | 7.07 | 76.8 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 3.81 | 1637311 | 54934608 | 75562.414 | 86.673 | 325.625 | 245.44 | -24.6 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 3.97 | 75562.414 | 2130806 |  | 75562.414 | 4 | 4.56 | 14 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 4.05 | 18668.201 | 676678 | 1482.366 | 50.374 | 325.625 | 301.21 | -7.5 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 4.05 | 1482.366 | 54721 |  | 1482.366 | 4 | 4.41 | 10.2 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 4.07 | 190048.672 | 6581417 | 43163.68 | 17.612 | 325.625 | 293.48 | -9.9 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 4.4 | 1052163.5 | 39965172 | 43163.68 | 97.505 | 325.625 | 337.92 | 3.8 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 4.4 | 43163.68 | 1662867 |  | 43163.68 | 4 | 4.83 | 20.7
20231106_HepCl_ReRun_065 | CC17 | Analyte | 3.88 | 90726.891 | 1682119 | 6843.28 | 53.031 | 325.625 | 193.15 | -40.7 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 3.78 | 20518.283 | 706717 | 6843.28 | 11.993 | 325.625 | 175.63 | -46.1 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 3.78 | 6843.28 | 238203 |  | 6843.28 | 4 | 8.46 | 111.5 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 3.81 | 1690942.625 | 56753616 | 73827.055 | 91.616 | 325.625 | 259.47 | -20.3 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 3.97 | 73827.055 | 2077619 |  | 73827.055 | 4 | 4.45 | 11.3 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 4.05 | 18651.377 | 669332 | 1484.605 | 50.253 | 325.625 | 300.48 | -7.7 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 4.05 | 1484.605 | 50568 |  | 1484.605 | 4 | 4.42 | 10.4 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 4.07 | 204025.5 | 7220709 | 38809.316 | 21.029 | 325.625 | 350.54 | 7.7 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 4.4 | 930327.5 | 35965604 | 38809.316 | 95.887 | 325.625 | 332.3 | 2.1 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 4.4 | 38809.316 | 1505012 |  | 38809.316 | 4 | 4.34 | 8.5
20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.9 | 7.372 | 255 | [5 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.82 | 33.317 | 899 | [5 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.86 | 475.916 | 8865 | [5 empty] | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.06 | 0.326 | 19 | 0.572 | 2.28 |  | 12.67 |  | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.06 | 0.572 | 55 |  | 0.572 | 4 | 0 | -100 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.01 | 0.655 | 65 | 3.127 | 0.838 |  | 13.36 |  | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.37 | 140.794 | 2229 | 3.127 | 180.101 |  | 624.63 |  | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.38 | 3.127 | 303 |  | 3.127 | 4 | 0 | -100
20231106_HepCl_ReRun_028 | Crash Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.82 | 17.326 | 404 | [5 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.86 | 244.642 | 4634 | 9.577 | 102.179 |  | 289.43 |  | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.98 | 9.577 | 349 |  | 9.577 | 4 | 0 | -100 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.08 | 0.433 | 25 | [5 empty] | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.03 | 3.303 | 150 | 4.24 | 3.116 |  | 51.41 |  | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.36 | 23.582 | 2277 | 4.24 | 22.247 |  | 76.68 |  | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.39 | 4.24 | 142 |  | 4.24 | 4 | 0 | -100
20231106_HepCl_ReRun_046 | Crash Blank | Blank | 3.82 | 2.362 | 85 | [5 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 3.81 | 26.536 | 537 | [5 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.02 | 0.155 | 15 | 0.087 | 7.126 |  | 41.75 |  | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.04 | 0.087 | 5 |  | 0.087 | 4 | 0 | -100 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.04 | 3.32 | 197 | 1.228 | 10.814 |  | 179.97 |  | 20231106_HepCl_ReRun_046 | Crash Blank | Blank |  |  |  | 1.228 |  |  |  |  | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.42 | 1.228 | 138 |  | 1.228 | 4 | 0 | -100
20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.328 |  |  |  |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.328 |  |  |  |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 3.81 | 1.328 | 48 |  | 1.328 | 4 | 0 | -100 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 3.86 | 250.409 | 4456 | 12.974 | 77.203 |  | 218.58 |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 3.97 | 12.974 | 336 |  | 12.974 | 4 | 0 | -100 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.05 | 0.073 | 18 | 0.348 | 0.839 |  | 4.03 |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.1 | 0.348 | 15 |  | 0.348 | 4 | 0 | -100 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.03 | 3.021 | 145 | 1.469 | 8.226 |  | 136.74 |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.469 |  |  |  |  | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.39 | 1.469 | 120 |  | 1.469 | 4 | 0 | -100
20231106_HepCl_ReRun_004 | Matrix Blank | Blank |  |  |  | 8.997 |  |  |  |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.81 | 29.492 | 926 | 8.997 | 13.112 |  | 192.16 |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.81 | 8.997 | 224 |  | 8.997 | 4 | 0.01 | -99.7 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.85 | 607.958 | 11669 | 0.638 | 3811.649 |  | 10811.72 |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.96 | 0.638 | 27 |  | 0.638 | 4 | 0 | -100 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.08 | 1.594 | 79 | 0.86 | 7.414 |  | 43.47 |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.05 | 0.86 | 50 |  | 0.86 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.07 | 1.874 | 69 | 5.073 | 1.478 |  | 24.05 |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.43 | 13.412 | 1130 | 5.073 | 10.575 |  | 36.17 |  | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.41 | 5.073 | 225 |  | 5.073 | 4 | 0 | -100
20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.87 | 35.183 | 628 | 1.072 | 131.28 |  | 479.23 |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.83 | 19.347 | 590 | 1.072 | 72.19 |  | 1065.19 |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.81 | 1.072 | 39 |  | 1.072 | 4 | 0 | -100 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank |  |  |  | 0.598 |  |  |  |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.07 | 0.598 | 35 |  | 0.598 | 4 | 0 | -100 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.07 | 4.099 | 215 | 16.466 | 0.996 |  | 16 |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.38 | 59.634 | 2149 | 16.466 | 14.487 |  | 49.74 |  | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.41 | 16.466 | 203 |  | 16.466 | 4 | 0 | -100
20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.9 | 2.746 | 126 | [5 empty] | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.81 | 16.706 | 473 | [5 empty] | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.86 | 183.34 | 3807 | 8.655 | 84.733 |  | 239.94 |  | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 4 | 8.655 | 211 |  | 8.655 | 4 | 0 | -100 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 0.278 |  |  |  |  | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 4.1 | 0.278 | 12 |  | 0.278 | 4 | 0 | -100 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 9.694 |  |  |  |  | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 9.694 |  |  |  |  | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 4.43 | 9.694 | 259 |  | 9.694 | 4 | 0 | -100
20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 0.37 |  |  |  |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.81 | 13.005 | 297 | 0.37 | 140.595 |  | 2076.03 |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.81 | 0.37 | 26 |  | 0.37 | 4 | 0 | -100 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.86 | 175.411 | 3826 | 0.454 | 1545.471 |  | 4383.47 |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4 | 0.454 | 37 |  | 0.454 | 4 | 0 | -100 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 3.149 |  |  |  |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4.02 | 3.149 | 105 |  | 3.149 | 4 | 0.01 | -99.8 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 6.959 |  |  |  |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4.43 | 0.86 | 403 | 6.959 | 0.494 |  | 1.17 |  | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4.37 | 6.959 | 192 |  | 6.959 | 4 | 0 | -100
20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.89 | 15743.503 | 247765 | 4776.272 | 13.185 |  | 47.47 |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.78 | 205.262 | 6517 | 4776.272 | 0.172 |  | 0.94 |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.78 | 4776.272 | 164687 |  | 4776.272 | 4 | 5.91 | 47.6 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.81 | 514.675 | 8646 | 127111.25 | 0.016 |  |  |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.97 | 127111.25 | 3554961 |  | 127111.25 | 4 | 7.67 | 91.7 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.05 | 455.547 | 17598 | 2610.134 | 0.698 |  | 3.18 |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.05 | 2610.134 | 101241 |  | 2610.134 | 4 | 7.76 | 94.1 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte |  |  |  | 98002.758 |  |  |  |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.4 | 288979.969 | 11903002 | 98002.758 | 11.795 |  | 40.4 |  | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.4 | 98002.758 | 4043580 |  | 98002.758 | 4 | 10.96 | 174.1
20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.88 | 13915.396 | 245491 | 3195.141 | 17.421 |  | 62.96 |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.78 | 292.673 | 9773 | 3195.141 | 0.366 |  | 3.81 |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.77 | 3195.141 | 107465 |  | 3195.141 | 4 | 3.95 | -1.2 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.81 | 993.376 | 26476 | 107706.938 | 0.037 |  |  |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.96 | 107706.938 | 3042009 |  | 107706.938 | 4 | 6.5 | 62.4 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.05 | 404.303 | 15376 | 2373.515 | 0.681 |  | 3.08 |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.05 | 2373.515 | 89417 |  | 2373.515 | 4 | 7.06 | 76.5 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte |  |  |  | 88635.797 |  |  |  |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.4 | 252387.266 | 10245247 | 88635.797 | 11.39 |  | 38.99 |  | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.4 | 88635.797 | 3597215 |  | 88635.797 | 4 | 9.92 | 147.9
20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.88 | 14715.912 | 240678 | 4462.737 | 13.19 |  | 47.49 |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.78 | 210.794 | 7030 | 4462.737 | 0.189 |  | 1.19 |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.78 | 4462.737 | 153013 |  | 4462.737 | 4 | 5.52 | 37.9 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.81 | 642.249 | 12421 | 113437.414 | 0.023 |  |  |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.96 | 113437.414 | 3224395 |  | 113437.414 | 4 | 6.84 | 71.1 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.05 | 382.016 | 15123 | 2397.312 | 0.637 |  | 2.82 |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.05 | 2397.312 | 91363 |  | 2397.312 | 4 | 7.13 | 78.3 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte |  |  |  | 89190.406 |  |  |  |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.4 | 232170.688 | 9274514 | 89190.406 | 10.412 |  | 35.6 |  | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.4 | 89190.406 | 3620419 |  | 89190.406 | 4 | 9.98 | 149.4
20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.87 | 9074.265 | 137502 | 3209.497 | 11.309 |  | 40.62 |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.78 | 108.098 | 3082 | 3209.497 | 0.135 |  | 0.39 |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.78 | 3209.497 | 108432 |  | 3209.497 | 4 | 3.97 | -0.8 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.81 | 412.521 | 6320 | 94312.25 | 0.017 |  |  |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.96 | 94312.25 | 2578495 |  | 94312.25 | 4 | 5.69 | 42.2 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.05 | 231.241 | 9183 | 2126.96 | 0.435 |  | 1.6 |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.05 | 2126.96 | 81491 |  | 2126.96 | 4 | 6.33 | 58.2 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.06 | 11.272 | 433 | 76076.914 | 0.001 |  |  |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.4 | 205967.813 | 8300433 | 76076.914 | 10.829 |  | 37.05 |  | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.4 | 76076.914 | 2958757 |  | 76076.914 | 4 | 8.51 | 112.8
20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.88 | 9561.769 | 153073 | 3877.044 | 9.865 |  | 35.34 |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.78 | 139.492 | 4379 | 3877.044 | 0.144 |  | 0.52 |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.78 | 3877.044 | 131675 |  | 3877.044 | 4 | 4.79 | 19.8 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.8 | 408.999 | 5523 | 101210.508 | 0.016 |  |  |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.96 | 101210.508 | 2857656 |  | 101210.508 | 4 | 6.11 | 52.6 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.05 | 357.515 | 13825 | 2142.216 | 0.668 |  | 3 |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.05 | 2142.216 | 78827 |  | 2142.216 | 4 | 6.37 | 59.3 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.07 | 7.606 | 361 | 76258 | 0 |  |  |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.4 | 206321.797 | 8091617 | 76258 | 10.822 |  | 37.02 |  | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.4 | 76258 | 2982523 |  | 76258 | 4 | 8.53 | 113.3
20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.88 | 9736.915 | 150985 | 3869.138 | 10.066 |  | 36.07 |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.78 | 106.808 | 3014 | 3869.138 | 0.11 |  | 0.03 |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.78 | 3869.138 | 133521 |  | 3869.138 | 4 | 4.78 | 19.6 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.81 | 356.726 | 5025 | 104400.188 | 0.014 |  |  |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.96 | 104400.188 | 2856369 |  | 104400.188 | 4 | 6.3 | 57.5 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.05 | 167.013 | 6154 | 2372.858 | 0.282 |  | 0.68 |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.05 | 2372.858 | 87556 |  | 2372.858 | 4 | 7.06 | 76.5 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.06 | 8.773 | 389 | 80176.891 | 0 |  |  |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.4 | 225792.75 | 9374209 | 80176.891 | 11.265 |  | 38.56 |  | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.4 | 80176.891 | 3248475 |  | 80176.891 | 4 | 8.97 | 124.2
20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.88 | 12368.906 | 206371 | 4447.937 | 11.123 |  | 39.94 |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.78 | 144.765 | 4211 | 4447.937 | 0.13 |  | 0.32 |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.77 | 4447.937 | 152114 |  | 4447.937 | 4 | 5.5 | 37.5 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.81 | 408.254 | 5604 | 110987.766 | 0.015 |  |  |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.96 | 110987.766 | 3056744 |  | 110987.766 | 4 | 6.7 | 67.4 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.05 | 153.025 | 6504 | 2359.211 | 0.259 |  | 0.55 |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.05 | 2359.211 | 88771 |  | 2359.211 | 4 | 7.02 | 75.4 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte |  |  |  | 89103.305 |  |  |  |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.4 | 267149.969 | 10917079 | 89103.305 | 11.993 |  | 41.09 |  | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.4 | 89103.305 | 3621246 |  | 89103.305 | 4 | 9.97 | 149.2
20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.88 | 12134.313 | 203497 | 4036.639 | 12.024 |  | 43.23 |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.78 | 99.272 | 2968 | 4036.639 | 0.098 |  |  |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.78 | 4036.639 | 139623 |  | 4036.639 | 4 | 4.99 | 24.8 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.81 | 337.916 | 5499 | 99300.031 | 0.014 |  |  |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.96 | 99300.031 | 2693969 |  | 99300.031 | 4 | 5.99 | 49.8 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.05 | 408.419 | 15643 | 2233.017 | 0.732 |  | 3.38 |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.05 | 2233.017 | 83972 |  | 2233.017 | 4 | 6.64 | 66.1 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.06 | 18.349 | 722 | 83877.867 | 0.001 |  |  |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.4 | 231363.469 | 9350834 | 83877.867 | 11.033 |  | 37.76 |  | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.4 | 83877.867 | 3228451 |  | 83877.867 | 4 | 9.38 | 134.6
20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.88 | 11903.52 | 204453 | 3155.708 | 15.088 |  | 54.43 |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.77 | 136.658 | 4376 | 3155.708 | 0.173 |  | 0.96 |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.78 | 3155.708 | 105566 |  | 3155.708 | 4 | 3.9 | -2.5 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.81 | 462.735 | 7215 | 105545.609 | 0.018 |  |  |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.96 | 105545.609 | 2893700 |  | 105545.609 | 4 | 6.37 | 59.2 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 164.305 | 6189 | 2234.758 | 0.294 |  | 0.76 |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 2234.758 | 81479 |  | 2234.758 | 4 | 6.65 | 66.2 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 7.83 | 256 | 86799.516 | 0 |  |  |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.39 | 257262.188 | 10705604 | 86799.516 | 11.855 |  | 40.61 |  | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.4 | 86799.516 | 3602688 |  | 86799.516 | 4 | 9.71 | 142.7
20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.87 | 6488.23 | 94317 | 3451.207 | 7.52 |  | 26.76 |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.77 | 94.632 | 2824 | 3451.207 | 0.11 |  | 0.02 |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.78 | 3451.207 | 117553 |  | 3451.207 | 4 | 4.27 | 6.7 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.81 | 404.812 | 5864 | 84769.352 | 0.019 |  |  |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.96 | 84769.352 | 2328900 |  | 84769.352 | 4 | 5.11 | 27.8 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.05 | 315.459 | 11990 | 1960.445 | 0.644 |  | 2.85 |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.05 | 1960.445 | 72293 |  | 1960.445 | 4 | 5.83 | 45.8 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.06 | 0.389 | 70 | 65362.707 | 0 |  |  |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.4 | 189770.188 | 7714787 | 65362.707 | 11.613 |  | 39.77 |  | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.4 | 65362.707 | 2663057 |  | 65362.707 | 4 | 7.31 | 82.8
20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.88 | 4897.804 | 78547 | 3384.839 | 5.788 |  | 20.43 |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.78 | 83.841 | 2740 | 3384.839 | 0.099 |  |  |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.77 | 3384.839 | 115881 |  | 3384.839 | 4 | 4.19 | 4.6 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.85 | 333.628 | 4606 | 88387.469 | 0.015 |  |  |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.96 | 88387.469 | 2483812 |  | 88387.469 | 4 | 5.33 | 33.3 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.05 | 136.141 | 5084 | 1928.79 | 0.282 |  | 0.69 |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.05 | 1928.79 | 72627 |  | 1928.79 | 4 | 5.74 | 43.4 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.06 | 4.735 | 316 | 66185.563 | 0 |  |  |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.4 | 188079.766 | 7687524 | 66185.563 | 11.367 |  | 38.91 |  | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.4 | 66185.563 | 2694615 |  | 66185.563 | 4 | 7.4 | 85.1
20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.88 | 7423.349 | 109685 | 3411.54 | 8.704 |  | 31.09 |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.78 | 78.491 | 2813 | 3411.54 | 0.092 |  |  |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.78 | 3411.54 | 117593 |  | 3411.54 | 4 | 4.22 | 5.5 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.86 | 371.009 | 4792 | 75575.227 | 0.02 |  |  |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.96 | 75575.227 | 2013143 |  | 75575.227 | 4 | 4.56 | 14 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.05 | 111.834 | 4367 | 1861.049 | 0.24 |  | 0.43 |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.05 | 1861.049 | 71569 |  | 1861.049 | 4 | 5.54 | 38.4 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.06 | 18.365 | 455 | 64907.961 | 0.001 |  |  |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.4 | 188022.875 | 7440712 | 64907.961 | 11.587 |  | 39.68 |  | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.4 | 64907.961 | 2536898 |  | 64907.961 | 4 | 7.26 | 81.5
20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.88 | 11315.929 | 169106 | 4054.992 | 11.162 |  | 40.08 |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.78 | 110.127 | 3275 | 4054.992 | 0.109 |  | 0 |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.78 | 4054.992 | 139514 |  | 4054.992 | 4 | 5.01 | 25.3 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.81 | 453.139 | 6478 | 104039.031 | 0.017 |  |  |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.96 | 104039.031 | 2808497 |  | 104039.031 | 4 | 6.28 | 56.9 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.05 | 142.847 | 6171 | 2519.11 | 0.227 |  | 0.35 |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.05 | 2519.11 | 93264 |  | 2519.11 | 4 | 7.49 | 87.3 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.07 | 3.867 | 251 | 93797.602 | 0 |  |  |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.4 | 279341.906 | 11378870 | 93797.602 | 11.913 |  | 40.81 |  | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.4 | 93797.602 | 3755169 |  | 93797.602 | 4 | 10.49 | 162.3
20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.87 | 12633.037 | 196143 | 4225.446 | 11.959 |  | 42.99 |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.78 | 98.905 | 2884 | 4225.446 | 0.094 |  |  |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.78 | 4225.446 | 146406 |  | 4225.446 | 4 | 5.22 | 30.6 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.86 | 381.047 | 4607 | 119178.984 | 0.013 |  |  |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.96 | 119178.984 | 3371710 |  | 119178.984 | 4 | 7.19 | 79.7 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.05 | 149.053 | 5522 | 2439.5 | 0.244 |  | 0.46 |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.05 | 2439.5 | 90275 |  | 2439.5 | 4 | 7.26 | 81.4 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.07 | 16.547 | 987 | 95630.289 | 0.001 |  |  |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.4 | 282657.531 | 11606108 | 95630.289 | 11.823 |  | 40.5 |  | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.4 | 95630.289 | 3956639 |  | 95630.289 | 4 | 10.7 | 167.4
20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.88 | 11869.145 | 199049 | 4254.767 | 11.158 |  | 40.06 |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.77 | 130.767 | 3718 | 4254.767 | 0.123 |  | 0.21 |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.78 | 4254.767 | 145612 |  | 4254.767 | 4 | 5.26 | 31.5 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.81 | 396.53 | 5932 | 113962.063 | 0.014 |  |  |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.96 | 113962.063 | 3231676 |  | 113962.063 | 4 | 6.87 | 71.9 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.05 | 261.987 | 10292 | 2235.031 | 0.469 |  | 1.81 |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.05 | 2235.031 | 83647 |  | 2235.031 | 4 | 6.65 | 66.2 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.06 | 8.845 | 502 | 88288.648 | 0 |  |  |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.4 | 274644.281 | 11201561 | 88288.648 | 12.443 |  | 42.65 |  | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.4 | 88288.648 | 3610306 |  | 88288.648 | 4 | 9.88 | 146.9
20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.88 | 11134.584 | 184254 | 4188.782 | 10.633 |  | 38.14 |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.78 | 137.994 | 4125 | 4188.782 | 0.132 |  | 0.34 |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.78 | 4188.782 | 143659 |  | 4188.782 | 4 | 5.18 | 29.5 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.81 | 322.722 | 5508 | 118969.555 | 0.011 |  |  |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.96 | 118969.555 | 3318540 |  | 118969.555 | 4 | 7.18 | 79.4 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.05 | 218.786 | 8347 | 2343.231 | 0.373 |  | 1.23 |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.05 | 2343.231 | 87432 |  | 2343.231 | 4 | 6.97 | 74.3 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.07 | 7.013 | 311 | 88574.773 | 0 |  |  |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.4 | 279947.906 | 11400369 | 88574.773 | 12.642 |  | 43.34 |  | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.4 | 88574.773 | 3453310 |  | 88574.773 | 4 | 9.91 | 147.7
20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.87 | 10594.069 | 161173 | 4496.683 | 9.424 |  | 33.72 |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.78 | 140.797 | 4509 | 4496.683 | 0.125 |  | 0.25 |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.78 | 4496.683 | 155205 |  | 4496.683 | 4 | 5.56 | 39 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.85 | 377.622 | 4486 | 122033.266 | 0.012 |  |  |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.96 | 122033.266 | 3442648 |  | 122033.266 | 4 | 7.36 | 84 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.05 | 139.54 | 5755 | 2400.613 | 0.233 |  | 0.39 |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.05 | 2400.613 | 90344 |  | 2400.613 | 4 | 7.14 | 78.5 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.07 | 9.717 | 489 | 95312.484 | 0 |  |  |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.4 | 300004.281 | 12091701 | 95312.484 | 12.59 |  | 43.16 |  | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.4 | 95312.484 | 3746577 |  | 95312.484 | 4 | 10.66 | 166.6
20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.88 | 10354.753 | 164488 | 4153.806 | 9.971 |  | 35.72 |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.78 | 121.203 | 3598 | 4153.806 | 0.117 |  | 0.12 |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.78 | 4153.806 | 139829 |  | 4153.806 | 4 | 5.14 | 28.4 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.85 | 390.746 | 5225 | 110901.977 | 0.014 |  |  |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.96 | 110901.977 | 3062770 |  | 110901.977 | 4 | 6.69 | 67.3 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.05 | 143.163 | 5696 | 2513.042 | 0.228 |  | 0.36 |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.05 | 2513.042 | 92824 |  | 2513.042 | 4 | 7.48 | 86.9 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte |  |  |  | 94841.609 |  |  |  |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.4 | 299780.094 | 11869348 | 94841.609 | 12.643 |  | 43.34 |  | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.4 | 94841.609 | 3775499 |  | 94841.609 | 4 | 10.61 | 165.2
20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.87 | 13138.3 | 187084 | 4056.419 | 12.956 |  | 46.64 |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.78 | 189.653 | 5817 | 4056.419 | 0.187 |  | 1.16 |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.78 | 4056.419 | 140110 |  | 4056.419 | 4 | 5.02 | 25.4 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.81 | 457.152 | 6935 | 96807.188 | 0.019 |  |  |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.97 | 96807.188 | 2668441 |  | 96807.188 | 4 | 5.84 | 46 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.05 | 207.095 | 7876 | 2059.017 | 0.402 |  | 1.41 |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.05 | 2059.017 | 77275 |  | 2059.017 | 4 | 6.12 | 53.1 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte |  |  |  | 76163.773 |  |  |  |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.4 | 198084.219 | 8400724 | 76163.773 | 10.403 |  | 35.57 |  | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.4 | 76163.773 | 3220525 |  | 76163.773 | 4 | 8.52 | 113
20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.87 | 14517.875 | 206204 | 4559.46 | 12.736 |  | 45.83 |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.78 | 241.631 | 8170 | 4559.46 | 0.212 |  | 1.53 |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.78 | 4559.46 | 157257 |  | 4559.46 | 4 | 5.64 | 40.9 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.81 | 475.469 | 8669 | 110827.492 | 0.017 |  |  |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.96 | 110827.492 | 3028426 |  | 110827.492 | 4 | 6.69 | 67.1 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.05 | 353.006 | 13268 | 2482.213 | 0.569 |  | 2.41 |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.05 | 2482.213 | 93606 |  | 2482.213 | 4 | 7.38 | 84.6 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.07 | 1.967 | 217 | 92251.773 | 0 |  |  |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.4 | 241833.688 | 9651093 | 92251.773 | 10.486 |  | 35.85 |  | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.4 | 92251.773 | 3670430 |  | 92251.773 | 4 | 10.32 | 158
20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.87 | 15841.149 | 227470 | 4493.075 | 14.103 |  | 50.83 |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.78 | 290.765 | 9335 | 4493.075 | 0.259 |  | 2.22 |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.78 | 4493.075 | 154908 |  | 4493.075 | 4 | 5.56 | 38.9 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.81 | 554.069 | 8920 | 106620.406 | 0.021 |  |  |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.96 | 106620.406 | 3040845 |  | 106620.406 | 4 | 6.43 | 60.8 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 232.939 | 8411 | 2293.903 | 0.406 |  | 1.43 |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 2293.903 | 85682 |  | 2293.903 | 4 | 6.82 | 70.6 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 4.014 | 281 | 83986.945 | 0 |  |  |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.4 | 218327.75 | 8751073 | 83986.945 | 10.398 |  | 35.55 |  | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.4 | 83986.945 | 3315792 |  | 83986.945 | 4 | 9.4 | 134.9
20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.87 | 12003.916 | 162939 | 4114.258 | 11.671 |  | 41.94 |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.78 | 190.399 | 6349 | 4114.258 | 0.185 |  | 1.13 |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.78 | 4114.258 | 141391 |  | 4114.258 | 4 | 5.09 | 27.2 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.81 | 469.105 | 6650 | 86723.508 | 0.022 |  |  |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.96 | 86723.508 | 2258546 |  | 86723.508 | 4 | 5.23 | 30.8 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.05 | 220.973 | 8680 | 2212.084 | 0.4 |  | 1.39 |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.05 | 2212.084 | 83409 |  | 2212.084 | 4 | 6.58 | 64.5 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.06 | 10.213 | 486 | 86446.32 | 0 |  |  |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.4 | 238265.75 | 9592995 | 86446.32 | 11.025 |  | 37.73 |  | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.4 | 86446.32 | 3432628 |  | 86446.32 | 4 | 9.67 | 141.8
20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.86 | 13315.447 | 183072 | 4004.459 | 13.301 |  | 47.9 |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.78 | 243.066 | 7101 | 4004.459 | 0.243 |  | 1.99 |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.78 | 4004.459 | 137439 |  | 4004.459 | 4 | 4.95 | 23.8 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte |  |  |  | 97677.023 |  |  |  |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.96 | 97677.023 | 2650361 |  | 97677.023 | 4 | 5.89 | 47.3 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.05 | 205.306 | 8231 | 2292.081 | 0.358 |  | 1.14 |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.05 | 2292.081 | 84135 |  | 2292.081 | 4 | 6.82 | 70.4 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte |  |  |  | 89926.844 |  |  |  |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.4 | 262954.25 | 10368118 | 89926.844 | 11.696 |  | 40.06 |  | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.4 | 89926.844 | 3532382 |  | 89926.844 | 4 | 10.06 | 151.5
20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.87 | 13054.155 | 177255 | 3985.281 | 13.102 |  | 47.17 |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.78 | 176.193 | 5558 | 3985.281 | 0.177 |  | 1.01 |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.77 | 3985.281 | 135627 |  | 3985.281 | 4 | 4.93 | 23.2 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.86 | 673.771 | 8906 | 100329.82 | 0.027 |  |  |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.96 | 100329.82 | 2834738 |  | 100329.82 | 4 | 6.05 | 51.3 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.05 | 193.423 | 7058 | 2233.479 | 0.346 |  | 1.07 |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.05 | 2233.479 | 87834 |  | 2233.479 | 4 | 6.64 | 66.1 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.09 | 16.609 | 464 | 91128.219 | 0.001 |  |  |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.4 | 266459.938 | 11257610 | 91128.219 | 11.696 |  | 40.06 |  | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.4 | 91128.219 | 3617198 |  | 91128.219 | 4 | 10.19 | 154.9
20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.87 | 13458.945 | 179811 | 3946.084 | 13.643 |  | 49.15 |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.78 | 372.564 | 13071 | 3946.084 | 0.378 |  | 3.98 |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.78 | 3946.084 | 135784 |  | 3946.084 | 4 | 4.88 | 22 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.81 | 1741.671 | 49424 | 98270.578 | 0.071 |  |  |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.96 | 98270.578 | 2691723 |  | 98270.578 | 4 | 5.93 | 48.2 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.05 | 1087.664 | 41574 | 2277.866 | 1.91 |  | 10.45 |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.05 | 2277.866 | 87482 |  | 2277.866 | 4 | 6.78 | 69.4 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.06 | 64.871 | 2246 | 78984.391 | 0.003 |  |  |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.4 | 228470.469 | 9286333 | 78984.391 | 11.57 |  | 39.62 |  | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.4 | 78984.391 | 3150051 |  | 78984.391 | 4 | 8.84 | 120.9
20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.88 | 13394.592 | 210981 | 4015.599 | 13.343 |  | 48.05 |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.78 | 110.375 | 3443 | 4015.599 | 0.11 |  | 0.02 |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.78 | 4015.599 | 139237 |  | 4015.599 | 4 | 4.96 | 24.1 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.85 | 359.241 | 4577 | 98628.742 | 0.015 |  |  |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.96 | 98628.742 | 2774183 |  | 98628.742 | 4 | 5.95 | 48.7 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.05 | 109.222 | 3935 | 1961.049 | 0.223 |  | 0.33 |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.05 | 1961.049 | 74969 |  | 1961.049 | 4 | 5.83 | 45.8 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte |  |  |  | 54687.828 |  |  |  |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.4 | 145300.688 | 5942828 | 54687.828 | 10.628 |  | 36.35 |  | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.4 | 54687.828 | 2268070 |  | 54687.828 | 4 | 6.12 | 52.9
20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.88 | 12774.548 | 184155 | 3722.044 | 13.729 |  | 49.46 |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.78 | 78.054 | 2592 | 3722.044 | 0.084 |  |  |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.78 | 3722.044 | 127148 |  | 3722.044 | 4 | 4.6 | 15.1 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.86 | 327.177 | 5240 | 106868.266 | 0.012 |  |  |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.96 | 106868.266 | 3033461 |  | 106868.266 | 4 | 6.45 | 61.2 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.05 | 134.736 | 5715 | 2366.501 | 0.228 |  | 0.36 |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.05 | 2366.501 | 90032 |  | 2366.501 | 4 | 7.04 | 76 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.07 | 6.07 | 347 | 63527.227 | 0 |  |  |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.4 | 171869.766 | 6608780 | 63527.227 | 10.822 |  | 37.02 |  | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.4 | 63527.227 | 2347999 |  | 63527.227 | 4 | 7.11 | 77.7
20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.87 | 13366.214 | 177634 | 3543.577 | 15.088 |  | 54.43 |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.78 | 86.154 | 2791 | 3543.577 | 0.097 |  |  |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.78 | 3543.577 | 120707 |  | 3543.577 | 4 | 4.38 | 9.5 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.81 | 381.625 | 5824 | 102934.406 | 0.015 |  |  |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.96 | 102934.406 | 2925716 |  | 102934.406 | 4 | 6.21 | 55.2 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.05 | 152.881 | 6501 | 2177.948 | 0.281 |  | 0.68 |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.05 | 2177.948 | 82550 |  | 2177.948 | 4 | 6.48 | 62 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte |  |  |  | 73346.039 |  |  |  |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.4 | 199021.969 | 7933975 | 73346.039 | 10.854 |  | 37.13 |  | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.4 | 73346.039 | 2836304 |  | 73346.039 | 4 | 8.2 | 105.1
20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.87 | 12986.323 | 189758 | 3555.943 | 14.608 |  | 52.68 |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.78 | 115.902 | 3813 | 3555.943 | 0.13 |  | 0.32 |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.78 | 3555.943 | 121173 |  | 3555.943 | 4 | 4.4 | 9.9 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.85 | 415.396 | 5471 | 95880.813 | 0.017 |  |  |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.96 | 95880.813 | 2743191 |  | 95880.813 | 4 | 5.78 | 44.6 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.05 | 356.513 | 14002 | 1966.663 | 0.725 |  | 3.34 |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.05 | 1966.663 | 74380 |  | 1966.663 | 4 | 5.85 | 46.2 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.06 | 14.537 | 793 | 68280.641 | 0.001 |  |  |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.39 | 195954.266 | 7834061 | 68280.641 | 11.479 |  | 39.3 |  | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.4 | 68280.641 | 2797197 |  | 68280.641 | 4 | 7.64 | 91
20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.86 | 13166.602 | 186700 | 3988.081 | 13.206 |  | 47.55 |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.78 | 137.6 | 4290 | 3988.081 | 0.138 |  | 0.44 |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.78 | 3988.081 | 136185 |  | 3988.081 | 4 | 4.93 | 23.3 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.81 | 350.126 | 4576 | 111986.016 | 0.013 |  |  |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.96 | 111986.016 | 3190599 |  | 111986.016 | 4 | 6.76 | 68.9 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.05 | 221.256 | 8132 | 2330.105 | 0.38 |  | 1.27 |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.05 | 2330.105 | 87302 |  | 2330.105 | 4 | 6.93 | 73.3 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.06 | 11.696 | 573 | 73499.289 | 0.001 |  |  |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.4 | 209089.75 | 8472446 | 73499.289 | 11.379 |  | 38.96 |  | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.4 | 73499.289 | 2960621 |  | 73499.289 | 4 | 8.22 | 105.5
20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.87 | 13906.821 | 198513 | 3719.442 | 14.956 |  | 53.95 |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.78 | 138.96 | 4052 | 3719.442 | 0.149 |  | 0.61 |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.78 | 3719.442 | 129467 |  | 3719.442 | 4 | 4.6 | 15 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.81 | 471.128 | 8876 | 104541.266 | 0.018 |  |  |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.96 | 104541.266 | 2937157 |  | 104541.266 | 4 | 6.31 | 57.7 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.05 | 200.381 | 7915 | 2257.612 | 0.355 |  | 1.12 |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.05 | 2257.612 | 86402 |  | 2257.612 | 4 | 6.72 | 67.9 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.07 | 6.733 | 385 | 81925.484 | 0 |  |  |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.4 | 232343.172 | 9388697 | 81925.484 | 11.344 |  | 38.83 |  | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.4 | 81925.484 | 3419540 |  | 81925.484 | 4 | 9.16 | 129.1
20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.88 | 12754.647 | 175426 | 3898.412 | 13.087 |  | 47.12 |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.78 | 126.729 | 4009 | 3898.412 | 0.13 |  | 0.32 |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.78 | 3898.412 | 134486 |  | 3898.412 | 4 | 4.82 | 20.5 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.81 | 416.801 | 6653 | 112556.063 | 0.015 |  |  |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.96 | 112556.063 | 3156701 |  | 112556.063 | 4 | 6.79 | 69.8 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 271.229 | 10772 | 2389.278 | 0.454 |  | 1.72 |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 2389.278 | 89739 |  | 2389.278 | 4 | 7.11 | 77.7 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 12.484 | 452 | 77283.281 | 0.001 |  |  |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.4 | 238910.734 | 9909917 | 77283.281 | 12.365 |  | 42.38 |  | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.4 | 77283.281 | 3151650 |  | 77283.281 | 4 | 8.65 | 116.1
20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.88 | 13496.068 | 184112 | 3840.235 | 14.058 |  | 50.66 |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.78 | 122.661 | 3841 | 3840.235 | 0.128 |  | 0.29 |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.78 | 3840.235 | 132810 |  | 3840.235 | 4 | 4.75 | 18.7 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.81 | 276.12 | 4768 | 104859.008 | 0.011 |  |  |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.96 | 104859.008 | 2910590 |  | 104859.008 | 4 | 6.33 | 58.1 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.05 | 147.119 | 5598 | 2282.505 | 0.258 |  | 0.54 |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.05 | 2282.505 | 85777 |  | 2282.505 | 4 | 6.79 | 69.7 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.06 | 2.375 | 131 | 77620.695 | 0 |  |  |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.4 | 228967.797 | 9634266 | 77620.695 | 11.799 |  | 40.41 |  | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.4 | 77620.695 | 3186330 |  | 77620.695 | 4 | 8.68 | 117.1
20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.87 | 11126.391 | 155600 | 3506.183 | 12.693 |  | 45.68 |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.78 | 56.173 | 1957 | 3506.183 | 0.064 |  |  |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.78 | 3506.183 | 119804 |  | 3506.183 | 4 | 4.34 | 8.4 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte |  |  |  | 84980.797 |  |  |  |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.96 | 84980.797 | 2345692 |  | 84980.797 | 4 | 5.13 | 28.2 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.05 | 118.698 | 4868 | 1788.846 | 0.265 |  | 0.58 |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.05 | 1788.846 | 68641 |  | 1788.846 | 4 | 5.32 | 33 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte |  |  |  | 55749.684 |  |  |  |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.4 | 163660.906 | 6955397 | 55749.684 | 11.743 |  | 40.22 |  | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.4 | 55749.684 | 2264068 |  | 55749.684 | 4 | 6.24 | 55.9
20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.87 | 9472.6 | 127395 | 2791.078 | 13.576 |  | 48.9 |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.77 | 80.953 | 2729 | 2791.078 | 0.116 |  | 0.11 |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.77 | 2791.078 | 94932 |  | 2791.078 | 4 | 3.45 | -13.7 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.85 | 292.296 | 3730 | 64520.02 | 0.018 |  |  |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.96 | 64520.02 | 1765655 |  | 64520.02 | 4 | 3.89 | -2.7 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.05 | 89.453 | 3749 | 1405.372 | 0.255 |  | 0.52 |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.05 | 1405.372 | 53963 |  | 1405.372 | 4 | 4.18 | 4.5 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte |  |  |  | 30788.695 |  |  |  |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.4 | 87596.305 | 3600304 | 30788.695 | 11.38 |  | 38.96 |  | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.4 | 30788.695 | 1220278 |  | 30788.695 | 4 | 3.44 | -13.9
20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.87 | 9704.444 | 128645 | 3117.41 | 12.452 |  | 44.79 |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.78 | 62.378 | 2190 | 3117.41 | 0.08 |  |  |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.78 | 3117.41 | 106241 |  | 3117.41 | 4 | 3.85 | -3.6 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.85 | 367.758 | 5330 | 77287.742 | 0.019 |  |  |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.96 | 77287.742 | 2053044 |  | 77287.742 | 4 | 4.66 | 16.6 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.05 | 130.351 | 5363 | 1826.592 | 0.285 |  | 0.7 |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.05 | 1826.592 | 70966 |  | 1826.592 | 4 | 5.43 | 35.8 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.07 | 7.596 | 468 | 54407.445 | 0.001 |  |  |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.4 | 144037.141 | 5669520 | 54407.445 | 10.59 |  | 36.21 |  | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.4 | 54407.445 | 2019079 |  | 54407.445 | 4 | 6.09 | 52.2
20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.88 | 14466.164 | 202071 | 3796.952 | 15.24 |  | 54.99 |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.78 | 115.056 | 3650 | 3796.952 | 0.121 |  | 0.19 |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.78 | 3796.952 | 130307 |  | 3796.952 | 4 | 4.69 | 17.4 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.81 | 414.222 | 6606 | 105383.109 | 0.016 |  |  |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.96 | 105383.109 | 2935355 |  | 105383.109 | 4 | 6.36 | 58.9 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.05 | 167.121 | 6732 | 2296.952 | 0.291 |  | 0.74 |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.05 | 2296.952 | 84946 |  | 2296.952 | 4 | 6.83 | 70.8 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.06 | 4.094 | 252 | 86903.203 | 0 |  |  |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.4 | 257647.172 | 10769629 | 86903.203 | 11.859 |  | 40.62 |  | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.4 | 86903.203 | 3634190 |  | 86903.203 | 4 | 9.72 | 143
20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.87 | 13676.497 | 198516 | 3988.584 | 13.716 |  | 49.41 |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.77 | 98.341 | 3529 | 3988.584 | 0.099 |  |  |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.77 | 3988.584 | 136537 |  | 3988.584 | 4 | 4.93 | 23.3 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.85 | 410.053 | 5828 | 99879.406 | 0.016 |  |  |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.96 | 99879.406 | 2791654 |  | 99879.406 | 4 | 6.03 | 50.6 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.05 | 147.658 | 5726 | 2078.18 | 0.284 |  | 0.7 |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.05 | 2078.18 | 78341 |  | 2078.18 | 4 | 6.18 | 54.5 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte |  |  |  | 64245.207 |  |  |  |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.4 | 179824.234 | 6961351 | 64245.207 | 11.196 |  | 38.32 |  | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.4 | 64245.207 | 2644733 |  | 64245.207 | 4 | 7.19 | 79.7
20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.87 | 13672.899 | 183759 | 4089.779 | 13.373 |  | 48.16 |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.78 | 149.401 | 5058 | 4089.779 | 0.146 |  | 0.56 |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.77 | 4089.779 | 139738 |  | 4089.779 | 4 | 5.06 | 26.4 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.81 | 350.653 | 7143 | 94358.172 | 0.015 |  |  |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.96 | 94358.172 | 2525923 |  | 94358.172 | 4 | 5.69 | 42.3 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.05 | 219.107 | 8462 | 2319.462 | 0.378 |  | 1.26 |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.05 | 2319.462 | 87625 |  | 2319.462 | 4 | 6.9 | 72.5 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.06 | 20.189 | 1028 | 84458.141 | 0.001 |  |  |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.4 | 246133.094 | 10120721 | 84458.141 | 11.657 |  | 39.92 |  | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.4 | 84458.141 | 3363683 |  | 84458.141 | 4 | 9.45 | 136.2
20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.87 | 13819.075 | 185596 | 4257.981 | 12.982 |  | 46.73 |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.78 | 185.739 | 6213 | 4257.981 | 0.174 |  | 0.98 |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.78 | 4257.981 | 146948 |  | 4257.981 | 4 | 5.26 | 31.6 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.8 | 772.868 | 20550 | 116692.453 | 0.026 |  |  |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.96 | 116692.453 | 3338579 |  | 116692.453 | 4 | 7.04 | 76 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.05 | 691.918 | 26391 | 2365.194 | 1.17 |  | 6.01 |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.05 | 2365.194 | 86293 |  | 2365.194 | 4 | 7.04 | 75.9 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.06 | 68.218 | 3260 | 87830.766 | 0.003 |  |  |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.4 | 278046.969 | 11063472 | 87830.766 | 12.663 |  | 43.41 |  | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.4 | 87830.766 | 3422934 |  | 87830.766 | 4 | 9.83 | 145.6
20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.86 | 14415.857 | 197477 | 4107.876 | 14.037 |  | 50.59 |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.78 | 297.333 | 9442 | 4107.876 | 0.29 |  | 2.68 |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.77 | 4107.876 | 139990 |  | 4107.876 | 4 | 5.08 | 27 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.81 | 947.536 | 23810 | 107368.117 | 0.035 |  |  |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.96 | 107368.117 | 3028922 |  | 107368.117 | 4 | 6.48 | 61.9 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.05 | 669.348 | 25473 | 2378.302 | 1.126 |  | 5.75 |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.05 | 2378.302 | 91198 |  | 2378.302 | 4 | 7.07 | 76.9 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.07 | 17.151 | 687 | 82120.75 | 0.001 |  |  |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.4 | 275360.031 | 10737526 | 82120.75 | 13.412 |  | 46.01 |  | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.4 | 82120.75 | 3182716 |  | 82120.75 | 4 | 9.19 | 129.7
20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.87 | 13589.694 | 179296 | 3864.161 | 14.067 |  | 50.7 |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.78 | 246.998 | 8342 | 3864.161 | 0.256 |  | 2.18 |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.78 | 3864.161 | 132419 |  | 3864.161 | 4 | 4.78 | 19.4 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.81 | 1016.708 | 27596 | 109868.875 | 0.037 |  |  |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.96 | 109868.875 | 3072305 |  | 109868.875 | 4 | 6.63 | 65.7 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.05 | 469.471 | 17560 | 2149.042 | 0.874 |  | 4.23 |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.05 | 2149.042 | 80393 |  | 2149.042 | 4 | 6.39 | 59.8 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.07 | 17.723 | 611 | 76499.594 | 0.001 |  |  |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.4 | 245384.563 | 10366516 | 76499.594 | 12.831 |  | 43.99 |  | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.4 | 76499.594 | 3151665 |  | 76499.594 | 4 | 8.56 | 113.9
20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.86 | 62.169 | 761 | 5093.202 | 0.049 |  |  |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.78 | 6391.303 | 220071 | 5093.202 | 5.019 |  | 72.57 |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.78 | 5093.202 | 174541 |  | 5093.202 | 4 | 6.3 | 57.4 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.81 | 1968.833 | 59138 | 120155.836 | 0.066 |  |  |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.96 | 120155.836 | 3332802 |  | 120155.836 | 4 | 7.25 | 81.2 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.05 | 4630.355 | 168087 | 2534.763 | 7.307 |  | 42.83 |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.05 | 2534.763 | 92974 |  | 2534.763 | 4 | 7.54 | 88.5 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.06 | 81.691 | 3339 | 107375.094 | 0.003 |  |  |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.4 | 27795.627 | 1163048 | 107375.094 | 1.035 |  | 3.05 |  | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.4 | 107375.094 | 4487870 |  | 107375.094 | 4 | 12.01 | 200.3
20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte |  |  |  | 4939.262 |  |  |  |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.78 | 5906.037 | 203264 | 4939.262 | 4.783 |  | 69.08 |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.78 | 4939.262 | 170266 |  | 4939.262 | 4 | 6.11 | 52.7 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.81 | 1659.804 | 49492 | 112531.641 | 0.059 |  |  |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.96 | 112531.641 | 3182520 |  | 112531.641 | 4 | 6.79 | 69.7 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.05 | 4343.149 | 162971 | 2424.725 | 7.165 |  | 41.98 |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.05 | 2424.725 | 87860 |  | 2424.725 | 4 | 7.21 | 80.3 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.06 | 71.185 | 2151 | 104266.414 | 0.003 |  |  |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.4 | 4972.751 | 197217 | 104266.414 | 0.191 |  | 0.12 |  | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.4 | 104266.414 | 4273133 |  | 104266.414 | 4 | 11.66 | 191.6
20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.85 | 18.73 | 503 | 5233.792 | 0.014 |  |  |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.78 | 6012.225 | 207327 | 5233.792 | 4.595 |  | 66.3 |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.78 | 5233.792 | 181293 |  | 5233.792 | 4 | 6.47 | 61.8 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.81 | 1071.008 | 26239 | 118805.617 | 0.036 |  |  |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.96 | 118805.617 | 3380132 |  | 118805.617 | 4 | 7.17 | 79.2 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.05 | 4442.118 | 170191 | 2412.854 | 7.364 |  | 43.17 |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.05 | 2412.854 | 89939 |  | 2412.854 | 4 | 7.18 | 79.4 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.06 | 105.631 | 3989 | 105538.766 | 0.004 |  |  |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.4 | 5474.621 | 219701 | 105538.766 | 0.207 |  | 0.18 |  | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.4 | 105538.766 | 4219927 |  | 105538.766 | 4 | 11.81 | 195.2
20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte |  |  |  | 4130.82 |  |  |  |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.78 | 4317.092 | 148947 | 4130.82 | 4.18 |  | 60.17 |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.78 | 4130.82 | 140590 |  | 4130.82 | 4 | 5.11 | 27.7 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.81 | 339.214 | 5143 | 85510.195 | 0.016 |  |  |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.96 | 85510.195 | 2336976 |  | 85510.195 | 4 | 5.16 | 29 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.05 | 3521.64 | 135865 | 2093.272 | 6.729 |  | 39.37 |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.05 | 2093.272 | 76894 |  | 2093.272 | 4 | 6.23 | 55.7 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.06 | 5.17 | 309 | 78967.586 | 0 |  |  |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.4 | 1691.597 | 72046 | 78967.586 | 0.086 |  |  |  | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.4 | 78967.586 | 3313229 |  | 78967.586 | 4 | 8.83 | 120.8
20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.87 | 11.221 | 400 | 4509.017 | 0.01 |  |  |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.78 | 5081.018 | 175617 | 4509.017 | 4.507 |  | 65.01 |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.78 | 4509.017 | 155305 |  | 4509.017 | 4 | 5.58 | 39.4 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.81 | 492.186 | 8924 | 92018.813 | 0.021 |  |  |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.96 | 92018.813 | 2410122 |  | 92018.813 | 4 | 5.55 | 38.8 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.05 | 3449.08 | 131631 | 2165.184 | 6.372 |  | 37.22 |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.05 | 2165.184 | 82844 |  | 2165.184 | 4 | 6.44 | 61 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.03 | 0.403 | 62 | 83951.203 | 0 |  |  |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.4 | 2561.379 | 93938 | 83951.203 | 0.122 |  |  |  | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.4 | 83951.203 | 3239414 |  | 83951.203 | 4 | 9.39 | 134.8
20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.87 | 13.982 | 290 | 4377.103 | 0.013 |  |  |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.78 | 4593.416 | 157865 | 4377.103 | 4.198 |  | 60.43 |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.78 | 4377.103 | 150190 |  | 4377.103 | 4 | 5.41 | 35.3 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.81 | 1344.97 | 39408 | 87029.469 | 0.062 |  |  |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.97 | 87029.469 | 2377545 |  | 87029.469 | 4 | 5.25 | 31.3 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.05 | 3022.028 | 109144 | 2151.335 | 5.619 |  | 32.7 |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.05 | 2151.335 | 79762 |  | 2151.335 | 4 | 6.4 | 60 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.06 | 68.997 | 2377 | 82722.727 | 0.003 |  |  |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.39 | 4321.598 | 173179 | 82722.727 | 0.209 |  | 0.18 |  | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.4 | 82722.727 | 3247101 |  | 82722.727 | 4 | 9.25 | 131.3
20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.9 | 39.308 | 527 | 4518.646 | 0.035 |  |  |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.78 | 4942.229 | 169810 | 4518.646 | 4.375 |  | 63.05 |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.78 | 4518.646 | 155494 |  | 4518.646 | 4 | 5.59 | 39.7 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.8 | 1369.747 | 38750 | 102610.383 | 0.053 |  |  |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.96 | 102610.383 | 2811674 |  | 102610.383 | 4 | 6.19 | 54.8 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.05 | 3689.881 | 140619 | 2272.698 | 6.494 |  | 37.95 |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.05 | 2272.698 | 84725 |  | 2272.698 | 4 | 6.76 | 69 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.07 | 57.227 | 2282 | 84474.836 | 0.003 |  |  |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.4 | 2002.274 | 89927 | 84474.836 | 0.095 |  |  |  | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.4 | 84474.836 | 3404028 |  | 84474.836 | 4 | 9.45 | 136.2
20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.86 | 0.79 | 104 | 4595.066 | 0.001 |  |  |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.78 | 4874.581 | 169161 | 4595.066 | 4.243 |  | 61.1 |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.78 | 4595.066 | 158691 |  | 4595.066 | 4 | 5.68 | 42 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.85 | 313.797 | 3798 | 102214.219 | 0.012 |  |  |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.96 | 102214.219 | 2882559 |  | 102214.219 | 4 | 6.17 | 54.2 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.05 | 3383.043 | 128491 | 2150.592 | 6.292 |  | 36.74 |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.05 | 2150.592 | 81038 |  | 2150.592 | 4 | 6.4 | 59.9 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.07 | 6.926 | 306 | 87116.578 | 0 |  |  |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.4 | 2397.686 | 94151 | 87116.578 | 0.11 |  |  |  | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.4 | 87116.578 | 3474043 |  | 87116.578 | 4 | 9.75 | 143.6
20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.85 | 9.745 | 292 | 4516.958 | 0.009 |  |  |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.78 | 3923.402 | 133112 | 4516.958 | 3.474 |  | 49.74 |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.78 | 4516.958 | 153628 |  | 4516.958 | 4 | 5.58 | 39.6 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.81 | 374.854 | 4879 | 99535.336 | 0.015 |  |  |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.96 | 99535.336 | 2741889 |  | 99535.336 | 4 | 6 | 50.1 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.05 | 3382.356 | 128897 | 2036.234 | 6.644 |  | 38.85 |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.05 | 2036.234 | 75303 |  | 2036.234 | 4 | 6.06 | 51.4 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.06 | 3.727 | 259 | 83468.359 | 0 |  |  |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.4 | 2073.612 | 89065 | 83468.359 | 0.099 |  |  |  | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.4 | 83468.359 | 3439578 |  | 83468.359 | 4 | 9.34 | 133.4
20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte |  |  |  | 3747.828 |  |  |  |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.78 | 3518.875 | 119857 | 3747.828 | 3.756 |  | 53.9 |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.78 | 3747.828 | 127274 |  | 3747.828 | 4 | 4.63 | 15.8 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.81 | 20960.096 | 702734 | 89270.031 | 0.939 |  | 2.25 |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.96 | 89270.031 | 2472065 |  | 89270.031 | 4 | 5.39 | 34.6 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.05 | 2831.834 | 110285 | 1989.004 | 5.695 |  | 33.16 |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.05 | 1989.004 | 74074 |  | 1989.004 | 4 | 5.92 | 47.9 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.06 | 0.798 | 93 | 80214.852 | 0 |  |  |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.4 | 2641.276 | 106059 | 80214.852 | 0.132 |  |  |  | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.4 | 80214.852 | 3212792 |  | 80214.852 | 4 | 8.97 | 124.3
20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte |  |  |  | 4038.739 |  |  |  |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.78 | 3942.875 | 135332 | 4038.739 | 3.905 |  | 56.1 |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.78 | 4038.739 | 139785 |  | 4038.739 | 4 | 4.99 | 24.8 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.86 | 364.35 | 4413 | 85731.852 | 0.017 |  |  |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.96 | 85731.852 | 2440697 |  | 85731.852 | 4 | 5.17 | 29.3 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.05 | 2885.134 | 111975 | 1868.877 | 6.175 |  | 36.04 |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.05 | 1868.877 | 72443 |  | 1868.877 | 4 | 5.56 | 39 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte |  |  |  | 77124.359 |  |  |  |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.4 | 2297.828 | 98469 | 77124.359 | 0.119 |  |  |  | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.4 | 77124.359 | 3314510 |  | 77124.359 | 4 | 8.63 | 115.7
20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.87 | 21.437 | 444 | 3789.513 | 0.023 |  |  |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.78 | 3520.303 | 121928 | 3789.513 | 3.716 |  | 53.31 |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.78 | 3789.513 | 131680 |  | 3789.513 | 4 | 4.69 | 17.1 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.8 | 304.308 | 3837 | 82351.266 | 0.015 |  |  |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.96 | 82351.266 | 2320337 |  | 82351.266 | 4 | 4.97 | 24.2 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.05 | 2826.412 | 111737 | 1927.043 | 5.867 |  | 34.19 |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.05 | 1927.043 | 72505 |  | 1927.043 | 4 | 5.73 | 43.3 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.07 | 2.654 | 165 | 74714.813 | 0 |  |  |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.4 | 1735.724 | 69695 | 74714.813 | 0.093 |  |  |  | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.4 | 74714.813 | 3064155 |  | 74714.813 | 4 | 8.36 | 108.9
20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte |  |  |  | 4562.208 |  |  |  |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.78 | 4543.188 | 154753 | 4562.208 | 3.983 |  | 57.26 |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.78 | 4562.208 | 155516 |  | 4562.208 | 4 | 5.64 | 41 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.8 | 544.265 | 9791 | 103807.086 | 0.021 |  |  |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.96 | 103807.086 | 2960972 |  | 103807.086 | 4 | 6.26 | 56.6 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.05 | 3519.728 | 133013 | 2160.931 | 6.515 |  | 38.08 |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.05 | 2160.931 | 80933 |  | 2160.931 | 4 | 6.43 | 60.7 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.06 | 12.387 | 330 | 89927.414 | 0.001 |  |  |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.4 | 2722.367 | 114256 | 89927.414 | 0.121 |  |  |  | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.4 | 89927.414 | 3486817 |  | 89927.414 | 4 | 10.06 | 151.5
20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.88 | 0.404 | 57 | 4277.345 | 0 |  |  |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.78 | 4894.958 | 166429 | 4277.345 | 4.578 |  | 66.04 |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.78 | 4277.345 | 144960 |  | 4277.345 | 4 | 5.29 | 32.2 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.8 | 482.408 | 8308 | 99959.609 | 0.019 |  |  |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.96 | 99959.609 | 2879778 |  | 99959.609 | 4 | 6.03 | 50.8 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.05 | 3685.783 | 137360 | 2170.902 | 6.791 |  | 39.74 |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.05 | 2170.902 | 80815 |  | 2170.902 | 4 | 6.46 | 61.4 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.06 | 10.963 | 199 | 86421.602 | 0.001 |  |  |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.4 | 2252.211 | 97347 | 86421.602 | 0.104 |  |  |  | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.4 | 86421.602 | 3416719 |  | 86421.602 | 4 | 9.67 | 141.7
20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.87 | 15.454 | 583 | 4028.417 | 0.015 |  |  |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.78 | 4958.275 | 168396 | 4028.417 | 4.923 |  | 71.15 |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.78 | 4028.417 | 137600 |  | 4028.417 | 4 | 4.98 | 24.5 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.81 | 1167.067 | 33017 | 100739.109 | 0.046 |  |  |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.96 | 100739.109 | 2878863 |  | 100739.109 | 4 | 6.08 | 51.9 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.05 | 3722.396 | 140864 | 2075.98 | 7.172 |  | 42.02 |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.05 | 2075.98 | 76015 |  | 2075.98 | 4 | 6.18 | 54.4 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.07 | 13.604 | 430 | 91154.789 | 0.001 |  |  |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.39 | 3439.174 | 128733 | 91154.789 | 0.151 |  |  |  | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.4 | 91154.789 | 3585221 |  | 91154.789 | 4 | 10.2 | 154.9
20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.84 | 4.492 | 168 | 4201.701 | 0.004 |  |  |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.78 | 5418.749 | 187049 | 4201.701 | 5.159 |  | 74.63 |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.78 | 4201.701 | 142567 |  | 4201.701 | 4 | 5.2 | 29.9 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.81 | 260.508 | 9475 | 95184.141 | 0.011 |  |  |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.96 | 95184.141 | 2623040 |  | 95184.141 | 4 | 5.74 | 43.6 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.05 | 3417.414 | 128123 | 2070.364 | 6.603 |  | 38.6 |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.05 | 2070.364 | 77569 |  | 2070.364 | 4 | 6.16 | 54 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.03 | 16.884 | 380 | 89026.438 | 0.001 |  |  |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.4 | 2381.84 | 88870 | 89026.438 | 0.107 |  |  |  | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.4 | 89026.438 | 3511176 |  | 89026.438 | 4 | 9.96 | 149
20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.85 | 3.701 | 145 | 4522.662 | 0.003 |  |  |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.78 | 5020.021 | 174032 | 4522.662 | 4.44 |  | 64.01 |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.78 | 4522.662 | 156282 |  | 4522.662 | 4 | 5.59 | 39.8 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.81 | 3652.715 | 113771 | 106147.617 | 0.138 |  |  |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.96 | 106147.617 | 3012649 |  | 106147.617 | 4 | 6.4 | 60.1 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.05 | 3877.35 | 146603 | 2153.034 | 7.204 |  | 42.21 |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.05 | 2153.034 | 81167 |  | 2153.034 | 4 | 6.4 | 60.1 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.06 | 13.921 | 612 | 89150.375 | 0.001 |  |  |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.4 | 6998.725 | 272006 | 89150.375 | 0.314 |  | 0.55 |  | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.4 | 89150.375 | 3327173 |  | 89150.375 | 4 | 9.97 | 149.3
20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte |  |  |  | 4864.867 |  |  |  |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.78 | 5217.081 | 179277 | 4864.867 | 4.29 |  | 61.79 |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.78 | 4864.867 | 169224 |  | 4864.867 | 4 | 6.02 | 50.4 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.81 | 1195.365 | 32755 | 98089.266 | 0.049 |  |  |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.96 | 98089.266 | 2678359 |  | 98089.266 | 4 | 5.92 | 47.9 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.05 | 4665.503 | 182306 | 2243.382 | 8.319 |  | 48.9 |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.05 | 2243.382 | 84356 |  | 2243.382 | 4 | 6.67 | 66.8 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.07 | 29.097 | 1028 | 89808.383 | 0.001 |  |  |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.4 | 14859.413 | 603578 | 89808.383 | 0.662 |  | 1.75 |  | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.4 | 89808.383 | 3663026 |  | 89808.383 | 4 | 10.05 | 151.2
20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.87 | 38.099 | 1027 | 4646.139 | 0.033 |  |  |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.78 | 4634.335 | 157508 | 4646.139 | 3.99 |  | 57.36 |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.78 | 4646.139 | 159487 |  | 4646.139 | 4 | 5.74 | 43.6 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.81 | 370.412 | 5965 | 102671.5 | 0.014 |  |  |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.96 | 102671.5 | 2832084 |  | 102671.5 | 4 | 6.19 | 54.8 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.05 | 3195.274 | 116858 | 2224.932 | 5.744 |  | 33.46 |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.05 | 2224.932 | 81925 |  | 2224.932 | 4 | 6.62 | 65.5 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte |  |  |  | 99776.836 |  |  |  |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.4 | 2355.618 | 92222 | 99776.836 | 0.094 |  |  |  | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.4 | 99776.836 | 4067719 |  | 99776.836 | 4 | 11.16 | 179
20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte |  |  |  | 4670.868 |  |  |  |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.78 | 4617.606 | 157457 | 4670.868 | 3.954 |  | 56.83 |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.78 | 4670.868 | 159954 |  | 4670.868 | 4 | 5.78 | 44.4 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.81 | 368.678 | 5262 | 101879.82 | 0.014 |  |  |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.96 | 101879.82 | 2750035 |  | 101879.82 | 4 | 6.15 | 53.7 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.05 | 3959.748 | 152075 | 2450.528 | 6.464 |  | 37.77 |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.05 | 2450.528 | 94042 |  | 2450.528 | 4 | 7.29 | 82.2 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.06 | 55.627 | 2127 | 101979.219 | 0.002 |  |  |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.4 | 3240.659 | 122106 | 101979.219 | 0.127 |  |  |  | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.4 | 101979.219 | 4124754 |  | 101979.219 | 4 | 11.41 | 185.2
20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.88 | 34.705 | 698 | 4760.106 | 0.029 |  |  |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.78 | 5855.985 | 201916 | 4760.106 | 4.921 |  | 71.12 |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.78 | 4760.106 | 162646 |  | 4760.106 | 4 | 5.89 | 47.1 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.81 | 929.424 | 25504 | 119917.711 | 0.031 |  |  |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.97 | 119917.711 | 3374345 |  | 119917.711 | 4 | 7.23 | 80.9 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.05 | 4245.902 | 161078 | 2422.206 | 7.012 |  | 41.06 |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.05 | 2422.206 | 92834 |  | 2422.206 | 4 | 7.21 | 80.1 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.07 | 3.523 | 218 | 114534.242 | 0 |  |  |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.4 | 16277.225 | 691768 | 114534.242 | 0.568 |  | 1.43 |  | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.4 | 114534.242 | 4796443 |  | 114534.242 | 4 | 12.81 | 220.3
20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.87 | 16.086 | 360 | 4087.066 | 0.016 |  |  |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.78 | 4184.689 | 143212 | 4087.066 | 4.096 |  | 58.92 |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.78 | 4087.066 | 139402 |  | 4087.066 | 4 | 5.05 | 26.3 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.81 | 397.327 | 5254 | 97705.281 | 0.016 |  |  |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.96 | 97705.281 | 2789053 |  | 97705.281 | 4 | 5.89 | 47.4 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.05 | 3195.347 | 124660 | 2046.279 | 6.246 |  | 36.47 |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.05 | 2046.279 | 77503 |  | 2046.279 | 4 | 6.09 | 52.2 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.06 | 8.106 | 373 | 79844.398 | 0 |  |  |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.4 | 3266.747 | 141507 | 79844.398 | 0.164 |  | 0.02 |  | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.4 | 79844.398 | 3250808 |  | 79844.398 | 4 | 8.93 | 123.3
20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.88 | 16.546 | 403 | 4565.343 | 0.014 |  |  |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.78 | 4455.831 | 152479 | 4565.343 | 3.904 |  | 56.09 |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.78 | 4565.343 | 156691 |  | 4565.343 | 4 | 5.64 | 41.1 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.81 | 355.733 | 5451 | 102602.477 | 0.014 |  |  |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.96 | 102602.477 | 2908390 |  | 102602.477 | 4 | 6.19 | 54.7 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.05 | 3478.008 | 137993 | 2164.459 | 6.427 |  | 37.55 |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.05 | 2164.459 | 83044 |  | 2164.459 | 4 | 6.44 | 61 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.06 | 9.177 | 373 | 86026.953 | 0 |  |  |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.4 | 3671.382 | 155139 | 86026.953 | 0.171 |  | 0.05 |  | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.4 | 86026.953 | 3398443 |  | 86026.953 | 4 | 9.62 | 140.6
20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.85 | 20.354 | 479 | 4533.528 | 0.018 |  |  |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.78 | 4756.871 | 162191 | 4533.528 | 4.197 |  | 60.42 |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.78 | 4533.528 | 154364 |  | 4533.528 | 4 | 5.61 | 40.1 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.81 | 345.67 | 4848 | 91857.102 | 0.015 |  |  |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.96 | 91857.102 | 2594836 |  | 91857.102 | 4 | 5.54 | 38.5 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.05 | 3191.466 | 119308 | 2047.973 | 6.233 |  | 36.39 |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.05 | 2047.973 | 77307 |  | 2047.973 | 4 | 6.09 | 52.3 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.06 | 3.263 | 202 | 90566.922 | 0 |  |  |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.4 | 3701.034 | 150082 | 90566.922 | 0.163 |  | 0.02 |  | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.4 | 90566.922 | 3678126 |  | 90566.922 | 4 | 10.13 | 153.3
20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.88 | 4.288 | 199 | 4403.589 | 0.004 |  |  |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.78 | 4863.822 | 167123 | 4403.589 | 4.418 |  | 63.69 |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.78 | 4403.589 | 152070 |  | 4403.589 | 4 | 5.44 | 36.1 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.8 | 331.232 | 5083 | 96676.672 | 0.014 |  |  |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.96 | 96676.672 | 2624358 |  | 96676.672 | 4 | 5.83 | 45.8 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.05 | 3666.114 | 142234 | 2335.834 | 6.278 |  | 36.66 |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.05 | 2335.834 | 89409 |  | 2335.834 | 4 | 6.95 | 73.7 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.06 | 13.18 | 449 | 99319.828 | 0.001 |  |  |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.4 | 4481.813 | 193533 | 99319.828 | 0.181 |  | 0.08 |  | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.4 | 99319.828 | 4043231 |  | 99319.828 | 4 | 11.11 | 177.8
20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.87 | 11.475 | 363 | 4598.22 | 0.01 |  |  |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.78 | 5016.161 | 172989 | 4598.22 | 4.364 |  | 62.88 |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.78 | 4598.22 | 158526 |  | 4598.22 | 4 | 5.69 | 42.1 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.8 | 363.727 | 6758 | 106880.938 | 0.014 |  |  |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.96 | 106880.938 | 3036519 |  | 106880.938 | 4 | 6.45 | 61.2 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.05 | 3464.063 | 129174 | 2227.993 | 6.219 |  | 36.3 |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.05 | 2227.993 | 82124 |  | 2227.993 | 4 | 6.63 | 65.7 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte |  |  |  | 94711.563 |  |  |  |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.4 | 7830.961 | 317929 | 94711.563 | 0.331 |  | 0.6 |  | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.4 | 94711.563 | 3750901 |  | 94711.563 | 4 | 10.59 | 164.9
20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.85 | 10.074 | 350 | 4689.85 | 0.009 |  |  |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.78 | 5046.894 | 172993 | 4689.85 | 4.305 |  | 62.01 |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.78 | 4689.85 | 160258 |  | 4689.85 | 4 | 5.8 | 45 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.8 | 1032.07 | 27234 | 109563.195 | 0.038 |  |  |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.96 | 109563.195 | 3144449 |  | 109563.195 | 4 | 6.61 | 65.2 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.05 | 4187.168 | 160122 | 2176.578 | 7.695 |  | 45.16 |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.05 | 2176.578 | 83746 |  | 2176.578 | 4 | 6.47 | 61.9 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.06 | 7.088 | 331 | 90822.211 | 0 |  |  |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.4 | 9123.27 | 397031 | 90822.211 | 0.402 |  | 0.85 |  | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.4 | 90822.211 | 3627143 |  | 90822.211 | 4 | 10.16 | 154
20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.83 | 9.293 | 285 | 3744.223 | 0.01 |  |  |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.78 | 4283.556 | 146551 | 3744.223 | 4.576 |  | 66.02 |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.78 | 3744.223 | 127381 |  | 3744.223 | 4 | 4.63 | 15.7 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.8 | 629.296 | 15328 | 105342.867 | 0.024 |  |  |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.96 | 105342.867 | 3033331 |  | 105342.867 | 4 | 6.35 | 58.9 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.05 | 3289.579 | 127687 | 2107.887 | 6.242 |  | 36.44 |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.05 | 2107.887 | 80161 |  | 2107.887 | 4 | 6.27 | 56.8 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.06 | 38.406 | 1610 | 87124.805 | 0.002 |  |  |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.4 | 2383.649 | 96799 | 87124.805 | 0.109 |  |  |  | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.4 | 87124.805 | 3374102 |  | 87124.805 | 4 | 9.75 | 143.7
20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte |  |  |  | 4034.394 |  |  |  |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.78 | 4557.383 | 156709 | 4034.394 | 4.519 |  | 65.17 |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.78 | 4034.394 | 138969 |  | 4034.394 | 4 | 4.99 | 24.7 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.85 | 294.01 | 4664 | 98521.195 | 0.012 |  |  |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.97 | 98521.195 | 2867227 |  | 98521.195 | 4 | 5.94 | 48.6 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.05 | 2823.59 | 105276 | 2019.512 | 5.593 |  | 32.54 |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.05 | 2019.512 | 74692 |  | 2019.512 | 4 | 6.01 | 50.2 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.02 | 2.931 | 191 | 77005.422 | 0 |  |  |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.4 | 1627.43 | 65057 | 77005.422 | 0.085 |  |  |  | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.4 | 77005.422 | 3313557 |  | 77005.422 | 4 | 8.61 | 115.4
20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.87 | 19.84 | 431 | 4172.714 | 0.019 |  |  |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.78 | 4135.498 | 140199 | 4172.714 | 3.964 |  | 56.98 |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.78 | 4172.714 | 143367 |  | 4172.714 | 4 | 5.16 | 29 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.81 | 312.739 | 4898 | 87807.195 | 0.014 |  |  |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.97 | 87807.195 | 2354495 |  | 87807.195 | 4 | 5.3 | 32.4 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.05 | 3351.329 | 126030 | 2174.987 | 6.163 |  | 35.97 |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.05 | 2174.987 | 82892 |  | 2174.987 | 4 | 6.47 | 61.7 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.08 | 9.604 | 240 | 88369.641 | 0 |  |  |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.4 | 3482.942 | 128910 | 88369.641 | 0.158 |  | 0 |  | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.4 | 88369.641 | 3280962 |  | 88369.641 | 4 | 9.89 | 147.1
20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.88 | 6.985 | 141 | 4132.267 | 0.007 |  |  |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.78 | 4255.095 | 146051 | 4132.267 | 4.119 |  | 59.26 |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.78 | 4132.267 | 141081 |  | 4132.267 | 4 | 5.11 | 27.7 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.81 | 433.407 | 8770 | 99443.953 | 0.017 |  |  |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.97 | 99443.953 | 2784842 |  | 99443.953 | 4 | 6 | 50 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.05 | 3326.948 | 127478 | 2189.358 | 6.078 |  | 35.46 |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.05 | 2189.358 | 78834 |  | 2189.358 | 4 | 6.51 | 62.8 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.07 | 5.175 | 302 | 84107.078 | 0 |  |  |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.4 | 2338.676 | 92326 | 84107.078 | 0.111 |  |  |  | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.4 | 84107.078 | 3258718 |  | 84107.078 | 4 | 9.41 | 135.2
20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.86 | 50.12 | 1129 | 4125.632 | 0.049 |  |  |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.78 | 3951.575 | 134733 | 4125.632 | 3.831 |  | 55.01 |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.78 | 4125.632 | 141685 |  | 4125.632 | 4 | 5.1 | 27.5 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.86 | 279.778 | 3623 | 106413.328 | 0.011 |  |  |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.96 | 106413.328 | 3078344 |  | 106413.328 | 4 | 6.42 | 60.5 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.05 | 3317.284 | 126490 | 2193.385 | 6.05 |  | 35.29 |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.05 | 2193.385 | 83567 |  | 2193.385 | 4 | 6.52 | 63.1 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte |  |  |  | 82892.469 |  |  |  |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.4 | 1933.72 | 74930 | 82892.469 | 0.093 |  |  |  | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.4 | 82892.469 | 3306039 |  | 82892.469 | 4 | 9.27 | 131.8
20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.88 | 5.404 | 192 | 4422.781 | 0.005 |  |  |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.78 | 4628.658 | 158950 | 4422.781 | 4.186 |  | 60.26 |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.78 | 4422.781 | 153183 |  | 4422.781 | 4 | 5.47 | 36.7 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte |  |  |  | 109031.305 |  |  |  |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.97 | 109031.305 | 3128109 |  | 109031.305 | 4 | 6.58 | 64.4 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.05 | 3435.218 | 132714 | 2207.604 | 6.224 |  | 36.33 |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.05 | 2207.604 | 80918 |  | 2207.604 | 4 | 6.57 | 64.2 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte |  |  |  | 89697.148 |  |  |  |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.4 | 2124.865 | 95781 | 89697.148 | 0.095 |  |  |  | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.4 | 89697.148 | 3831743 |  | 89697.148 | 4 | 10.03 | 150.8
20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte |  |  |  | 3922.801 |  |  |  |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.78 | 3757.041 | 129864 | 3922.801 | 3.831 |  | 55.01 |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.78 | 3922.801 | 137534 |  | 3922.801 | 4 | 4.85 | 21.3 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.86 | 289.661 | 4097 | 84955.523 | 0.014 |  |  |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.97 | 84955.523 | 2334746 |  | 84955.523 | 4 | 5.13 | 28.1 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.05 | 2774.165 | 104724 | 1885.64 | 5.885 |  | 34.3 |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.05 | 1885.64 | 69423 |  | 1885.64 | 4 | 5.61 | 40.2 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.06 | 0.168 | 31 | 76210.016 | 0 |  |  |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.4 | 1741.286 | 67916 | 76210.016 | 0.091 |  |  |  | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.4 | 76210.016 | 3155056 |  | 76210.016 | 4 | 8.53 | 113.1
20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.85 | 29.131 | 366 | 3270.064 | 0.036 |  |  |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.78 | 3527.55 | 121269 | 3270.064 | 4.315 |  | 62.16 |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.78 | 3270.064 | 113843 |  | 3270.064 | 4 | 4.04 | 1.1 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte |  |  |  | 83049.625 |  |  |  |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.96 | 83049.625 | 2385833 |  | 83049.625 | 4 | 5.01 | 25.3 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.05 | 2413.183 | 95545 | 1554.347 | 6.21 |  | 36.25 |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.05 | 1554.347 | 59618 |  | 1554.347 | 4 | 4.62 | 15.6 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.09 | 0.105 | 18 | 52721.07 | 0 |  |  |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.4 | 1187.428 | 48088 | 52721.07 | 0.09 |  |  |  | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.4 | 52721.07 | 2267876 |  | 52721.07 | 4 | 5.9 | 47.4
20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte |  |  |  | 4049.156 |  |  |  |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.78 | 4092.351 | 141067 | 4049.156 | 4.043 |  | 58.14 |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.78 | 4049.156 | 140243 |  | 4049.156 | 4 | 5.01 | 25.2 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.85 | 285.751 | 3659 | 75152.219 | 0.015 |  |  |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.97 | 75152.219 | 2040597 |  | 75152.219 | 4 | 4.53 | 13.3 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.05 | 2684.16 | 102251 | 1868.361 | 5.747 |  | 33.47 |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.05 | 1868.361 | 70814 |  | 1868.361 | 4 | 5.56 | 38.9 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.07 | 9.736 | 352 | 55817.023 | 0.001 |  |  |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.4 | 1416.868 | 47558 | 55817.023 | 0.102 |  |  |  | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.4 | 55817.023 | 2277350 |  | 55817.023 | 4 | 6.24 | 56.1
20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.85 | 8.917 | 245 | 4699.705 | 0.008 |  |  |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.78 | 4871.632 | 168775 | 4699.705 | 4.146 |  | 59.67 |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.78 | 4699.705 | 162117 |  | 4699.705 | 4 | 5.81 | 45.3 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.81 | 346.74 | 5135 | 100056 | 0.014 |  |  |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.96 | 100056 | 2828201 |  | 100056 | 4 | 6.04 | 50.9 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.05 | 3227.542 | 125488 | 2186.832 | 5.904 |  | 34.41 |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.05 | 2186.832 | 81592 |  | 2186.832 | 4 | 6.5 | 62.6 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.07 | 8.217 | 261 | 84959.609 | 0 |  |  |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.4 | 2743.245 | 102340 | 84959.609 | 0.129 |  |  |  | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.4 | 84959.609 | 3222198 |  | 84959.609 | 4 | 9.5 | 137.6
20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.88 | 44.112 | 881 | 4633.228 | 0.038 |  |  |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.78 | 4619.891 | 158711 | 4633.228 | 3.988 |  | 57.34 |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.78 | 4633.228 | 160681 |  | 4633.228 | 4 | 5.73 | 43.2 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.8 | 493.649 | 10680 | 89716.656 | 0.022 |  |  |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.96 | 89716.656 | 2419546 |  | 89716.656 | 4 | 5.41 | 35.3 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.05 | 3479.264 | 131384 | 2072.819 | 6.714 |  | 39.27 |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.05 | 2072.819 | 77608 |  | 2072.819 | 4 | 6.17 | 54.1 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte |  |  |  | 86283.852 |  |  |  |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.4 | 3430.022 | 134885 | 86283.852 | 0.159 |  | 0.01 |  | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.4 | 86283.852 | 3504027 |  | 86283.852 | 4 | 9.65 | 141.3
20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.88 | 9.015 | 235 | 4931.43 | 0.007 |  |  |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.78 | 5559.856 | 190760 | 4931.43 | 4.51 |  | 65.04 |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.78 | 4931.43 | 170328 |  | 4931.43 | 4 | 6.1 | 52.4 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.85 | 335.928 | 4351 | 109559.086 | 0.012 |  |  |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.97 | 109559.086 | 2959963 |  | 109559.086 | 4 | 6.61 | 65.2 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.05 | 3257.239 | 117830 | 2269.267 | 5.741 |  | 33.44 |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.05 | 2269.267 | 85774 |  | 2269.267 | 4 | 6.75 | 68.8 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.07 | 13.263 | 417 | 100046.359 | 0.001 |  |  |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.4 | 2008.583 | 88863 | 100046.359 | 0.08 |  |  |  | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.4 | 100046.359 | 4210013 |  | 100046.359 | 4 | 11.19 | 179.8
20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.85 | 10.279 | 210 | 4408.869 | 0.009 |  |  |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.78 | 4127.725 | 143210 | 4408.869 | 3.745 |  | 53.74 |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.78 | 4408.869 | 152599 |  | 4408.869 | 4 | 5.45 | 36.3 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.81 | 556.298 | 12640 | 90430.555 | 0.025 |  |  |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.96 | 90430.555 | 2440787 |  | 90430.555 | 4 | 5.46 | 36.4 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.05 | 3758.575 | 140254 | 2124.282 | 7.077 |  | 41.45 |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.05 | 2124.282 | 80191 |  | 2124.282 | 4 | 6.32 | 58 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.07 | 7.463 | 338 | 95674.383 | 0 |  |  |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.4 | 14150.011 | 605117 | 95674.383 | 0.592 |  | 1.51 |  | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.4 | 95674.383 | 4084776 |  | 95674.383 | 4 | 10.7 | 167.6
20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte |  |  |  | 4910.336 |  |  |  |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.78 | 4334.655 | 148483 | 4910.336 | 3.531 |  | 50.58 |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.78 | 4910.336 | 169194 |  | 4910.336 | 4 | 6.07 | 51.8 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.81 | 745.601 | 16891 | 102643.82 | 0.029 |  |  |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.97 | 102643.82 | 2885297 |  | 102643.82 | 4 | 6.19 | 54.8 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.05 | 3681.078 | 139784 | 2198.032 | 6.699 |  | 39.18 |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.05 | 2198.032 | 83709 |  | 2198.032 | 4 | 6.54 | 63.5 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte |  |  |  | 93804.273 |  |  |  |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.4 | 6172.626 | 255818 | 93804.273 | 0.263 |  | 0.37 |  | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.4 | 93804.273 | 3910266 |  | 93804.273 | 4 | 10.49 | 162.3
20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.85 | 10.55 | 300 | 4659.491 | 0.009 |  |  |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.78 | 4811.418 | 164882 | 4659.491 | 4.13 |  | 59.43 |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.78 | 4659.491 | 160745 |  | 4659.491 | 4 | 5.76 | 44 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.81 | 662.704 | 14954 | 96604.727 | 0.027 |  |  |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.96 | 96604.727 | 2642798 |  | 96604.727 | 4 | 5.83 | 45.7 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.05 | 3579.895 | 132544 | 2234.033 | 6.41 |  | 37.45 |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.05 | 2234.033 | 83528 |  | 2234.033 | 4 | 6.65 | 66.1 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.11 | 1.75 | 98 | 75229.672 | 0 |  |  |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.4 | 5040.559 | 197915 | 75229.672 | 0.268 |  | 0.39 |  | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.4 | 75229.672 | 2819570 |  | 75229.672 | 4 | 8.42 | 110.4
20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte |  |  |  | 4314.995 |  |  |  |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.78 | 65.511 | 2011 | 4314.995 | 0.061 |  |  |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.78 | 4314.995 | 149357 |  | 4314.995 | 4 | 5.34 | 33.4 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.81 | 211102.141 | 6984337 | 122806.492 | 6.876 |  | 19.09 |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.97 | 122806.492 | 3541354 |  | 122806.492 | 4 | 7.41 | 85.2 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.05 | 444.989 | 16010 | 2325.502 | 0.765 |  | 3.58 |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.05 | 2325.502 | 83837 |  | 2325.502 | 4 | 6.92 | 72.9 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.07 | 84765.813 | 3149459 | 97325.922 | 3.484 |  | 57.55 |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.4 | 4395.498 | 171055 | 97325.922 | 0.181 |  | 0.08 |  | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.4 | 97325.922 | 4022685 |  | 97325.922 | 4 | 10.89 | 172.2
20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.9 | 8.644 | 217 | 4169.22 | 0.008 |  |  |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.78 | 90.844 | 2571 | 4169.22 | 0.087 |  |  |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.78 | 4169.22 | 142828 |  | 4169.22 | 4 | 5.15 | 28.9 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.81 | 308503.531 | 10401536 | 105540.742 | 11.692 |  | 32.75 |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.96 | 105540.742 | 2950542 |  | 105540.742 | 4 | 6.37 | 59.2 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.05 | 169.976 | 6440 | 2318.157 | 0.293 |  | 0.75 |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.05 | 2318.157 | 89454 |  | 2318.157 | 4 | 6.9 | 72.4 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.06 | 89158.578 | 3440956 | 102662.727 | 3.474 |  | 57.38 |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.4 | 4938.487 | 182785 | 102662.727 | 0.192 |  | 0.12 |  | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.4 | 102662.727 | 3738008 |  | 102662.727 | 4 | 11.48 | 187.1
20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.88 | 5.067 | 195 | 4343.747 | 0.005 |  |  |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.78 | 59.138 | 1829 | 4343.747 | 0.054 |  |  |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.78 | 4343.747 | 149701 |  | 4343.747 | 4 | 5.37 | 34.3 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.8 | 356524.719 | 12119818 | 108226.414 | 13.177 |  | 36.97 |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.96 | 108226.414 | 2970949 |  | 108226.414 | 4 | 6.53 | 63.2 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.05 | 77.56 | 3248 | 2446.042 | 0.127 |  |  |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.05 | 2446.042 | 96028 |  | 2446.042 | 4 | 7.28 | 81.9 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.06 | 89658.922 | 3423202 | 99524.109 | 3.604 |  | 59.55 |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.4 | 3017.959 | 122122 | 99524.109 | 0.121 |  |  |  | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.4 | 99524.109 | 3978316 |  | 99524.109 | 4 | 11.13 | 178.3
20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.84 | 3.194 | 60 | 3898.605 | 0.003 |  |  |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.78 | 229.48 | 8090 | 3898.605 | 0.235 |  | 1.88 |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.78 | 3898.605 | 133591 |  | 3898.605 | 4 | 4.82 | 20.5 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.81 | 283991.438 | 9699914 | 103418.438 | 10.984 |  | 30.75 |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.96 | 103418.438 | 3004115 |  | 103418.438 | 4 | 6.24 | 56 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.05 | 71.583 | 3014 | 2244.717 | 0.128 |  |  |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.05 | 2244.717 | 87807 |  | 2244.717 | 4 | 6.68 | 66.9 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.06 | 82879.844 | 2973760 | 93904.141 | 3.53 |  | 58.33 |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.4 | 2398.448 | 106629 | 93904.141 | 0.102 |  |  |  | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.4 | 93904.141 | 4091944 |  | 93904.141 | 4 | 10.5 | 162.6
20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.88 | 14.696 | 311 | 4061.239 | 0.014 |  |  |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.78 | 54.196 | 1918 | 4061.239 | 0.053 |  |  |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.78 | 4061.239 | 140995 |  | 4061.239 | 4 | 5.02 | 25.5 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.81 | 303358.25 | 10312001 | 106341.625 | 11.411 |  | 31.95 |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.97 | 106341.625 | 3036697 |  | 106341.625 | 4 | 6.42 | 60.4 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.05 | 45.632 | 2059 | 2253.734 | 0.081 |  |  |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.05 | 2253.734 | 90122 |  | 2253.734 | 4 | 6.7 | 67.6 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.06 | 83693.422 | 3138311 | 99092.422 | 3.378 |  | 55.79 |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.4 | 2452.679 | 97232 | 99092.422 | 0.099 |  |  |  | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.4 | 99092.422 | 3960975 |  | 99092.422 | 4 | 11.08 | 177.1
20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte |  |  |  | 3517.021 |  |  |  |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.78 | 73.433 | 2244 | 3517.021 | 0.084 |  |  |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.78 | 3517.021 | 121705 |  | 3517.021 | 4 | 4.35 | 8.7 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.81 | 279995.094 | 9373390 | 92977.766 | 12.046 |  | 33.76 |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.97 | 92977.766 | 2623362 |  | 92977.766 | 4 | 5.61 | 40.2 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.05 | 68.436 | 2773 | 2225.482 | 0.123 |  |  |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.05 | 2225.482 | 85507 |  | 2225.482 | 4 | 6.62 | 65.5 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.07 | 85406.219 | 3482632 | 102824.828 | 3.322 |  | 54.85 |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.4 | 3060.941 | 112753 | 102824.828 | 0.119 |  |  |  | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.4 | 102824.828 | 3514491 |  | 102824.828 | 4 | 11.5 | 187.6
20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte |  |  |  | 3994.895 |  |  |  |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.78 | 42.996 | 1394 | 3994.895 | 0.043 |  |  |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.77 | 3994.895 | 137003 |  | 3994.895 | 4 | 4.94 | 23.5 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.81 | 384793.406 | 13078994 | 98816.094 | 15.576 |  | 43.77 |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.96 | 98816.094 | 2714665 |  | 98816.094 | 4 | 5.96 | 49 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.05 | 65.056 | 3165 | 2371.027 | 0.11 |  |  |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.05 | 2371.027 | 93883 |  | 2371.027 | 4 | 7.05 | 76.3 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.06 | 83136.398 | 3056431 | 104638.984 | 3.178 |  | 52.44 |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.4 | 2407.487 | 92725 | 104638.984 | 0.092 |  |  |  | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.4 | 104638.984 | 4058627 |  | 104638.984 | 4 | 11.71 | 192.6
20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte |  |  |  | 3869.739 |  |  |  |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.78 | 39.354 | 1325 | 3869.739 | 0.041 |  |  |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.78 | 3869.739 | 132127 |  | 3869.739 | 4 | 4.78 | 19.6 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.81 | 283668.125 | 9436496 | 114010.602 | 9.952 |  | 27.82 |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.96 | 114010.602 | 3199621 |  | 114010.602 | 4 | 6.88 | 71.9 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.05 | 59.29 | 2368 | 2403.863 | 0.099 |  |  |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.05 | 2403.863 | 91589 |  | 2403.863 | 4 | 7.15 | 78.8 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.06 | 73273.953 | 2781368 | 95722.523 | 3.062 |  | 50.5 |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.4 | 2357.775 | 96783 | 95722.523 | 0.099 |  |  |  | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.4 | 95722.523 | 3923375 |  | 95722.523 | 4 | 10.71 | 167.7
20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte |  |  |  | 4094.355 |  |  |  |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.78 | 65.363 | 1799 | 4094.355 | 0.064 |  |  |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.78 | 4094.355 | 141335 |  | 4094.355 | 4 | 5.06 | 26.6 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.81 | 270666.656 | 8983324 | 111625.508 | 9.699 |  | 27.1 |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.97 | 111625.508 | 3109118 |  | 111625.508 | 4 | 6.73 | 68.3 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.05 | 51.517 | 2109 | 2416.499 | 0.085 |  |  |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.05 | 2416.499 | 92311 |  | 2416.499 | 4 | 7.19 | 79.7 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.07 | 91674.398 | 3524289 | 104475.281 | 3.51 |  | 57.98 |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.4 | 2659.664 | 107623 | 104475.281 | 0.102 |  |  |  | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.4 | 104475.281 | 4375782 |  | 104475.281 | 4 | 11.69 | 192.2
20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte |  |  |  | 3202.326 |  |  |  |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.78 | 29.845 | 1106 | 3202.326 | 0.037 |  |  |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.78 | 3202.326 | 110024 |  | 3202.326 | 4 | 3.96 | -1 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.81 | 210909.141 | 7091518 | 79338.781 | 10.633 |  | 29.75 |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.97 | 79338.781 | 2209920 |  | 79338.781 | 4 | 4.79 | 19.7 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.05 | 68.165 | 2481 | 1927.321 | 0.141 |  |  |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.05 | 1927.321 | 76090 |  | 1927.321 | 4 | 5.73 | 43.3 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.07 | 60181.426 | 2252242 | 74625.852 | 3.226 |  | 53.24 |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.4 | 2650.645 | 109312 | 74625.852 | 0.142 |  |  |  | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.4 | 74625.852 | 2961373 |  | 74625.852 | 4 | 8.35 | 108.7
20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte |  |  |  | 3661.565 |  |  |  |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.78 | 51.159 | 1557 | 3661.565 | 0.056 |  |  |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.78 | 3661.565 | 126574 |  | 3661.565 | 4 | 4.53 | 13.2 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.81 | 219247.406 | 7365930 | 92511.57 | 9.48 |  | 26.48 |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.97 | 92511.57 | 2666977 |  | 92511.57 | 4 | 5.58 | 39.5 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.05 | 68.579 | 2800 | 1956.9 | 0.14 |  |  |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.05 | 1956.9 | 75761 |  | 1956.9 | 4 | 5.82 | 45.5 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.07 | 66792.984 | 2494126 | 78271.125 | 3.413 |  | 56.37 |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.4 | 2487.767 | 106960 | 78271.125 | 0.127 |  |  |  | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.4 | 78271.125 | 3317273 |  | 78271.125 | 4 | 8.76 | 118.9
20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.85 | 3.96 | 193 | 2804.06 | 0.006 |  |  |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.78 | 58.809 | 1433 | 2804.06 | 0.084 |  |  |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.78 | 2804.06 | 95771 |  | 2804.06 | 4 | 3.47 | -13.3 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.81 | 177725.781 | 5986646 | 85295.422 | 8.335 |  | 23.23 |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.97 | 85295.422 | 2390381 |  | 85295.422 | 4 | 5.15 | 28.6 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.05 | 44.038 | 1395 | 1911.096 | 0.092 |  |  |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.05 | 1911.096 | 76727 |  | 1911.096 | 4 | 5.68 | 42.1 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.07 | 63163.129 | 2425594 | 82232.203 | 3.072 |  | 50.68 |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.4 | 1859.349 | 76758 | 82232.203 | 0.09 |  |  |  | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.4 | 82232.203 | 3383647 |  | 82232.203 | 4 | 9.2 | 130
20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.86 | 51.112 | 1346 | 4088.134 | 0.05 |  |  |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.78 | 164.218 | 5442 | 4088.134 | 0.161 |  | 0.77 |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.78 | 4088.134 | 141052 |  | 4088.134 | 4 | 5.05 | 26.4 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.8 | 298650.938 | 9971744 | 111956.539 | 10.67 |  | 29.85 |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.96 | 111956.539 | 3172459 |  | 111956.539 | 4 | 6.75 | 68.8 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.05 | 134.341 | 5716 | 2279.891 | 0.236 |  | 0.41 |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.05 | 2279.891 | 87427 |  | 2279.891 | 4 | 6.78 | 69.5 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.06 | 69076.906 | 2644574 | 96895.539 | 2.852 |  | 46.99 |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.4 | 6274.098 | 254249 | 96895.539 | 0.259 |  | 0.35 |  | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.4 | 96895.539 | 3608709 |  | 96895.539 | 4 | 10.84 | 171
20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.82 | 8.796 | 183 | 3656.892 | 0.01 |  |  |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.78 | 66.106 | 1966 | 3656.892 | 0.072 |  |  |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.78 | 3656.892 | 126601 |  | 3656.892 | 4 | 4.52 | 13 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.81 | 333344 | 11203420 | 104588.813 | 12.749 |  | 35.75 |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.96 | 104588.813 | 2995732 |  | 104588.813 | 4 | 6.31 | 57.7 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.05 | 120.986 | 4204 | 2295.882 | 0.211 |  | 0.26 |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.05 | 2295.882 | 91146 |  | 2295.882 | 4 | 6.83 | 70.7 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.06 | 80887.133 | 3075320 | 90586.703 | 3.572 |  | 59.02 |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.4 | 2641.203 | 106814 | 90586.703 | 0.117 |  |  |  | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.4 | 90586.703 | 3675424 |  | 90586.703 | 4 | 10.13 | 153.3
20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.9 | 3.336 | 179 | 4115.619 | 0.003 |  |  |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.77 | 67.688 | 2189 | 4115.619 | 0.066 |  |  |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.78 | 4115.619 | 141480 |  | 4115.619 | 4 | 5.09 | 27.2 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.8 | 274019.813 | 9070789 | 87999.492 | 12.456 |  | 34.92 |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.96 | 87999.492 | 2312692 |  | 87999.492 | 4 | 5.31 | 32.7 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.05 | 66.306 | 2725 | 2274.982 | 0.117 |  |  |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.05 | 2274.982 | 87382 |  | 2274.982 | 4 | 6.77 | 69.2 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.06 | 88501.672 | 3452261 | 85220.258 | 4.154 |  | 68.74 |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.4 | 2226.196 | 79075 | 85220.258 | 0.104 |  |  |  | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.4 | 85220.258 | 3006224 |  | 85220.258 | 4 | 9.53 | 138.3
20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.89 | 5.356 | 204 | 3668.877 | 0.006 |  |  |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.78 | 58.988 | 1991 | 3668.877 | 0.064 |  |  |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.78 | 3668.877 | 125748 |  | 3668.877 | 4 | 4.54 | 13.4 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.81 | 299645.75 | 10020487 | 97360.305 | 12.311 |  | 34.51 |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.97 | 97360.305 | 2799244 |  | 97360.305 | 4 | 5.87 | 46.8 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.05 | 85.625 | 3468 | 1978.307 | 0.173 |  | 0.03 |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.05 | 1978.307 | 76393 |  | 1978.307 | 4 | 5.88 | 47.1 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.06 | 83890.813 | 3278919 | 87958.188 | 3.815 |  | 63.08 |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.4 | 2858.172 | 106053 | 87958.188 | 0.13 |  |  |  | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.4 | 87958.188 | 3322711 |  | 87958.188 | 4 | 9.84 | 146
20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.9 | 4.782 | 169 | 3204.234 | 0.006 |  |  |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.78 | 114.769 | 3960 | 3204.234 | 0.143 |  | 0.51 |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.78 | 3204.234 | 110651 |  | 3204.234 | 4 | 3.96 | -1 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.81 | 235749.125 | 7878546 | 84831.938 | 11.116 |  | 31.12 |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.96 | 84831.938 | 2370620 |  | 84831.938 | 4 | 5.12 | 27.9 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.05 | 98.728 | 3906 | 1924.4 | 0.205 |  | 0.22 |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.05 | 1924.4 | 73973 |  | 1924.4 | 4 | 5.72 | 43.1 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.06 | 74294.75 | 2865533 | 73946.117 | 4.019 |  | 66.48 |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.4 | 3113.741 | 112647 | 73946.117 | 0.168 |  | 0.04 |  | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.4 | 73946.117 | 2831281 |  | 73946.117 | 4 | 8.27 | 106.8
20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.84 | 3.179 | 97 | 3915.398 | 0.003 |  |  |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.77 | 205.511 | 7108 | 3915.398 | 0.21 |  | 1.5 |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.78 | 3915.398 | 135372 |  | 3915.398 | 4 | 4.84 | 21 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.8 | 323462.594 | 11000032 | 92662.234 | 13.963 |  | 39.19 |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.96 | 92662.234 | 2609777 |  | 92662.234 | 4 | 5.59 | 39.7 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.05 | 162.785 | 6639 | 1966.173 | 0.331 |  | 0.98 |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.05 | 1966.173 | 76220 |  | 1966.173 | 4 | 5.85 | 46.2 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.06 | 88411.766 | 3434985 | 88629.414 | 3.99 |  | 66 |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.4 | 5364.324 | 218230 | 88629.414 | 0.242 |  | 0.3 |  | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.4 | 88629.414 | 3533247 |  | 88629.414 | 4 | 9.91 | 147.9
20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte |  |  |  | 4269.086 |  |  |  |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.78 | 110.167 | 3794 | 4269.086 | 0.103 |  |  |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.78 | 4269.086 | 147221 |  | 4269.086 | 4 | 5.28 | 32 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.81 | 352025.063 | 11860594 | 120083.727 | 11.726 |  | 32.85 |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.97 | 120083.727 | 3464551 |  | 120083.727 | 4 | 7.24 | 81.1 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.05 | 189.058 | 7842 | 2702.162 | 0.28 |  | 0.67 |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.05 | 2702.162 | 105538 |  | 2702.162 | 4 | 8.04 | 100.9 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.06 | 98320.117 | 3693365 | 180164.406 | 2.183 |  | 35.82 |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.4 | 6702.993 | 251062 | 180164.406 | 0.149 |  |  |  | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.4 | 180164.406 | 6972059 |  | 180164.406 | 4 | 20.15 | 403.9
20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.9 | 0.276 | 37 | 4165.854 | 0 |  |  |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.78 | 102.365 | 3701 | 4165.854 | 0.098 |  |  |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.78 | 4165.854 | 143861 |  | 4165.854 | 4 | 5.15 | 28.8 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.81 | 261868.188 | 8873026 | 120913.375 | 8.663 |  | 24.16 |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.97 | 120913.375 | 3529813 |  | 120913.375 | 4 | 7.29 | 82.4 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.05 | 105.691 | 4814 | 2421.035 | 0.175 |  | 0.04 |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.05 | 2421.035 | 92806 |  | 2421.035 | 4 | 7.2 | 80 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.07 | 84453.852 | 3030661 | 128731.063 | 2.624 |  | 43.19 |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.4 | 3526.294 | 137227 | 128731.063 | 0.11 |  |  |  | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.4 | 128731.063 | 5090614 |  | 128731.063 | 4 | 14.4 | 260
20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.87 | 3.55 | 130 | 4105.743 | 0.003 |  |  |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.78 | 144.584 | 5057 | 4105.743 | 0.141 |  | 0.48 |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.78 | 4105.743 | 139697 |  | 4105.743 | 4 | 5.08 | 26.9 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.81 | 358415.75 | 12138892 | 120947.359 | 11.854 |  | 33.21 |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.97 | 120947.359 | 3381191 |  | 120947.359 | 4 | 7.3 | 82.4 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.05 | 149.801 | 6328 | 2636.327 | 0.227 |  | 0.36 |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.05 | 2636.327 | 101850 |  | 2636.327 | 4 | 7.84 | 96 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.06 | 98009.695 | 3569567 | 119917.117 | 3.269 |  | 53.96 |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.4 | 3763.688 | 148501 | 119917.117 | 0.126 |  |  |  | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.4 | 119917.117 | 4419736 |  | 119917.117 | 4 | 13.41 | 235.4
20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.91 | 5.588 | 229 | 4091.521 | 0.005 |  |  |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.78 | 301.788 | 9552 | 4091.521 | 0.295 |  | 2.76 |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.78 | 4091.521 | 141866 |  | 4091.521 | 4 | 5.06 | 26.5 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.81 | 284973.594 | 9606363 | 98304.648 | 11.596 |  | 32.48 |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.97 | 98304.648 | 2728535 |  | 98304.648 | 4 | 5.93 | 48.3 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.05 | 311.631 | 12419 | 2181.149 | 0.571 |  | 2.42 |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.05 | 2181.149 | 82489 |  | 2181.149 | 4 | 6.49 | 62.2 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.06 | 66279.32 | 2477570 | 88866.891 | 2.983 |  | 49.19 |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.4 | 9193.139 | 358467 | 88866.891 | 0.414 |  | 0.89 |  | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.4 | 88866.891 | 3528454 |  | 88866.891 | 4 | 9.94 | 148.5
20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.84 | 5.883 | 249 | 3699.73 | 0.006 |  |  |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.78 | 299.284 | 9875 | 3699.73 | 0.324 |  | 3.18 |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.78 | 3699.73 | 127092 |  | 3699.73 | 4 | 4.57 | 14.4 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.81 | 236682.141 | 7911366 | 108873.719 | 8.696 |  | 24.25 |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.97 | 108873.719 | 3186634 |  | 108873.719 | 4 | 6.57 | 64.2 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.05 | 312.361 | 11999 | 2069.581 | 0.604 |  | 2.61 |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.05 | 2069.581 | 81333 |  | 2069.581 | 4 | 6.16 | 53.9 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.06 | 66022.867 | 2635104 | 88157.313 | 2.996 |  | 49.4 |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.4 | 9416.714 | 361946 | 88157.313 | 0.427 |  | 0.94 |  | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.4 | 88157.313 | 3428768 |  | 88157.313 | 4 | 9.86 | 146.5
20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.88 | 3.841 | 109 | 3781.379 | 0.004 |  |  |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.78 | 306.603 | 9765 | 3781.379 | 0.324 |  | 3.19 |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.78 | 3781.379 | 130088 |  | 3781.379 | 4 | 4.68 | 16.9 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.81 | 301589.063 | 10239744 | 104880.836 | 11.502 |  | 32.21 |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.97 | 104880.836 | 3105835 |  | 104880.836 | 4 | 6.33 | 58.2 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.05 | 359.079 | 14109 | 2232.48 | 0.643 |  | 2.85 |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.05 | 2232.48 | 87859 |  | 2232.48 | 4 | 6.64 | 66 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.06 | 59553.82 | 2256071 | 98938.063 | 2.408 |  | 39.58 |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.4 | 10473.873 | 423184 | 98938.063 | 0.423 |  | 0.93 |  | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.4 | 98938.063 | 4006047 |  | 98938.063 | 4 | 11.07 | 176.7
20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte |  |  |  | 4173.079 |  |  |  |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.78 | 77.001 | 2803 | 4173.079 | 0.074 |  |  |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.78 | 4173.079 | 145040 |  | 4173.079 | 4 | 5.16 | 29 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.81 | 226021.344 | 7559059 | 110140.234 | 8.208 |  | 22.87 |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.96 | 110140.234 | 3188337 |  | 110140.234 | 4 | 6.64 | 66.1 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.05 | 86.69 | 3498 | 2183.897 | 0.159 |  |  |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.05 | 2183.897 | 87172 |  | 2183.897 | 4 | 6.5 | 62.4 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.06 | 83874.047 | 3258557 | 68976.289 | 4.864 |  | 80.6 |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.4 | 2407.742 | 103954 | 68976.289 | 0.14 |  |  |  | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.4 | 68976.289 | 2929824 |  | 68976.289 | 4 | 7.72 | 92.9
20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.89 | 7.057 | 258 | 3930.435 | 0.007 |  |  |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.78 | 41.344 | 1442 | 3930.435 | 0.042 |  |  |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.78 | 3930.435 | 136738 |  | 3930.435 | 4 | 4.86 | 21.5 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.81 | 278933.406 | 9405930 | 108602.188 | 10.274 |  | 28.73 |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.97 | 108602.188 | 3147145 |  | 108602.188 | 4 | 6.55 | 63.8 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.05 | 47.479 | 2095 | 2102.246 | 0.09 |  |  |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.05 | 2102.246 | 79293 |  | 2102.246 | 4 | 6.25 | 56.3 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.07 | 72234.156 | 2659945 | 80912.125 | 3.571 |  | 59 |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.4 | 1676.365 | 65562 | 80912.125 | 0.083 |  |  |  | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.4 | 80912.125 | 3194553 |  | 80912.125 | 4 | 9.05 | 126.3
20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte |  |  |  | 3689.984 |  |  |  |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.77 | 39.448 | 947 | 3689.984 | 0.043 |  |  |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.78 | 3689.984 | 126097 |  | 3689.984 | 4 | 4.56 | 14.1 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.81 | 276028.344 | 9283592 | 90797.094 | 12.16 |  | 34.08 |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.96 | 90797.094 | 2563937 |  | 90797.094 | 4 | 5.48 | 36.9 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.05 | 49.284 | 2089 | 2034.973 | 0.097 |  |  |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.05 | 2034.973 | 79528 |  | 2034.973 | 4 | 6.05 | 51.3 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.06 | 76960.344 | 2921204 | 64799.211 | 4.751 |  | 78.7 |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.4 | 1451.845 | 55029 | 64799.211 | 0.09 |  |  |  | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.4 | 64799.211 | 2501513 |  | 64799.211 | 4 | 7.25 | 81.2
20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.85 | 3.359 | 200 | 3804.575 | 0.004 |  |  |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.78 | 79.311 | 2437 | 3804.575 | 0.083 |  |  |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.77 | 3804.575 | 130640 |  | 3804.575 | 4 | 4.7 | 17.6 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.81 | 216459.453 | 6929789 | 118122.609 | 7.33 |  | 20.38 |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.97 | 118122.609 | 3489290 |  | 118122.609 | 4 | 7.13 | 78.1 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.05 | 63.795 | 2704 | 2285.38 | 0.112 |  |  |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.05 | 2285.38 | 88611 |  | 2285.38 | 4 | 6.8 | 70 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.06 | 82758.57 | 3205700 | 100054.281 | 3.309 |  | 54.62 |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.4 | 2658.295 | 114635 | 100054.281 | 0.106 |  |  |  | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.4 | 100054.281 | 4122043 |  | 100054.281 | 4 | 11.19 | 179.8
20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.85 | 0.614 | 56 | 3650.367 | 0.001 |  |  |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.78 | 63.602 | 1984 | 3650.367 | 0.07 |  |  |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.78 | 3650.367 | 127037 |  | 3650.367 | 4 | 4.51 | 12.8 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.81 | 338452.781 | 11647518 | 102091.367 | 13.261 |  | 37.2 |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.97 | 102091.367 | 2883097 |  | 102091.367 | 4 | 6.16 | 54 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.05 | 58.385 | 2570 | 2275.792 | 0.103 |  |  |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.05 | 2275.792 | 87791 |  | 2275.792 | 4 | 6.77 | 69.2 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.07 | 84048.07 | 3297586 | 91328.695 | 3.681 |  | 60.84 |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.4 | 2295.008 | 85333 | 91328.695 | 0.101 |  |  |  | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.4 | 91328.695 | 3601070 |  | 91328.695 | 4 | 10.22 | 155.4
20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.82 | 3.82 | 148 | 4028.423 | 0.004 |  |  |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.78 | 57.151 | 1930 | 4028.423 | 0.057 |  |  |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.78 | 4028.423 | 139652 |  | 4028.423 | 4 | 4.98 | 24.5 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.81 | 282383.156 | 9499879 | 125547.516 | 8.997 |  | 25.11 |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.97 | 125547.516 | 3727889 |  | 125547.516 | 4 | 7.57 | 89.3 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.05 | 60.019 | 2065 | 2367.659 | 0.101 |  |  |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.05 | 2367.659 | 92560 |  | 2367.659 | 4 | 7.04 | 76.1 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.07 | 90018.523 | 3535425 | 102183.391 | 3.524 |  | 58.22 |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.4 | 2382.871 | 91499 | 102183.391 | 0.093 |  |  |  | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.4 | 102183.391 | 3971211 |  | 102183.391 | 4 | 11.43 | 185.8
20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.9 | 4.366 | 159 | 3838.459 | 0.005 |  |  |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.78 | 74.386 | 1977 | 3838.459 | 0.078 |  |  |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.78 | 3838.459 | 133137 |  | 3838.459 | 4 | 4.75 | 18.6 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.81 | 367082 | 12520329 | 111401.727 | 13.18 |  | 36.98 |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.97 | 111401.727 | 3155190 |  | 111401.727 | 4 | 6.72 | 68 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.05 | 49.724 | 2144 | 2392.802 | 0.083 |  |  |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.05 | 2392.802 | 92804 |  | 2392.802 | 4 | 7.12 | 77.9 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.07 | 93146.758 | 3627092 | 95344.883 | 3.908 |  | 64.63 |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.4 | 2270.83 | 89051 | 95344.883 | 0.095 |  |  |  | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.4 | 95344.883 | 3645451 |  | 95344.883 | 4 | 10.67 | 166.6
20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.85 | 2.995 | 159 | 3972.163 | 0.003 |  |  |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.77 | 43.105 | 1522 | 3972.163 | 0.043 |  |  |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.78 | 3972.163 | 136332 |  | 3972.163 | 4 | 4.91 | 22.8 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.81 | 353533.938 | 11858321 | 110531.492 | 12.794 |  | 35.88 |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.96 | 110531.492 | 3088762 |  | 110531.492 | 4 | 6.67 | 66.7 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.05 | 54.339 | 2333 | 2360.479 | 0.092 |  |  |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.05 | 2360.479 | 94249 |  | 2360.479 | 4 | 7.02 | 75.5 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.07 | 77323.891 | 2759422 | 98936.914 | 3.126 |  | 51.58 |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.4 | 2288.021 | 85144 | 98936.914 | 0.093 |  |  |  | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.4 | 98936.914 | 3873731 |  | 98936.914 | 4 | 11.07 | 176.7
20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.9 | 0.645 | 36 | 3842.832 | 0.001 |  |  |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.78 | 58.642 | 2012 | 3842.832 | 0.061 |  |  |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.78 | 3842.832 | 133074 |  | 3842.832 | 4 | 4.75 | 18.8 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.81 | 323413.156 | 10911557 | 112650.75 | 11.484 |  | 32.16 |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.97 | 112650.75 | 3159743 |  | 112650.75 | 4 | 6.8 | 69.9 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.05 | 71.881 | 2863 | 2468 | 0.117 |  |  |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.05 | 2468 | 99081 |  | 2468 | 4 | 7.34 | 83.5 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.06 | 88055.625 | 3283423 | 96992.188 | 3.631 |  | 60.01 |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.4 | 2269.482 | 96527 | 96992.188 | 0.094 |  |  |  | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.4 | 96992.188 | 3737935 |  | 96992.188 | 4 | 10.85 | 171.3
20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.88 | 8.494 | 187 | 3605.647 | 0.009 |  |  |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.78 | 34.578 | 1251 | 3605.647 | 0.038 |  |  |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.78 | 3605.647 | 124797 |  | 3605.647 | 4 | 4.46 | 11.5 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.81 | 194233.328 | 6480172 | 96524.773 | 8.049 |  | 22.42 |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.97 | 96524.773 | 2796460 |  | 96524.773 | 4 | 5.82 | 45.6 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.05 | 43.619 | 1630 | 2073.876 | 0.084 |  |  |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.05 | 2073.876 | 84779 |  | 2073.876 | 4 | 6.17 | 54.2 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.07 | 68525.758 | 2553414 | 95108.883 | 2.882 |  | 47.5 |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.4 | 2309.284 | 84392 | 95108.883 | 0.097 |  |  |  | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.4 | 95108.883 | 3551584 |  | 95108.883 | 4 | 10.64 | 166
20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte |  |  |  | 3280.838 |  |  |  |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.78 | 55.165 | 1589 | 3280.838 | 0.067 |  |  |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.78 | 3280.838 | 113487 |  | 3280.838 | 4 | 4.06 | 1.4 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.81 | 237021.109 | 8063289 | 88051.555 | 10.767 |  | 30.13 |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.97 | 88051.555 | 2541873 |  | 88051.555 | 4 | 5.31 | 32.8 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.05 | 49.147 | 2150 | 1820.366 | 0.108 |  |  |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.05 | 1820.366 | 71864 |  | 1820.366 | 4 | 5.41 | 35.4 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.07 | 58218.891 | 2193878 | 53921.77 | 4.319 |  | 71.49 |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.4 | 1318.323 | 47264 | 53921.77 | 0.098 |  |  |  | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.4 | 53921.77 | 2013740 |  | 53921.77 | 4 | 6.03 | 50.8
20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte |  |  |  | 3179.154 |  |  |  |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.78 | 42.574 | 1442 | 3179.154 | 0.054 |  |  |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.78 | 3179.154 | 110380 |  | 3179.154 | 4 | 3.93 | -1.7 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.81 | 243039.016 | 8258455 | 96476.289 | 10.077 |  | 28.17 |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.97 | 96476.289 | 2806186 |  | 96476.289 | 4 | 5.82 | 45.5 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.05 | 49.617 | 1955 | 1823.648 | 0.109 |  |  |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.05 | 1823.648 | 69186 |  | 1823.648 | 4 | 5.42 | 35.6 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.07 | 65500.477 | 2518845 | 61000.227 | 4.295 |  | 71.1 |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.4 | 1408.505 | 61221 | 61000.227 | 0.092 |  |  |  | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.4 | 61000.227 | 2395311 |  | 61000.227 | 4 | 6.82 | 70.6
20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.87 | 0.965 | 116 | 3893.189 | 0.001 |  |  |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.77 | 87.68 | 2747 | 3893.189 | 0.09 |  |  |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.78 | 3893.189 | 135899 |  | 3893.189 | 4 | 4.81 | 20.3 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.81 | 345005.031 | 11796750 | 113099.43 | 12.202 |  | 34.2 |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.97 | 113099.43 | 3265188 |  | 113099.43 | 4 | 6.82 | 70.6 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.05 | 64.293 | 2544 | 2278.214 | 0.113 |  |  |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.05 | 2278.214 | 91468 |  | 2278.214 | 4 | 6.78 | 69.4 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.07 | 96175.375 | 3901113 | 108030.414 | 3.561 |  | 58.84 |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.4 | 2530.248 | 100776 | 108030.414 | 0.094 |  |  |  | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.4 | 108030.414 | 4106710 |  | 108030.414 | 4 | 12.08 | 202.1
20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.9 | 5.351 | 219 | 3559.379 | 0.006 |  |  |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.78 | 85.112 | 3077 | 3559.379 | 0.096 |  |  |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.78 | 3559.379 | 121498 |  | 3559.379 | 4 | 4.4 | 10 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.81 | 336113.375 | 11298503 | 91099.672 | 14.758 |  | 41.45 |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.96 | 91099.672 | 2589675 |  | 91099.672 | 4 | 5.5 | 37.4 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.05 | 73.579 | 3336 | 2180.041 | 0.135 |  |  |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.05 | 2180.041 | 83876 |  | 2180.041 | 4 | 6.48 | 62.1 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.06 | 85629.828 | 3304683 | 100591.141 | 3.405 |  | 56.23 |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.4 | 3093.332 | 115462 | 100591.141 | 0.123 |  |  |  | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.4 | 100591.141 | 3646289 |  | 100591.141 | 4 | 11.25 | 181.3
20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.84 | 2.901 | 108 | 4010.53 | 0.003 |  |  |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.78 | 56.909 | 2038 | 4010.53 | 0.057 |  |  |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.78 | 4010.53 | 139009 |  | 4010.53 | 4 | 4.96 | 24 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.81 | 369345.031 | 12524720 | 95262.484 | 15.509 |  | 43.58 |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.97 | 95262.484 | 2696060 |  | 95262.484 | 4 | 5.75 | 43.7 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.05 | 48.467 | 1868 | 2026.516 | 0.096 |  |  |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.05 | 2026.516 | 78913 |  | 2026.516 | 4 | 6.03 | 50.7 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.07 | 81606.148 | 3017306 | 69586.211 | 4.691 |  | 77.71 |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.4 | 1622.044 | 59654 | 69586.211 | 0.093 |  |  |  | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.4 | 69586.211 | 2536447 |  | 69586.211 | 4 | 7.78 | 94.6
20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.87 | 6.052 | 191 | 4683.644 | 0.005 |  |  |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.78 | 6827.96 | 236140 | 4683.644 | 5.831 |  | 84.57 |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.78 | 4683.644 | 161988 |  | 4683.644 | 4 | 5.79 | 44.8 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.81 | 278402.094 | 9377985 | 95167.969 | 11.702 |  | 32.78 |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.96 | 95167.969 | 2738875 |  | 95167.969 | 4 | 5.74 | 43.5 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.05 | 4290.784 | 170039 | 1951.082 | 8.797 |  | 51.77 |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.05 | 1951.082 | 75634 |  | 1951.082 | 4 | 5.8 | 45.1 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.06 | 87272.586 | 3324550 | 104467.281 | 3.342 |  | 55.17 |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.4 | 121838.508 | 4179704 | 104467.281 | 4.665 |  | 15.65 |  | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.4 | 104467.281 | 3438160 |  | 104467.281 | 4 | 11.69 | 192.2
20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte |  |  |  | 3473.603 |  |  |  |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.78 | 217.342 | 7810 | 3473.603 | 0.25 |  | 2.1 |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.78 | 3473.603 | 119313 |  | 3473.603 | 4 | 4.29 | 7.4 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.81 | 253162.359 | 8413451 | 101755.492 | 9.952 |  | 27.82 |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.97 | 101755.492 | 2993356 |  | 101755.492 | 4 | 6.14 | 53.5 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.05 | 187.424 | 7368 | 1931.021 | 0.388 |  | 1.32 |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.05 | 1931.021 | 76802 |  | 1931.021 | 4 | 5.74 | 43.6 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.06 | 83227.961 | 3148725 | 48669.789 | 6.84 |  | 113.6 |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.4 | 4144.071 | 183465 | 48669.789 | 0.341 |  | 0.64 |  | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.4 | 48669.789 | 2086193 |  | 48669.789 | 4 | 5.44 | 36.1
20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte |  |  |  | 3760.475 |  |  |  |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.78 | 2704.803 | 93865 | 3760.475 | 2.877 |  | 40.91 |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.78 | 3760.475 | 129538 |  | 3760.475 | 4 | 4.65 | 16.2 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.81 | 245721.672 | 8129133 | 87558.133 | 11.226 |  | 31.43 |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.97 | 87558.133 | 2520853 |  | 87558.133 | 4 | 5.28 | 32.1 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.05 | 1550.709 | 60551 | 1544.025 | 4.017 |  | 23.09 |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.05 | 1544.025 | 60954 |  | 1544.025 | 4 | 4.59 | 14.8 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.06 | 73634.305 | 2678980 | 39606.652 | 7.437 |  | 123.56 |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.4 | 34674.074 | 1378169 | 39606.652 | 3.502 |  | 11.61 |  | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.4 | 39606.652 | 1617206 |  | 39606.652 | 4 | 4.43 | 10.8
20231106_HepCl_ReRun_041 | QC A | QC | 3.88 | 165.842 | 3159 | 3334.889 | 0.199 | 0.625 |  |  | 20231106_HepCl_ReRun_041 | QC A | QC | 3.78 | 94.863 | 3071 | 3334.889 | 0.114 | 0.625 | 0.08 | -87.4 | 20231106_HepCl_ReRun_041 | QC A | QC | 3.78 | 3334.889 | 113286 |  | 3334.889 | 4 | 4.12 | 3.1 | 20231106_HepCl_ReRun_041 | QC A | QC | 3.81 | 5815.193 | 189001 | 87338.508 | 0.266 | 0.625 | 0.34 | -45.2 | 20231106_HepCl_ReRun_041 | QC A | QC | 3.96 | 87338.508 | 2479133 |  | 87338.508 | 4 | 5.27 | 31.7 | 20231106_HepCl_ReRun_041 | QC A | QC | 4.05 | 83.107 | 3237 | 1704.962 | 0.195 | 0.625 | 0.16 | -74.1 | 20231106_HepCl_ReRun_041 | QC A | QC | 4.05 | 1704.962 | 62062 |  | 1704.962 | 4 | 5.07 | 26.8 | 20231106_HepCl_ReRun_041 | QC A | QC | 4.07 | 508.329 | 19978 | 49355.238 | 0.041 | 0.625 | 0.06 | -90.8 | 20231106_HepCl_ReRun_041 | QC A | QC | 4.4 | 3405.577 | 139548 | 49355.238 | 0.276 | 0.625 | 0.41 | -33.8 | 20231106_HepCl_ReRun_041 | QC A | QC | 4.4 | 49355.238 | 1984290 |  | 49355.238 | 4 | 5.52 | 38
20231106_HepCl_ReRun_080 | QC A | QC | 3.87 | 234.077 | 3816 | 3603.385 | 0.26 | 0.625 | 0.22 | -64.9 | 20231106_HepCl_ReRun_080 | QC A | QC | 3.78 | 122.98 | 3746 | 3603.385 | 0.137 | 0.625 | 0.41 | -33.6 | 20231106_HepCl_ReRun_080 | QC A | QC | 3.78 | 3603.385 | 124910 |  | 3603.385 | 4 | 4.46 | 11.4 | 20231106_HepCl_ReRun_080 | QC A | QC | 3.81 | 4988.641 | 158149 | 88138.531 | 0.226 | 0.625 | 0.23 | -63.3 | 20231106_HepCl_ReRun_080 | QC A | QC | 3.96 | 88138.531 | 2488217 |  | 88138.531 | 4 | 5.32 | 32.9 | 20231106_HepCl_ReRun_080 | QC A | QC | 4.05 | 74.95 | 2549 | 1712.896 | 0.175 | 0.625 | 0.04 | -93.2 | 20231106_HepCl_ReRun_080 | QC A | QC | 4.05 | 1712.896 | 62704 |  | 1712.896 | 4 | 5.1 | 27.4 | 20231106_HepCl_ReRun_080 | QC A | QC | 4.07 | 502.869 | 19829 | 36659.156 | 0.055 | 0.625 | 0.29 | -54.3 | 20231106_HepCl_ReRun_080 | QC A | QC | 4.4 | 2607.647 | 110211 | 36659.156 | 0.285 | 0.625 | 0.44 | -29 | 20231106_HepCl_ReRun_080 | QC A | QC | 4.4 | 36659.156 | 1533681 |  | 36659.156 | 4 | 4.1 | 2.5
20231106_HepCl_ReRun_042 | QC B | QC | 3.87 | 710.963 | 12145 | 3127.347 | 0.909 | 2.506 | 2.59 | 3.5 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.78 | 191.108 | 5682 | 3127.347 | 0.244 | 2.506 | 2.01 | -19.8 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.78 | 3127.347 | 113458 |  | 3127.347 | 4 | 3.87 | -3.3 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.81 | 22068.494 | 733400 | 76308.461 | 1.157 | 2.506 | 2.87 | 14.5 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.96 | 76308.461 | 2025701 |  | 76308.461 | 4 | 4.6 | 15.1 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.05 | 231.524 | 9385 | 1697.607 | 0.546 | 2.506 | 2.27 | -9.6 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.05 | 1697.607 | 63218 |  | 1697.607 | 4 | 5.05 | 26.2 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.06 | 1949.468 | 71825 | 51642.734 | 0.151 | 2.506 | 1.89 | -24.6 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.4 | 10822.062 | 445443 | 51642.734 | 0.838 | 2.506 | 2.37 | -5.6 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.4 | 51642.734 | 2142571 |  | 51642.734 | 4 | 5.78 | 44.4
20231106_HepCl_ReRun_081 | QC B | QC | 3.87 | 752.553 | 12162 | 2887.97 | 1.042 | 2.506 | 3.08 | 22.9 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.78 | 213.334 | 7274 | 2887.97 | 0.295 | 2.506 | 2.76 | 10.3 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.77 | 2887.97 | 97800 |  | 2887.97 | 4 | 3.57 | -10.7 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.8 | 19008.291 | 624404 | 84922.633 | 0.895 | 2.506 | 2.13 | -15.1 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.96 | 84922.633 | 2354939 |  | 84922.633 | 4 | 5.12 | 28.1 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.05 | 222.404 | 8528 | 1702.264 | 0.523 | 2.506 | 2.13 | -15.1 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.05 | 1702.264 | 67418 |  | 1702.264 | 4 | 5.06 | 26.6 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.06 | 2080.625 | 81819 | 37154.508 | 0.224 | 2.506 | 3.11 | 24.1 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.4 | 7522.85 | 298162 | 37154.508 | 0.81 | 2.506 | 2.27 | -9.5 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.4 | 37154.508 | 1560233 |  | 37154.508 | 4 | 4.16 | 3.9
20231106_HepCl_ReRun_043 | QC C | QC | 3.88 | 1686.431 | 30672 | 3371.759 | 2.001 | 6.269 | 6.58 | 5 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.78 | 449.471 | 14963 | 3371.759 | 0.533 | 6.269 | 6.28 | 0.1 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.78 | 3371.759 | 115251 |  | 3371.759 | 4 | 4.17 | 4.2 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.81 | 51954.094 | 1766232 | 84343.641 | 2.464 | 6.269 | 6.58 | 4.9 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.96 | 84343.641 | 2347158 |  | 84343.641 | 4 | 5.09 | 27.2 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.05 | 564.021 | 21717 | 1749.339 | 1.29 | 6.269 | 6.73 | 7.4 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.05 | 1749.339 | 68129 |  | 1749.339 | 4 | 5.2 | 30.1 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.06 | 5233.741 | 199717 | 53094.781 | 0.394 | 6.269 | 5.95 | -5 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.4 | 26934.248 | 1136498 | 53094.781 | 2.029 | 6.269 | 6.5 | 3.7 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.4 | 53094.781 | 2272471 |  | 53094.781 | 4 | 5.94 | 48.5
20231106_HepCl_ReRun_082 | QC C | QC | 3.88 | 2095.343 | 37365 | 3610.88 | 2.321 | 6.269 | 7.76 | 23.7 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.78 | 390.65 | 13088 | 3610.88 | 0.433 | 6.269 | 4.79 | -23.5 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.78 | 3610.88 | 124330 |  | 3610.88 | 4 | 4.46 | 11.6 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.81 | 56246.949 | 1905280 | 79141.828 | 2.843 | 6.269 | 7.65 | 22.1 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.96 | 79141.828 | 2191526 |  | 79141.828 | 4 | 4.77 | 19.4 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.05 | 532.346 | 19977 | 1712.841 | 1.243 | 6.269 | 6.45 | 2.9 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.05 | 1712.841 | 65873 |  | 1712.841 | 4 | 5.09 | 27.4 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.06 | 4570.148 | 205965 | 41341.777 | 0.442 | 6.269 | 6.75 | 7.7 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.4 | 21356.891 | 869820 | 41341.777 | 2.066 | 6.269 | 6.63 | 5.7 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.4 | 41341.777 | 1699979 |  | 41341.777 | 4 | 4.62 | 15.6
20231106_HepCl_ReRun_044 | QC D | QC | 3.88 | 6550.212 | 126329 | 3225.131 | 8.124 | 25 | 28.97 | 15.9 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.78 | 1519.89 | 52572 | 3225.131 | 1.885 | 25 | 26.25 | 5 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.78 | 3225.131 | 110928 |  | 3225.131 | 4 | 3.99 | -0.3 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.81 | 182690.922 | 6144534 | 70742.273 | 10.33 | 25 | 28.89 | 15.6 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.96 | 70742.273 | 1886520 |  | 70742.273 | 4 | 4.27 | 6.7 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.05 | 1745.07 | 65258 | 1516.869 | 4.602 | 25 | 26.6 | 6.4 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.05 | 1516.869 | 56991 |  | 1516.869 | 4 | 4.51 | 12.8 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.07 | 18617.66 | 686901 | 46975.297 | 1.585 | 25 | 25.84 | 3.4 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.4 | 85835.969 | 3549459 | 46975.297 | 7.309 | 25 | 24.83 | -0.7 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.4 | 46975.297 | 1959518 |  | 46975.297 | 4 | 5.25 | 31.4
20231106_HepCl_ReRun_083 | QC D | QC | 3.88 | 7293.705 | 119126 | 3571.287 | 8.169 | 25 | 29.14 | 16.5 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.78 | 1626.875 | 56148 | 3571.287 | 1.822 | 25 | 25.32 | 1.3 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.78 | 3571.287 | 123058 |  | 3571.287 | 4 | 4.42 | 10.4 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.8 | 186631.609 | 6350371 | 68663.195 | 10.872 | 25 | 30.43 | 21.7 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.96 | 68663.195 | 1862578 |  | 68663.195 | 4 | 4.14 | 3.6 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.05 | 1766.013 | 67718 | 1490.915 | 4.738 | 25 | 27.42 | 9.7 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.05 | 1490.915 | 57447 |  | 1490.915 | 4 | 4.43 | 10.9 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.06 | 15100.404 | 616116 | 32931.387 | 1.834 | 25 | 30 | 20 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.4 | 60284.25 | 2577048 | 32931.387 | 7.322 | 25 | 24.87 | -0.5 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.4 | 32931.387 | 1374550 |  | 32931.387 | 4 | 3.68 | -7.9
20231106_HepCl_ReRun_001 | Blank | Blank | 3.86 | 7.841 | 320 | [5 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | 3.82 | 149.304 | 2967 | [5 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_001 | Blank | Blank | 3.86 | 1311.774 | 28378 | [5 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_001 | Blank | Blank |  |  |  | 2.685 |  |  |  |  | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.07 | 2.685 | 157 |  | 2.685 | 4 | 0.01 | -99.8 | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.04 | 5.602 | 132 | [5 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.42 | 18.895 | 1390 | [5 empty] | 20231106_HepCl_ReRun_001 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_002 | Blank | Blank | 3.88 | 1.966 | 103 | 2.377 | 3.308 |  | 11.37 |  | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.81 | 70.268 | 1184 | 2.377 | 118.247 |  | 1745.78 |  | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.79 | 2.377 | 87 |  | 2.377 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.86 | 881.125 | 16293 | [5 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.06 | 1.228 | 71 | 3.534 | 1.39 |  | 7.33 |  | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.04 | 3.534 | 118 |  | 3.534 | 4 | 0.01 | -99.7 | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.05 | 4.978 | 128 | [5 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.41 | 62.547 | 1907 | [5 empty] | 20231106_HepCl_ReRun_002 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_003 | Blank | Blank |  |  |  | 0.237 |  |  |  |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.83 | 51.466 | 878 | 0.237 | 868.624 |  | 12834.47 |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.81 | 0.237 | 33 |  | 0.237 | 4 | 0 | -100 | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.85 | 721.975 | 14600 | 5.359 | 538.888 |  | 1528.2 |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.95 | 5.359 | 177 |  | 5.359 | 4 | 0 | -100 | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.04 | 4.574 | 189 | [5 empty] | 20231106_HepCl_ReRun_003 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.07 | 0.235 | 18 | 6.982 | 0.135 |  | 1.62 |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.39 | 1.965 | 356 | 6.982 | 1.126 |  | 3.36 |  | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.39 | 6.982 | 235 |  | 6.982 | 4 | 0 | -100
20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 8.596 |  |  |  |  | 20231106_HepCl_ReRun_006 | Blank | Blank | 3.81 | 46.309 | 811 | 8.596 | 21.549 |  | 316.84 |  | 20231106_HepCl_ReRun_006 | Blank | Blank | 3.82 | 8.596 | 316 |  | 8.596 | 4 | 0.01 | -99.7 | 20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 13.111 |  |  |  |  | 20231106_HepCl_ReRun_006 | Blank | Blank | 3.96 | 13.111 | 178 |  | 13.111 | 4 | 0 | -100 | 20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 0.473 |  |  |  |  | 20231106_HepCl_ReRun_006 | Blank | Blank | 4.04 | 0.473 | 28 |  | 0.473 | 4 | 0 | -100 | 20231106_HepCl_ReRun_006 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_006 | Blank | Blank | 4.39 | 12.852 | 1235 | [5 empty] | 20231106_HepCl_ReRun_006 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_007 | Blank | Blank | 3.88 | 6.992 | 172 | [5 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | 3.81 | 31.905 | 854 | [5 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_007 | Blank | Blank | 3.85 | 443.105 | 9101 | [5 empty] | 20231106_HepCl_ReRun_007 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.05 | 0.796 | 47 | 1.479 | 2.153 |  | 11.91 |  | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.09 | 1.479 | 80 |  | 1.479 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.03 | 2.4 | 141 | 7.306 | 1.314 |  | 21.31 |  | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.35 | 36.801 | 1623 | 7.306 | 20.148 |  | 69.4 |  | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.4 | 7.306 | 320 |  | 7.306 | 4 | 0 | -100
20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.04 |  |  |  |  | 20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.04 |  |  |  |  | 20231106_HepCl_ReRun_014 | Blank | Blank | 3.74 | 3.04 | 110 |  | 3.04 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.806 |  |  |  |  | 20231106_HepCl_ReRun_014 | Blank | Blank | 3.96 | 3.806 | 133 |  | 3.806 | 4 | 0 | -100 | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.06 | 0.207 | 12 | 0.151 | 5.483 |  | 31.89 |  | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.01 | 0.151 | 15 |  | 0.151 | 4 | 0 | -100 | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.05 | 2.265 | 131 | 0.639 | 14.178 |  | 236.14 |  | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.42 | 1.451 | 372 | 0.639 | 9.083 |  | 30.99 |  | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.37 | 0.639 | 71 |  | 0.639 | 4 | 0 | -100
20231106_HepCl_ReRun_021 | Blank | Blank | 3.9 | 10.663 | 141 | 2.584 | 16.506 |  | 59.62 |  | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.82 | 11.18 | 423 | 2.584 | 17.307 |  | 254.14 |  | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.82 | 2.584 | 93 |  | 2.584 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.85 | 265.891 | 4687 | 7.869 | 135.159 |  | 382.98 |  | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.99 | 7.869 | 293 |  | 7.869 | 4 | 0 | -100 | 20231106_HepCl_ReRun_021 | Blank | Blank |  |  |  | 3.142 |  |  |  |  | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.01 | 3.142 | 118 |  | 3.142 | 4 | 0.01 | -99.8 | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.09 | 6.331 | 205 | [5 empty] | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.43 | 18.559 | 1477 | [5 empty] | 20231106_HepCl_ReRun_021 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_032 | Blank | Blank | 3.87 | 5.334 | 155 | 0.567 | 37.63 |  | 136.84 |  | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.82 | 4.34 | 147 | 0.567 | 30.617 |  | 450.84 |  | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.8 | 0.567 | 13 |  | 0.567 | 4 | 0 | -100 | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.85 | 231.762 | 4646 | [5 empty] | 20231106_HepCl_ReRun_032 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 0.372 |  |  |  |  | 20231106_HepCl_ReRun_032 | Blank | Blank | 4.01 | 0.372 | 22 |  | 0.372 | 4 | 0 | -100 | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 4.114 |  |  |  |  | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 4.114 |  |  |  |  | 20231106_HepCl_ReRun_032 | Blank | Blank | 4.35 | 4.114 | 195 |  | 4.114 | 4 | 0 | -100
20231106_HepCl_ReRun_036 | Blank | Blank | 3.92 | 10.411 | 310 | [5 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank | 3.82 | 14.094 | 247 | [5 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_036 | Blank | Blank | 3.85 | 224.475 | 4100 | [5 empty] | 20231106_HepCl_ReRun_036 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_036 | Blank | Blank | 4.06 | 1.437 | 84 | 0.326 | 17.632 |  | 104.77 |  | 20231106_HepCl_ReRun_036 | Blank | Blank | 4.07 | 0.326 | 19 |  | 0.326 | 4 | 0 | -100 | 20231106_HepCl_ReRun_036 | Blank | Blank |  |  |  | 7.725 |  |  |  |  | 20231106_HepCl_ReRun_036 | Blank | Blank |  |  |  | 7.725 |  |  |  |  | 20231106_HepCl_ReRun_036 | Blank | Blank | 4.41 | 7.725 | 173 |  | 7.725 | 4 | 0 | -100
20231106_HepCl_ReRun_040 | Blank | Blank |  |  |  | 2.239 |  |  |  |  | 20231106_HepCl_ReRun_040 | Blank | Blank |  |  |  | 2.239 |  |  |  |  | 20231106_HepCl_ReRun_040 | Blank | Blank | 3.75 | 2.239 | 81 |  | 2.239 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_040 | Blank | Blank | 3.85 | 223.767 | 3692 | [5 empty] | 20231106_HepCl_ReRun_040 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.05 | 0.854 | 50 | [5 empty] | 20231106_HepCl_ReRun_040 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.07 | 5.768 | 153 | 5.681 | 4.061 |  | 67.19 |  | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.44 | 65.969 | 2411 | 5.681 | 46.449 |  | 160.69 |  | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.4 | 5.681 | 201 |  | 5.681 | 4 | 0 | -100
20231106_HepCl_ReRun_053 | Blank | Blank | 3.9 | 1.849 | 69 | 1.263 | 5.856 |  | 20.68 |  | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 1.263 |  |  |  |  | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.79 | 1.263 | 40 |  | 1.263 | 4 | 0 | -100 | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.85 | 234.037 | 4499 | 12.536 | 74.677 |  | 211.42 |  | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.97 | 12.536 | 315 |  | 12.536 | 4 | 0 | -100 | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 0.376 |  |  |  |  | 20231106_HepCl_ReRun_053 | Blank | Blank | 4.09 | 0.376 | 22 |  | 0.376 | 4 | 0 | -100 | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 3.035 |  |  |  |  | 20231106_HepCl_ReRun_053 | Blank | Blank | 4.43 | 2.767 | 722 | 3.035 | 3.647 |  | 12.11 |  | 20231106_HepCl_ReRun_053 | Blank | Blank | 4.39 | 3.035 | 208 |  | 3.035 | 4 | 0 | -100
20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 1.55 |  |  |  |  | 20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 1.55 |  |  |  |  | 20231106_HepCl_ReRun_060 | Blank | Blank | 3.76 | 1.55 | 56 |  | 1.55 | 4 | 0 | -100 | 20231106_HepCl_ReRun_060 | Blank | Blank | 3.85 | 243.211 | 4301 | [5 empty] | 20231106_HepCl_ReRun_060 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 0.214 |  |  |  |  | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.05 | 0.214 | 13 |  | 0.214 | 4 | 0 | -100 | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.06 | 11.194 | 371 | 2.979 | 15.031 |  | 250.38 |  | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.43 | 61.475 | 1351 | 2.979 | 82.544 |  | 285.99 |  | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.41 | 2.979 | 184 |  | 2.979 | 4 | 0 | -100
20231106_HepCl_ReRun_071 | Blank | Blank |  |  |  | 5.703 |  |  |  |  | 20231106_HepCl_ReRun_071 | Blank | Blank |  |  |  | 5.703 |  |  |  |  | 20231106_HepCl_ReRun_071 | Blank | Blank | 3.79 | 5.703 | 133 |  | 5.703 | 4 | 0.01 | -99.8 | 20231106_HepCl_ReRun_071 | Blank | Blank | 3.86 | 178.541 | 3006 | 7.057 | 101.199 |  | 286.65 |  | 20231106_HepCl_ReRun_071 | Blank | Blank | 3.97 | 7.057 | 248 |  | 7.057 | 4 | 0 | -100 | 20231106_HepCl_ReRun_071 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_071 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.04 | 0.246 | 35 | 13.529 | 0.073 |  | 0.58 |  | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.45 | 47.228 | 2391 | 13.529 | 13.963 |  | 47.93 |  | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.37 | 13.529 | 246 |  | 13.529 | 4 | 0 | -100
20231106_HepCl_ReRun_075 | Blank | Blank |  |  |  | 0.322 |  |  |  |  | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.79 | 0.88 | 95 | 0.322 | 10.932 |  | 159.94 |  | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.81 | 0.322 | 12 |  | 0.322 | 4 | 0 | -100 | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.86 | 219.784 | 4284 | 12.113 | 72.578 |  | 205.46 |  | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.97 | 12.113 | 359 |  | 12.113 | 4 | 0 | -100 | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.06 | 0.203 | 12 | [5 empty] | 20231106_HepCl_ReRun_075 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_075 | Blank | Blank |  |  |  | 8.646 |  |  |  |  | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.38 | 69.248 | 2207 | 8.646 | 32.037 |  | 110.66 |  | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.4 | 8.646 | 340 |  | 8.646 | 4 | 0 | -100
20231106_HepCl_ReRun_079 | Blank | Blank | 3.83 | 0.96 | 30 | 4.597 | 0.835 |  | 2.32 |  | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 4.597 |  |  |  |  | 20231106_HepCl_ReRun_079 | Blank | Blank | 3.81 | 4.597 | 132 |  | 4.597 | 4 | 0.01 | -99.9 | 20231106_HepCl_ReRun_079 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_079 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 0.232 |  |  |  |  | 20231106_HepCl_ReRun_079 | Blank | Blank | 4.05 | 0.232 | 36 |  | 0.232 | 4 | 0 | -100 | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 10.156 |  |  |  |  | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 10.156 |  |  |  |  | 20231106_HepCl_ReRun_079 | Blank | Blank | 4.36 | 10.156 | 202 |  | 10.156 | 4 | 0 | -100
20231106_HepCl_ReRun_084 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.83 | 20.354 | 553 | [5 empty] | 20231106_HepCl_ReRun_084 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.86 | 108.338 | 2639 | 2.32 | 186.79 |  | 529.44 |  | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.93 | 2.32 | 130 |  | 2.32 | 4 | 0 | -100 | 20231106_HepCl_ReRun_084 | Blank | Blank |  |  |  | 0.156 |  |  |  |  | 20231106_HepCl_ReRun_084 | Blank | Blank | 4.05 | 0.156 | 9 |  | 0.156 | 4 | 0 | -100 | 20231106_HepCl_ReRun_084 | Blank | Blank | 4.06 | 4.429 | 252 | 4.803 | 3.689 |  | 60.97 |  | 20231106_HepCl_ReRun_084 | Blank | Blank |  |  |  | 4.803 |  |  |  |  | 20231106_HepCl_ReRun_084 | Blank | Blank | 4.4 | 4.803 | 181 |  | 4.803 | 4 | 0 | -100
20231106_HepCl_ReRun_085 | Blank | Blank | 3.86 | 1.076 | 86 | [5 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | 3.83 | 1.805 | 126 | [5 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_085 | Blank | Blank | 3.86 | 243.378 | 5428 | [5 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_085 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_085 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.06 | 5.79 | 191 | 0.242 | 95.702 |  | 1597.58 |  | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.36 | 4.824 | 525 | 0.242 | 79.736 |  | 276.24 |  | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.42 | 0.242 | 58 |  | 0.242 | 4 | 0 | -100
20231106_HepCl_ReRun_087 | Blank | Blank | 3.89 | 6.497 | 168 | [5 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | 3.82 | 12.764 | 267 | [5 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_087 | Blank | Blank |  |  |  | 3.535 |  |  |  |  | 20231106_HepCl_ReRun_087 | Blank | Blank | 3.91 | 3.535 | 119 |  | 3.535 | 4 | 0 | -100 | 20231106_HepCl_ReRun_087 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_087 | Blank | Blank | 4.09 | 9.932 | 215 | [5 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | 4.36 | 22.948 | 1505 | [5 empty] | 20231106_HepCl_ReRun_087 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 1.329 |  |  |  |  | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.83 | 17.262 | 362 | 1.329 | 51.955 |  | 766.16 |  | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.81 | 1.329 | 48 |  | 1.329 | 4 | 0 | -100 | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.86 | 135.857 | 3418 | 4.894 | 111.04 |  | 314.56 |  | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.97 | 4.894 | 185 |  | 4.894 | 4 | 0 | -100 | 20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 1.338 |  |  |  |  | 20231106_HepCl_ReRun_094 | Blank | Blank | 4.1 | 1.338 | 61 |  | 1.338 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_094 | Blank | Blank | 4.08 | 2.282 | 84 | 0.119 | 76.706 |  | 1280.34 |  | 20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 0.119 |  |  |  |  | 20231106_HepCl_ReRun_094 | Blank | Blank | 4.4 | 0.119 | 22 |  | 0.119 | 4 | 0 | -100
20231106_HepCl_ReRun_101 | Blank | Blank | 3.83 | 17.593 | 249 | [5 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | 3.82 | 6.098 | 182 | [5 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_101 | Blank | Blank | 3.85 | 166.996 | 3172 | [5 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.06 | 1.23 | 72 | [5 empty] | 20231106_HepCl_ReRun_101 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.04 | 13.821 | 319 | 1.22 | 45.315 |  | 756.11 |  | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.42 | 49.545 | 2242 | 1.22 | 162.443 |  | 563.34 |  | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.39 | 1.22 | 126 |  | 1.22 | 4 | 0 | -100
20231106_HepCl_ReRun_108 | Blank | Blank | 3.85 | 6.219 | 242 | [5 empty] | 20231106_HepCl_ReRun_108 | Blank | Blank | 3.82 | 2.795 | 99 | [5 empty] | 20231106_HepCl_ReRun_108 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_108 | Blank | Blank |  |  |  | 3.622 |  |  |  |  | 20231106_HepCl_ReRun_108 | Blank | Blank | 3.97 | 3.622 | 139 |  | 3.622 | 4 | 0 | -100 | 20231106_HepCl_ReRun_108 | Blank | Blank |  |  |  | 0.289 |  |  |  |  | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.1 | 0.289 | 13 |  | 0.289 | 4 | 0 | -100 | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.07 | 2.447 | 167 | 7.197 | 1.36 |  | 22.08 |  | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.44 | 120.509 | 2214 | 7.197 | 66.977 |  | 231.95 |  | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.41 | 7.197 | 270 |  | 7.197 | 4 | 0 | -100
20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 3.311 |  |  |  |  | 20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 3.311 |  |  |  |  | 20231106_HepCl_ReRun_110 | Blank | Blank | 3.81 | 3.311 | 119 |  | 3.311 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_110 | Blank | Blank | 3.85 | 202.452 | 3781 | 1.239 | 653.598 |  | 1853.58 |  | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.01 | 1.239 | 97 |  | 1.239 | 4 | 0 | -100 | 20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 0.254 |  |  |  |  | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.04 | 0.254 | 45 |  | 0.254 | 4 | 0 | -100 | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.06 | 3.336 | 134 | 1.133 | 11.778 |  | 196.05 |  | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.42 | 11.766 | 1260 | 1.133 | 41.539 |  | 143.65 |  | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.38 | 1.133 | 113 |  | 1.133 | 4 | 0 | -100
20231106_HepCl_ReRun_117 | Blank | Blank | 3.91 | 20.044 | 377 | [5 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_117 | Blank | Blank | 3.86 | 219.288 | 4057 | [5 empty] | 20231106_HepCl_ReRun_117 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_117 | Blank | Blank |  |  |  | 0.204 |  |  |  |  | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.06 | 0.204 | 16 |  | 0.204 | 4 | 0 | -100 | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.02 | 3.587 | 216 | 1.785 | 8.038 |  | 133.6 |  | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.4 | 73.177 | 2401 | 1.785 | 163.982 |  | 568.68 |  | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.35 | 1.785 | 169 |  | 1.785 | 4 | 0 | -100
20231106_HepCl_ReRun_124 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank | 3.82 | 10.597 | 266 | [5 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_124 | Blank | Blank | 3.85 | 189.908 | 3607 | [5 empty] | 20231106_HepCl_ReRun_124 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_124 | Blank | Blank |  |  |  | 0.326 |  |  |  |  | 20231106_HepCl_ReRun_124 | Blank | Blank | 4.09 | 0.326 | 19 |  | 0.326 | 4 | 0 | -100 | 20231106_HepCl_ReRun_124 | Blank | Blank | 4.06 | 3.813 | 215 | 8.867 | 1.72 |  | 28.09 |  | 20231106_HepCl_ReRun_124 | Blank | Blank |  |  |  | 8.867 |  |  |  |  | 20231106_HepCl_ReRun_124 | Blank | Blank | 4.39 | 8.867 | 284 |  | 8.867 | 4 | 0 | -100
20231106_HepCl_ReRun_131 | Blank | Blank |  |  |  | 1.14 |  |  |  |  | 20231106_HepCl_ReRun_131 | Blank | Blank |  |  |  | 1.14 |  |  |  |  | 20231106_HepCl_ReRun_131 | Blank | Blank | 3.8 | 1.14 | 41 |  | 1.14 | 4 | 0 | -100 | 20231106_HepCl_ReRun_131 | Blank | Blank | 3.85 | 214.173 | 4079 | 14.521 | 58.997 |  | 166.94 |  | 20231106_HepCl_ReRun_131 | Blank | Blank | 3.92 | 14.521 | 393 |  | 14.521 | 4 | 0 | -100 | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.05 | 0.165 | 30 | 0.336 | 1.964 |  | 10.78 |  | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.09 | 0.336 | 18 |  | 0.336 | 4 | 0 | -100 | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.06 | 0.554 | 114 | 0.21 | 10.552 |  | 175.59 |  | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.43 | 5.209 | 616 | 0.21 | 99.219 |  | 343.87 |  | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.44 | 0.21 | 68 |  | 0.21 | 4 | 0 | -100
20231106_HepCl_ReRun_133 | Blank | Blank | 3.82 | 1.416 | 98 | 0.164 | 34.537 |  | 125.54 |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.81 | 10.285 | 369 | 0.164 | 250.854 |  | 3705.38 |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.81 | 0.164 | 6 |  | 0.164 | 4 | 0 | -100 | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.85 | 208.384 | 3838 | 7.49 | 111.287 |  | 315.26 |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.97 | 7.49 | 223 |  | 7.49 | 4 | 0 | -100 | 20231106_HepCl_ReRun_133 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_133 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.09 | 2.798 | 175 | 1.004 | 11.147 |  | 185.53 |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.39 | 20.099 | 1631 | 1.004 | 80.076 |  | 277.42 |  | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.38 | 1.004 | 93 |  | 1.004 | 4 | 0 | -100
20231106_HepCl_ReRun_140 | Blank | Blank | 3.88 | 3.301 | 150 | [5 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | 3.81 | 26.678 | 650 | [5 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_140 | Blank | Blank | 3.85 | 363.167 | 8123 | [5 empty] | 20231106_HepCl_ReRun_140 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_140 | Blank | Blank |  |  |  | 0.47 |  |  |  |  | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.06 | 0.47 | 28 |  | 0.47 | 4 | 0 | -100 | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.06 | 0.123 | 16 | 3.553 | 0.138 |  | 1.68 |  | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.43 | 18.612 | 2037 | 3.553 | 20.954 |  | 72.19 |  | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.4 | 3.553 | 153 |  | 3.553 | 4 | 0 | -100
20231106_HepCl_ReRun_142 | Blank | Blank |  |  |  | 2.523 |  |  |  |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.81 | 20.394 | 384 | 2.523 | 32.333 |  | 476.2 |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.72 | 2.523 | 91 |  | 2.523 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.86 | 295.344 | 5078 | [5 empty] | 20231106_HepCl_ReRun_142 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.02 | 0.165 | 18 | 0.21 | 3.143 |  | 17.85 |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.09 | 0.21 | 11 |  | 0.21 | 4 | 0 | -100 | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.1 | 3.614 | 178 | 4.462 | 3.24 |  | 53.47 |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.41 | 27.864 | 1977 | 4.462 | 24.979 |  | 86.16 |  | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.39 | 4.462 | 185 |  | 4.462 | 4 | 0 | -100
20231106_HepCl_ReRun_143 | Blank | Blank |  |  |  | 5.162 |  |  |  |  | 20231106_HepCl_ReRun_143 | Blank | Blank |  |  |  | 5.162 |  |  |  |  | 20231106_HepCl_ReRun_143 | Blank | Blank | 3.81 | 5.162 | 216 |  | 5.162 | 4 | 0.01 | -99.8 | 20231106_HepCl_ReRun_143 | Blank | Blank | 3.85 | 306.383 | 5571 | 2.895 | 423.327 |  | 1200.4 |  | 20231106_HepCl_ReRun_143 | Blank | Blank | 3.92 | 2.895 | 130 |  | 2.895 | 4 | 0 | -100 | 20231106_HepCl_ReRun_143 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_143 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_143 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_152 | Blank | Blank | 3.89 | 12.933 | 398 | 0.579 | 89.347 |  | 325.92 |  | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.81 | 16.621 | 419 | 0.579 | 114.826 |  | 1695.23 |  | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.79 | 0.579 | 34 |  | 0.579 | 4 | 0 | -100 | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.86 | 275.404 | 5049 | [5 empty] | 20231106_HepCl_ReRun_152 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.05 | 0.254 | 15 | [5 empty] | 20231106_HepCl_ReRun_152 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.05 | 4.853 | 276 | 6.476 | 2.998 |  | 49.43 |  | 20231106_HepCl_ReRun_152 | Blank | Blank |  |  |  | 6.476 |  |  |  |  | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.4 | 6.476 | 159 |  | 6.476 | 4 | 0 | -100
20231106_HepCl_ReRun_159 | Blank | Blank | 3.83 | 4.698 | 88 | [5 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.8 | 12.358 | 293 | [5 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.85 | 188.909 | 3291 | 4.77 | 158.414 |  | 448.95 |  | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.95 | 4.77 | 219 |  | 4.77 | 4 | 0 | -100 | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.08 | 0.303 | 18 | [5 empty] | 20231106_HepCl_ReRun_159 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.02 | 4.052 | 238 | 0.75 | 21.611 |  | 360.26 |  | 20231106_HepCl_ReRun_159 | Blank | Blank |  |  |  | 0.75 |  |  |  |  | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.4 | 0.75 | 88 |  | 0.75 | 4 | 0 | -100
20231106_HepCl_ReRun_166 | Blank | Blank |  |  |  | 0.413 |  |  |  |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 3.78 | 0.364 | 40 | 0.413 | 3.525 |  | 50.49 |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 3.82 | 0.413 | 15 |  | 0.413 | 4 | 0 | -100 | 20231106_HepCl_ReRun_166 | Blank | Blank |  |  |  | 6.376 |  |  |  |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 3.96 | 6.376 | 199 |  | 6.376 | 4 | 0 | -100 | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.05 | 0.883 | 52 | 0.343 | 10.297 |  | 60.77 |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.06 | 0.343 | 20 |  | 0.343 | 4 | 0 | -100 | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.04 | 8.785 | 356 | 3.511 | 10.009 |  | 166.51 |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.41 | 1.558 | 302 | 3.511 | 1.775 |  | 5.62 |  | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.41 | 3.511 | 188 |  | 3.511 | 4 | 0 | -100
20231106_HepCl_ReRun_168 | Blank | Blank |  |  |  | 2.772 |  |  |  |  | 20231106_HepCl_ReRun_168 | Blank | Blank | 3.82 | 10.882 | 381 | 2.772 | 15.703 |  | 230.44 |  | 20231106_HepCl_ReRun_168 | Blank | Blank | 3.81 | 2.772 | 100 |  | 2.772 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_168 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_168 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_168 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_168 | Blank | Blank | [5 empty] | 4
20231106_HepCl_ReRun_175 | Blank | Blank | 3.89 | 4.209 | 179 | [5 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | 3.82 | 11.069 | 300 | [5 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_175 | Blank | Blank | 3.85 | 215.745 | 4778 | [5 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_175 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_175 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.02 | 0.679 | 63 | 6.395 | 0.425 |  | 6.46 |  | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.43 | 88.735 | 1658 | 6.395 | 55.503 |  | 192.12 |  | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.4 | 6.395 | 288 |  | 6.395 | 4 | 0 | -100
20231106_HepCl_ReRun_182 | Blank | Blank | 3.84 | 7.722 | 217 | 1.668 | 18.518 |  | 66.97 |  | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.81 | 0.8 | 62 | 1.668 | 1.918 |  | 26.75 |  | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.81 | 1.668 | 38 |  | 1.668 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.86 | 188.246 | 4310 | 14.625 | 51.486 |  | 145.63 |  | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.96 | 14.625 | 344 |  | 14.625 | 4 | 0 | -100 | 20231106_HepCl_ReRun_182 | Blank | Blank | 4.02 | 0.146 | 9 | 0.546 | 1.07 |  | 5.41 |  | 20231106_HepCl_ReRun_182 | Blank | Blank | 4.09 | 0.546 | 33 |  | 0.546 | 4 | 0 | -100 | 20231106_HepCl_ReRun_182 | Blank | Blank |  |  |  | 0.43 |  |  |  |  | 20231106_HepCl_ReRun_182 | Blank | Blank |  |  |  | 0.43 |  |  |  |  | 20231106_HepCl_ReRun_182 | Blank | Blank | 4.4 | 0.43 | 81 |  | 0.43 | 4 | 0 | -100
20231106_HepCl_ReRun_189 | Blank | Blank | 3.87 | 10.259 | 349 | [5 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | 3.81 | 19.428 | 557 | [5 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_189 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_189 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.09 | 0.168 | 9 | 0.191 | 3.518 |  | 20.1 |  | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.06 | 0.191 | 11 |  | 0.191 | 4 | 0 | -100 | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.1 | 4.923 | 175 | 2.904 | 6.781 |  | 112.61 |  | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.43 | 14.778 | 1378 | 2.904 | 20.355 |  | 70.11 |  | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.42 | 2.904 | 87 |  | 2.904 | 4 | 0 | -100
20231106_HepCl_ReRun_191 | Blank | Blank | 3.83 | 5.461 | 215 | [5 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_191 | Blank | Blank | 3.85 | 107.023 | 2703 | 6.924 | 61.827 |  | 174.97 |  | 20231106_HepCl_ReRun_191 | Blank | Blank | 3.92 | 6.924 | 263 |  | 6.924 | 4 | 0 | -100 | 20231106_HepCl_ReRun_191 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_191 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.04 | 7.196 | 210 | 2.254 | 12.77 |  | 212.63 |  | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.41 | 3.016 | 787 | 2.254 | 5.352 |  | 18.03 |  | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.39 | 2.254 | 110 |  | 2.254 | 4 | 0 | -100
20231106_HepCl_ReRun_198 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.81 | 18.713 | 586 | [5 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.86 | 207.931 | 4316 | 3.35 | 248.276 |  | 703.85 |  | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.92 | 3.35 | 148 |  | 3.35 | 4 | 0 | -100 | 20231106_HepCl_ReRun_198 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_198 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_198 | Blank | Blank |  |  |  | 0.96 |  |  |  |  | 20231106_HepCl_ReRun_198 | Blank | Blank | 4.35 | 76.1 | 1422 | 0.96 | 317.083 |  | 1100.13 |  | 20231106_HepCl_ReRun_198 | Blank | Blank | 4.37 | 0.96 | 105 |  | 0.96 | 4 | 0 | -100
20231106_HepCl_ReRun_200 | Blank | Blank |  |  |  | 0.525 |  |  |  |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.82 | 9.87 | 225 | 0.525 | 75.2 |  | 1109.66 |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.81 | 0.525 | 61 |  | 0.525 | 4 | 0 | -100 | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.85 | 181.452 | 3449 | 2.811 | 258.203 |  | 732.01 |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 4 | 2.811 | 117 |  | 2.811 | 4 | 0 | -100 | 20231106_HepCl_ReRun_200 | Blank | Blank |  |  |  | 0.632 |  |  |  |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.04 | 0.632 | 37 |  | 0.632 | 4 | 0 | -100 | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.04 | 2.128 | 95 | 1.102 | 7.724 |  | 128.36 |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.43 | 185.614 | 2410 | 1.102 | 673.735 |  | 2338.17 |  | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.4 | 1.102 | 122 |  | 1.102 | 4 | 0 | -100
20231106_HepCl_ReRun_207 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank | 3.81 | 22.558 | 431 | [5 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_207 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_207 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.05 | 0.506 | 30 | 1.208 | 1.675 |  | 9.04 |  | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.08 | 1.208 | 32 |  | 1.208 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.1 | 3.972 | 181 | 1.83 | 8.682 |  | 144.36 |  | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.4 | 24.258 | 1634 | 1.83 | 53.023 |  | 183.51 |  | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.4 | 1.83 | 193 |  | 1.83 | 4 | 0 | -100
20231106_HepCl_ReRun_214 | Blank | Blank | 3.87 | 9.578 | 198 | [5 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | 3.82 | 21.328 | 482 | [5 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_214 | Blank | Blank | 3.85 | 12.237 | 909 | [5 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.04 | 0.323 | 19 | [5 empty] | 20231106_HepCl_ReRun_214 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.07 | 12.603 | 469 | 1.23 | 40.985 |  | 683.81 |  | 20231106_HepCl_ReRun_214 | Blank | Blank |  |  |  | 1.23 |  |  |  |  | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.43 | 1.23 | 122 |  | 1.23 | 4 | 0 | -100
20231106_HepCl_ReRun_221 | Blank | Blank | 3.87 | 8.066 | 252 | [5 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_221 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_221 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_221 | Blank | Blank |  |  |  | 0.465 |  |  |  |  | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.09 | 0.465 | 26 |  | 0.465 | 4 | 0 | -100 | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.02 | 10.347 | 233 | 0.419 | 98.778 |  | 1648.94 |  | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.36 | 269.191 | 2916 | 0.419 | 2569.842 |  | 8920.05 |  | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.38 | 0.419 | 76 |  | 0.419 | 4 | 0 | -100
20231106_HepCl_ReRun_223 | Blank | Blank | 3.87 | 3.376 | 132 | [5 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | 3.82 | 17.31 | 342 | [5 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_223 | Blank | Blank | 3.85 | 209.716 | 4694 | [5 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_223 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_223 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_223 | Blank | Blank | 4.05 | 3.697 | 187 | 5.807 | 2.547 |  | 41.9 |  | 20231106_HepCl_ReRun_223 | Blank | Blank |  |  |  | 5.807 |  |  |  |  | 20231106_HepCl_ReRun_223 | Blank | Blank | 4.39 | 5.807 | 229 |  | 5.807 | 4 | 0 | -100
20231106_HepCl_ReRun_230 | Blank | Blank | 3.83 | 4.137 | 166 | 4.215 | 3.926 |  | 13.62 |  | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.81 | 5.699 | 139 | 4.215 | 5.408 |  | 78.32 |  | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.82 | 4.215 | 142 |  | 4.215 | 4 | 0.01 | -99.9 | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.85 | 165.117 | 3099 | 9.549 | 69.166 |  | 195.78 |  | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.96 | 9.549 | 291 |  | 9.549 | 4 | 0 | -100 | 20231106_HepCl_ReRun_230 | Blank | Blank |  |  |  | 0.182 |  |  |  |  | 20231106_HepCl_ReRun_230 | Blank | Blank | 4.04 | 0.182 | 11 |  | 0.182 | 4 | 0 | -100 | 20231106_HepCl_ReRun_230 | Blank | Blank | 4.08 | 13.86 | 401 | 2.226 | 24.906 |  | 415.29 |  | 20231106_HepCl_ReRun_230 | Blank | Blank |  |  |  | 2.226 |  |  |  |  | 20231106_HepCl_ReRun_230 | Blank | Blank | 4.39 | 2.226 | 148 |  | 2.226 | 4 | 0 | -100
20231106_HepCl_ReRun_237 | Blank | Blank | 3.85 | 11.589 | 261 | 1.939 | 23.907 |  | 86.67 |  | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 1.939 |  |  |  |  | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.74 | 1.939 | 70 |  | 1.939 | 4 | 0 | -99.9 | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.86 | 148.115 | 2568 | 0.334 | 1773.832 |  | 5031.24 |  | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.98 | 0.334 | 38 |  | 0.334 | 4 | 0 | -100 | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 0.132 |  |  |  |  | 20231106_HepCl_ReRun_237 | Blank | Blank | 4.04 | 0.132 | 21 |  | 0.132 | 4 | 0 | -100 | 20231106_HepCl_ReRun_237 | Blank | Blank | 4.06 | 2.31 | 145 | 8.198 | 1.127 |  | 18.19 |  | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 8.198 |  |  |  |  | 20231106_HepCl_ReRun_237 | Blank | Blank | 4.37 | 8.198 | 144 |  | 8.198 | 4 | 0 | -100
20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 0.26 |  |  |  |  | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.82 | 13.663 | 294 | 0.26 | 210.2 |  | 3104.62 |  | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.79 | 0.26 | 9 |  | 0.26 | 4 | 0 | -100 | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.85 | 184.859 | 3229 | [5 empty] | 20231106_HepCl_ReRun_244 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 0.507 |  |  |  |  | 20231106_HepCl_ReRun_244 | Blank | Blank | 4.09 | 0.507 | 30 |  | 0.507 | 4 | 0 | -100 | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 7.8 |  |  |  |  | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 7.8 |  |  |  |  | 20231106_HepCl_ReRun_244 | Blank | Blank | 4.44 | 7.8 | 151 |  | 7.8 | 4 | 0 | -100
20231106_HepCl_ReRun_246 | Blank | Blank | 3.87 | 0.829 | 74 | [5 empty] | 20231106_HepCl_ReRun_246 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_246 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_246 | Blank | Blank | 3.86 | 213.496 | 3757 | 1.2 | 711.653 |  | 2018.26 |  | 20231106_HepCl_ReRun_246 | Blank | Blank | 3.96 | 1.2 | 64 |  | 1.2 | 4 | 0 | -100 | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.03 | 0.616 | 36 | 0.149 | 16.537 |  | 98.2 |  | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.04 | 0.149 | 9 |  | 0.149 | 4 | 0 | -100 | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.08 | 0.727 | 65 | 4.036 | 0.721 |  | 11.4 |  | 20231106_HepCl_ReRun_246 | Blank | Blank |  |  |  | 4.036 |  |  |  |  | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.4 | 4.036 | 240 |  | 4.036 | 4 | 0 | -100
20231106_HepCl_ReRun_253 | Blank | Blank | 3.84 | 5.148 | 173 | [5 empty] | 20231106_HepCl_ReRun_253 | Blank | Blank | 3.81 | 8.129 | 294 | [5 empty] | 20231106_HepCl_ReRun_253 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_253 | Blank | Blank |  |  |  | 12.938 |  |  |  |  | 20231106_HepCl_ReRun_253 | Blank | Blank | 3.96 | 12.938 | 311 |  | 12.938 | 4 | 0 | -100 | 20231106_HepCl_ReRun_253 | Blank | Blank |  |  |  | 0.066 |  |  |  |  | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.04 | 0.066 | 4 |  | 0.066 | 4 | 0 | -100 | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.06 | 0.597 | 60 | 1.238 | 1.929 |  | 31.58 |  | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.43 | 12.221 | 1633 | 1.238 | 39.486 |  | 136.52 |  | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.4 | 1.238 | 124 |  | 1.238 | 4 | 0 | -100
20231106_HepCl_ReRun_254 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | 3.81 | 13.784 | 295 | [5 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_254 | Blank | Blank | 3.86 | 153.336 | 3207 | [5 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_254 | Blank | Blank | [8 empty] | 20231106_HepCl_ReRun_254 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_254 | Blank | Blank | 4.08 | 4.24 | 210 | 7.609 | 2.229 |  | 36.59 |  | 20231106_HepCl_ReRun_254 | Blank | Blank |  |  |  | 7.609 |  |  |  |  | 20231106_HepCl_ReRun_254 | Blank | Blank | 4.39 | 7.609 | 217 |  | 7.609 | 4 | 0 | -100
20231106_HepCl_ReRun_255 | Blank | Blank | 3.86 | 7.01 | 252 | [5 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | 3.82 | 17.656 | 487 | [5 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_255 | Blank | Blank | 3.85 | 181.382 | 3786 | [5 empty] | 20231106_HepCl_ReRun_255 | Blank | Blank | [5 empty] | 4 |  |  | 20231106_HepCl_ReRun_255 | Blank | Blank |  |  |  | 0.647 |  |  |  |  | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.06 | 0.647 | 38 |  | 0.647 | 4 | 0 | -100 | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.01 | 0.367 | 74 | 3 | 0.489 |  | 7.54 |  | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.43 | 36.339 | 1486 | 3 | 48.452 |  | 167.65 |  | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.41 | 3 | 179 |  | 3 | 4 | 0 | -100

## Raw Data
Quantify Compound Summary Report
Printed Mon Nov 13 11:43:41 2023
Compound 1: 6:3 FTCA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20231106_HepCl_ReRun_001 | Blank | Blank | 3.86 | 7.841 | 320 | [5 empty] | bb | 2023-11-06
2 | 2 | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.88 | 1.966 | 103 | 2.377 | 3.308 |  | 11.37 |  | bb | 2023-11-06
3 | 3 | 20231106_HepCl_ReRun_003 | Blank | Blank |  |  |  | 0.237 | [5 empty] | 2023-11-06
4 | 4 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank |  |  |  | 8.997 | [5 empty] | 2023-11-06
5 | 5 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.9 | 7.372 | 255 | [5 empty] | bb | 2023-11-06
6 | 6 | 20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 8.596 | [5 empty] | 2023-11-06
7 | 7 | 20231106_HepCl_ReRun_007 | Blank | Blank | 3.88 | 6.992 | 172 | [5 empty] | bb | 2023-11-06
8 | 8 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 3.89 | 67.971 | 1011 | 3119.796 | 0.087 | 0.175 |  |  | bbI | 2023-11-06
9 | 9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 3.87 | 21.358 | 767 | 3010.686 | 0.028 | 0.281 |  |  | bbI | 2023-11-06
10 | 10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 3.86 | 120.148 | 2098 | 2911.38 | 0.165 | 0.444 |  |  | bbI | 2023-11-06
11 | 11 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 3.88 | 179.938 | 3117 | 3209.637 | 0.224 | 0.713 | 0.09 | -87.4 | bb | 2023-11-06
12 | 12 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 3.85 | 264.826 | 5128 | 3047.267 | 0.348 | 1.138 | 0.54 | -52.5 | bb | 2023-11-06
13 | 13 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.88 | 447.461 | 8778 | 3034.11 | 0.59 | 1.825 | 1.43 | -21.8 | bb | 2023-11-06
14 | 14 | 20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.04 | [5 empty] | 2023-11-06
15 | 15 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.88 | 779.956 | 11885 | 2993.711 | 1.042 | 2.919 | 3.08 | 5.5 | bb | 2023-11-06
16 | 16 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.88 | 1201.809 | 19563 | 3101.963 | 1.55 | 4.669 | 4.94 | 5.7 | bb | 2023-11-06
17 | 17 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.88 | 1982.509 | 35805 | 3138.186 | 2.527 | 7.45 | 8.51 | 14.2 | bb | 2023-11-06
18 | 18 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.87 | 2973.767 | 47181 | 3163.276 | 3.76 | 11.875 | 13.02 | 9.6 | bb | 2023-11-06
19 | 19 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.88 | 5696.95 | 94449 | 3229.261 | 7.057 | 19.125 | 25.07 | 31.1 | bb | 2023-11-06
20 | 20 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.87 | 6181.179 | 122226 | 2767.938 | 8.933 | 30.5 | 31.93 | 4.7 | MM | 2023-11-06
21 | 21 | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.9 | 10.663 | 141 | 2.584 | 16.506 |  | 59.62 |  | bb | 2023-11-06
22 | 22 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.88 | 11224.8 | 178932 | 3439.733 | 13.053 | 48.75 | 46.99 | -3.6 | bb | 2023-11-06
23 | 23 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.88 | 22367.092 | 397198 | 3919.615 | 22.826 | 78.125 | 82.72 | 5.9 | bb | 2023-11-06
24 | 24 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.88 | 30114.336 | 533010 | 3888.294 | 30.979 | 125 | 112.53 | -10 | bb | 2023-11-06
25 | 25 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 3.88 | 51578.672 | 891223 | 4835.771 | 42.664 | 200 | 155.25 | -22.4 | bb | 2023-11-06
26 | 26 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 3.88 | 91458.242 | 1662166 | 5718.93 | 63.969 | 325.625 | 233.14 | -28.4 | bb | 2023-11-06
27 | 27 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.87 | 35.183 | 628 | 1.072 | 131.28 |  | 479.23 |  | bb | 2023-11-06
28 | 28 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | [9 empty] | 2023-11-06
29 | 29 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 3.88 | 221.882 | 3836 | 2902.343 | 0.306 | 0.713 | 0.39 | -45.6 | bb | 2023-11-06
30 | 30 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 3.88 | 379.244 | 6609 | 2148.496 | 0.706 | 1.138 | 1.85 | 62.7 | bb | 2023-11-06
31 | 31 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.88 | 478.13 | 8539 | 2368.634 | 0.807 | 1.825 | 2.22 | 21.7 | bb | 2023-11-06
32 | 32 | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.87 | 5.334 | 155 | 0.567 | 37.63 |  | 136.84 |  | bb | 2023-11-06
33 | 33 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 3.88 | 215.362 | 4509 | 2884.729 | 0.299 | 0.713 | 0.36 | -49.3 | bb | 2023-11-06
34 | 34 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 3.88 | 245.912 | 4041 | 3311.456 | 0.297 | 1.138 | 0.36 | -68.7 | bb | 2023-11-06
35 | 35 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.87 | 418.264 | 7300 | 2974.036 | 0.563 | 1.825 | 1.33 | -27.3 | bb | 2023-11-06
36 | 36 | 20231106_HepCl_ReRun_036 | Blank | Blank | 3.92 | 10.411 | 310 | [5 empty] | bb | 2023-11-06
37 | 37 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 3.88 | 192.081 | 3259 | 2898.57 | 0.265 | 0.713 | 0.24 | -66.5 | bb | 2023-11-06
38 | 38 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 3.86 | 288.183 | 4636 | 2369.425 | 0.487 | 1.138 | 1.05 | -7.8 | bb | 2023-11-06
39 | 39 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.88 | 523.429 | 9106 | 2930.851 | 0.714 | 1.825 | 1.88 | 3.1 | bb | 2023-11-06
40 | 40 | 20231106_HepCl_ReRun_040 | Blank | Blank |  |  |  | 2.239 | [5 empty] | 2023-11-06
41 | 41 | 20231106_HepCl_ReRun_041 | QC A | QC | 3.88 | 165.842 | 3159 | 3334.889 | 0.199 | 0.625 |  |  | bbI | 2023-11-06
42 | 42 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.87 | 710.963 | 12145 | 3127.347 | 0.909 | 2.506 | 2.59 | 3.5 | bb | 2023-11-06
43 | 43 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.88 | 1686.431 | 30672 | 3371.759 | 2.001 | 6.269 | 6.58 | 5 | MM | 2023-11-06
44 | 44 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.88 | 6550.212 | 126329 | 3225.131 | 8.124 | 25 | 28.97 | 15.9 | MM | 2023-11-06
45 | 45 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.9 | 2.746 | 126 | [5 empty] | bb | 2023-11-06
46 | 46 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 3.82 | 2.362 | 85 | [5 empty] | bb | 2023-11-06
47 | 47 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 3.88 | 43.373 | 826 | 2990.745 | 0.058 | 0.175 |  |  | bbI | 2023-11-06
48 | 48 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 3.88 | 30.829 | 1001 | 3253.131 | 0.038 | 0.281 |  |  | bbI | 2023-11-06
49 | 49 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 3.88 | 110.652 | 1988 | 2768.087 | 0.16 | 0.444 |  |  | bbI | 2023-11-06
50 | 50 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 3.88 | 205.072 | 3755 | 3070.196 | 0.267 | 0.713 | 0.25 | -65.4 | bb | 2023-11-06
51 | 51 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 3.88 | 305.641 | 5035 | 2763.465 | 0.442 | 1.138 | 0.89 | -22 | bb | 2023-11-06
52 | 52 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.88 | 452.235 | 7649 | 2982.801 | 0.606 | 1.825 | 1.49 | -18.5 | bb | 2023-11-06
53 | 53 | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.9 | 1.849 | 69 | 1.263 | 5.856 |  | 20.68 |  | bb | 2023-11-06
54 | 54 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.88 | 712.941 | 14349 | 3350.341 | 0.851 | 2.919 | 2.38 | -18.4 | bb | 2023-11-06
55 | 55 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.87 | 1077.497 | 18622 | 2541.916 | 1.696 | 4.669 | 5.47 | 17.1 | bb | 2023-11-06
56 | 56 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.87 | 2015.593 | 34908 | 3151.321 | 2.558 | 7.45 | 8.62 | 15.7 | bb | 2023-11-06
57 | 57 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.88 | 3366.109 | 59308 | 3307.122 | 4.071 | 11.875 | 14.15 | 19.2 | bb | 2023-11-06
58 | 58 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.88 | 5013.81 | 76829 | 3103.845 | 6.461 | 19.125 | 22.89 | 19.7 | bb | 2023-11-06
59 | 59 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.88 | 8168.249 | 142368 | 3365.248 | 9.709 | 30.5 | 34.77 | 14 | bb | 2023-11-06
60 | 60 | 20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 1.55 | [5 empty] | 2023-11-06
61 | 61 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.88 | 12265.34 | 222118 | 3673.549 | 13.355 | 48.75 | 48.1 | -1.3 | bb | 2023-11-06
62 | 62 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.88 | 21361.063 | 375481 | 4066.942 | 21.009 | 78.125 | 76.08 | -2.6 | bb | 2023-11-06
63 | 63 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.88 | 32258.174 | 560369 | 3980.426 | 32.417 | 125 | 117.79 | -5.8 | bb | 2023-11-06
64 | 64 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 3.88 | 47232.074 | 763949 | 4885.239 | 38.673 | 200 | 140.66 | -29.7 | bb | 2023-11-06
65 | 65 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 3.88 | 90726.891 | 1682119 | 6843.28 | 53.031 | 325.625 | 193.15 | -40.7 | bb | 2023-11-07
66 | 66 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 0.37 | [5 empty] | 2023-11-07
67 | 67 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.328 | [5 empty] | 2023-11-07
68 | 68 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 3.88 | 218.265 | 4061 | 3113.354 | 0.28 | 0.713 | 0.29 | -58.6 | bb | 2023-11-07
69 | 69 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 3.88 | 284.915 | 4827 | 2642.521 | 0.431 | 1.138 | 0.85 | -25.6 | bb | 2023-11-07
70 | 70 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.88 | 472.381 | 7598 | 3064.16 | 0.617 | 1.825 | 1.52 | -16.5 | bb | 2023-11-07
71 | 71 | 20231106_HepCl_ReRun_071 | Blank | Blank |  |  |  | 5.703 | [5 empty] | 2023-11-07
72 | 72 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 3.88 | 154.409 | 2781 | 2424.344 | 0.255 | 0.713 | 0.2 | -71.8 | bb | 2023-11-07
73 | 73 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 3.88 | 256.321 | 5150 | 3320.192 | 0.309 | 1.138 | 0.4 | -65 | bb | 2023-11-07
74 | 74 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.88 | 432.128 | 7512 | 3262.376 | 0.53 | 1.825 | 1.21 | -33.9 | bb | 2023-11-07
75 | 75 | 20231106_HepCl_ReRun_075 | Blank | Blank |  |  |  | 0.322 | [5 empty] | 2023-11-07
76 | 76 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 3.88 | 181.043 | 2968 | 3100.597 | 0.234 | 0.713 | 0.12 | -82.7 | bb | 2023-11-07
77 | 77 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 3.87 | 303.546 | 4891 | 3376.185 | 0.36 | 1.138 | 0.58 | -48.6 | bb | 2023-11-07
78 | 78 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.88 | 465.458 | 7889 | 3313.979 | 0.562 | 1.825 | 1.32 | -27.5 | bb | 2023-11-07
79 | 79 | 20231106_HepCl_ReRun_079 | Blank | Blank | 3.83 | 0.96 | 30 | 4.597 | 0.835 |  | 2.32 |  | bb | 2023-11-07
80 | 80 | 20231106_HepCl_ReRun_080 | QC A | QC | 3.87 | 234.077 | 3816 | 3603.385 | 0.26 | 0.625 | 0.22 | -64.9 | bb | 2023-11-07
81 | 81 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.87 | 752.553 | 12162 | 2887.97 | 1.042 | 2.506 | 3.08 | 22.9 | bb | 2023-11-07
82 | 82 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.88 | 2095.343 | 37365 | 3610.88 | 2.321 | 6.269 | 7.76 | 23.7 | bb | 2023-11-07
83 | 83 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.88 | 7293.705 | 119126 | 3571.287 | 8.169 | 25 | 29.14 | 16.5 | bb | 2023-11-07
84 | 84 | 20231106_HepCl_ReRun_084 | Blank | Blank | [9 empty] | 2023-11-07
85 | 85 | 20231106_HepCl_ReRun_085 | Blank | Blank | 3.86 | 1.076 | 86 | [5 empty] | bb | 2023-11-07
86 | 86 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 3.88 | 8971.841 | 138720 | 3342.74 | 10.736 | 30.5 | 38.52 | 26.3 | bb | 2023-11-07
87 | 87 | 20231106_HepCl_ReRun_087 | Blank | Blank | 3.89 | 6.497 | 168 | [5 empty] | bb | 2023-11-07
88 | 88 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.89 | 15743.503 | 247765 | 4776.272 | 13.185 |  | 47.47 |  | bb | 2023-11-07
89 | 89 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.88 | 13915.396 | 245491 | 3195.141 | 17.421 |  | 62.96 |  | bb | 2023-11-07
90 | 90 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.88 | 14715.912 | 240678 | 4462.737 | 13.19 |  | 47.49 |  | bb | 2023-11-07
91 | 91 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.88 | 12368.906 | 206371 | 4447.937 | 11.123 |  | 39.94 |  | bb | 2023-11-07
92 | 92 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.88 | 12134.313 | 203497 | 4036.639 | 12.024 |  | 43.23 |  | bb | 2023-11-07
93 | 93 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.88 | 11903.52 | 204453 | 3155.708 | 15.088 |  | 54.43 |  | bb | 2023-11-07
94 | 94 | 20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 1.329 | [5 empty] | 2023-11-07
95 | 95 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.88 | 11315.929 | 169106 | 4054.992 | 11.162 |  | 40.08 |  | bb | 2023-11-07
96 | 96 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.87 | 12633.037 | 196143 | 4225.446 | 11.959 |  | 42.99 |  | bb | 2023-11-07
97 | 97 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.88 | 11869.145 | 199049 | 4254.767 | 11.158 |  | 40.06 |  | bb | 2023-11-07
98 | 98 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.88 | 11134.584 | 184254 | 4188.782 | 10.633 |  | 38.14 |  | bb | 2023-11-07
99 | 99 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.87 | 10594.069 | 161173 | 4496.683 | 9.424 |  | 33.72 |  | bb | 2023-11-07
100 | 100 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.88 | 10354.753 | 164488 | 4153.806 | 9.971 |  | 35.72 |  | bb | 2023-11-07
101 | 101 | 20231106_HepCl_ReRun_101 | Blank | Blank | 3.83 | 17.593 | 249 | [5 empty] | bb | 2023-11-07
102 | 102 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.87 | 9074.265 | 137502 | 3209.497 | 11.309 |  | 40.62 |  | bb | 2023-11-07
103 | 103 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.88 | 9561.769 | 153073 | 3877.044 | 9.865 |  | 35.34 |  | bb | 2023-11-07
104 | 104 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.88 | 9736.915 | 150985 | 3869.138 | 10.066 |  | 36.07 |  | bb | 2023-11-07
105 | 105 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.87 | 6488.23 | 94317 | 3451.207 | 7.52 |  | 26.76 |  | bb | 2023-11-07
106 | 106 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.88 | 4897.804 | 78547 | 3384.839 | 5.788 |  | 20.43 |  | bb | 2023-11-07
107 | 107 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.88 | 7423.349 | 109685 | 3411.54 | 8.704 |  | 31.09 |  | bb | 2023-11-07
108 | 108 | 20231106_HepCl_ReRun_108 | Blank | Blank | 3.85 | 6.219 | 242 | [5 empty] | bb | 2023-11-07
109 | 109 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 3.88 | 7950.584 | 118420 | 3304.299 | 9.625 | 30.5 | 34.46 | 13 | bb | 2023-11-07
110 | 110 | 20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 3.311 | [5 empty] | 2023-11-07
111 | 111 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.87 | 13458.945 | 179811 | 3946.084 | 13.643 |  | 49.15 |  | bb | 2023-11-07
112 | 112 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.88 | 13394.592 | 210981 | 4015.599 | 13.343 |  | 48.05 |  | bb | 2023-11-07
113 | 113 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.88 | 12774.548 | 184155 | 3722.044 | 13.729 |  | 49.46 |  | bb | 2023-11-07
114 | 114 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.87 | 13906.821 | 198513 | 3719.442 | 14.956 |  | 53.95 |  | bb | 2023-11-07
115 | 115 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.88 | 12754.647 | 175426 | 3898.412 | 13.087 |  | 47.12 |  | bb | 2023-11-07
116 | 116 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.88 | 13496.068 | 184112 | 3840.235 | 14.058 |  | 50.66 |  | bb | 2023-11-07
117 | 117 | 20231106_HepCl_ReRun_117 | Blank | Blank | 3.91 | 20.044 | 377 | [5 empty] | bb | 2023-11-07
118 | 118 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.88 | 14466.164 | 202071 | 3796.952 | 15.24 |  | 54.99 |  | bb | 2023-11-07
119 | 119 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.87 | 13676.497 | 198516 | 3988.584 | 13.716 |  | 49.41 |  | bb | 2023-11-07
120 | 120 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.87 | 13672.899 | 183759 | 4089.779 | 13.373 |  | 48.16 |  | bb | 2023-11-07
121 | 121 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.87 | 13819.075 | 185596 | 4257.981 | 12.982 |  | 46.73 |  | bb | 2023-11-07
122 | 122 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.86 | 14415.857 | 197477 | 4107.876 | 14.037 |  | 50.59 |  | bb | 2023-11-07
123 | 123 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.87 | 13589.694 | 179296 | 3864.161 | 14.067 |  | 50.7 |  | bb | 2023-11-07
124 | 124 | 20231106_HepCl_ReRun_124 | Blank | Blank | [9 empty] | 2023-11-07
125 | 125 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.87 | 13366.214 | 177634 | 3543.577 | 15.088 |  | 54.43 |  | bb | 2023-11-07
126 | 126 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.87 | 12986.323 | 189758 | 3555.943 | 14.608 |  | 52.68 |  | bb | 2023-11-07
127 | 127 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.86 | 13166.602 | 186700 | 3988.081 | 13.206 |  | 47.55 |  | bb | 2023-11-07
128 | 128 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.87 | 11126.391 | 155600 | 3506.183 | 12.693 |  | 45.68 |  | bb | 2023-11-07
129 | 129 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.87 | 9472.6 | 127395 | 2791.078 | 13.576 |  | 48.9 |  | bb | 2023-11-07
130 | 130 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.87 | 9704.444 | 128645 | 3117.41 | 12.452 |  | 44.79 |  | bb | 2023-11-07
131 | 131 | 20231106_HepCl_ReRun_131 | Blank | Blank |  |  |  | 1.14 | [5 empty] | 2023-11-07
132 | 132 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 3.87 | 7783.227 | 105462 | 3604.25 | 8.638 | 30.5 | 30.85 | 1.1 | bb | 2023-11-07
133 | 133 | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.82 | 1.416 | 98 | 0.164 | 34.537 |  | 125.54 |  | bb | 2023-11-07
134 | 134 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.87 | 13138.3 | 187084 | 4056.419 | 12.956 |  | 46.64 |  | bb | 2023-11-07
135 | 135 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.87 | 14517.875 | 206204 | 4559.46 | 12.736 |  | 45.83 |  | bb | 2023-11-07
136 | 136 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.87 | 15841.149 | 227470 | 4493.075 | 14.103 |  | 50.83 |  | bb | 2023-11-07
137 | 137 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.87 | 12003.916 | 162939 | 4114.258 | 11.671 |  | 41.94 |  | bb | 2023-11-07
138 | 138 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.86 | 13315.447 | 183072 | 4004.459 | 13.301 |  | 47.9 |  | bb | 2023-11-07
139 | 139 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.87 | 13054.155 | 177255 | 3985.281 | 13.102 |  | 47.17 |  | bb | 2023-11-07
140 | 140 | 20231106_HepCl_ReRun_140 | Blank | Blank | 3.88 | 3.301 | 150 | [5 empty] | bb | 2023-11-07
141 | 141 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 3.87 | 7408.023 | 99582 | 3653.36 | 8.111 | 30.5 | 28.92 | -5.2 | bb | 2023-11-07
142 | 142 | 20231106_HepCl_ReRun_142 | Blank | Blank |  |  |  | 2.523 | [5 empty] | 2023-11-07
143 | 143 | 20231106_HepCl_ReRun_143 | Blank | Blank |  |  |  | 5.162 | [5 empty] | 2023-11-07
144 | 144 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.86 | 62.169 | 761 | 5093.202 | 0.049 |  |  |  | bbI | 2023-11-07
145 | 145 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte |  |  |  | 4939.262 | [5 empty] | 2023-11-07
146 | 146 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.85 | 18.73 | 503 | 5233.792 | 0.014 |  |  |  | bbI | 2023-11-07
147 | 147 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.9 | 39.308 | 527 | 4518.646 | 0.035 |  |  |  | bbI | 2023-11-07
148 | 148 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.86 | 0.79 | 104 | 4595.066 | 0.001 |  |  |  | bbI | 2023-11-07
149 | 149 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.85 | 9.745 | 292 | 4516.958 | 0.009 |  |  |  | bbI | 2023-11-07
150 | 150 | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.89 | 12.933 | 398 | 0.579 | 89.347 |  | 325.92 |  | bb | 2023-11-07
151 | 151 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte |  |  |  | 4562.208 | [5 empty] | 2023-11-07
152 | 152 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.88 | 0.404 | 57 | 4277.345 | 0 |  |  |  | bbI | 2023-11-07
153 | 153 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.87 | 15.454 | 583 | 4028.417 | 0.015 |  |  |  | bbI | 2023-11-07
154 | 154 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.84 | 4.492 | 168 | 4201.701 | 0.004 |  |  |  | bbI | 2023-11-07
155 | 155 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.85 | 3.701 | 145 | 4522.662 | 0.003 |  |  |  | bbI | 2023-11-07
156 | 156 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte |  |  |  | 4864.867 | [5 empty] | 2023-11-07
157 | 157 | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.83 | 4.698 | 88 | [5 empty] | bb | 2023-11-07
158 | 158 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte |  |  |  | 4130.82 | [5 empty] | 2023-11-07
159 | 159 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.87 | 11.221 | 400 | 4509.017 | 0.01 |  |  |  | bbI | 2023-11-07
160 | 160 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.87 | 13.982 | 290 | 4377.103 | 0.013 |  |  |  | bbI | 2023-11-07
161 | 161 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte |  |  |  | 3747.828 | [5 empty] | 2023-11-07
162 | 162 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte |  |  |  | 4038.739 | [5 empty] | 2023-11-07
163 | 163 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.87 | 21.437 | 444 | 3789.513 | 0.023 |  |  |  | bbI | 2023-11-07
164 | 164 | 20231106_HepCl_ReRun_166 | Blank | Blank |  |  |  | 0.413 | [5 empty] | 2023-11-07
165 | 165 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 3.87 | 6318.074 | 98306 | 3208.256 | 7.877 | 30.5 | 28.07 | -8 | bb | 2023-11-07
166 | 166 | 20231106_HepCl_ReRun_168 | Blank | Blank |  |  |  | 2.772 | [5 empty] | 2023-11-07
167 | 167 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.88 | 4.288 | 199 | 4403.589 | 0.004 |  |  |  | bbI | 2023-11-07
168 | 168 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.87 | 11.475 | 363 | 4598.22 | 0.01 |  |  |  | bbI | 2023-11-07
169 | 169 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.85 | 10.074 | 350 | 4689.85 | 0.009 |  |  |  | bbI | 2023-11-07
170 | 170 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.88 | 6.985 | 141 | 4132.267 | 0.007 |  |  |  | bbI | 2023-11-07
171 | 171 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.86 | 50.12 | 1129 | 4125.632 | 0.049 |  |  |  | bbI | 2023-11-07
172 | 172 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.88 | 5.404 | 192 | 4422.781 | 0.005 |  |  |  | bbI | 2023-11-07
173 | 173 | 20231106_HepCl_ReRun_175 | Blank | Blank | 3.89 | 4.209 | 179 | [5 empty] | bb | 2023-11-07
174 | 174 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.85 | 8.917 | 245 | 4699.705 | 0.008 |  |  |  | bbI | 2023-11-07
175 | 175 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.88 | 44.112 | 881 | 4633.228 | 0.038 |  |  |  | bbI | 2023-11-07
176 | 176 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.88 | 9.015 | 235 | 4931.43 | 0.007 |  |  |  | bbI | 2023-11-07
177 | 177 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.85 | 10.279 | 210 | 4408.869 | 0.009 |  |  |  | bbI | 2023-11-07
178 | 178 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte |  |  |  | 4910.336 | [5 empty] | 2023-11-07
179 | 179 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.85 | 10.55 | 300 | 4659.491 | 0.009 |  |  |  | bbI | 2023-11-07
180 | 180 | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.84 | 7.722 | 217 | 1.668 | 18.518 |  | 66.97 |  | bb | 2023-11-07
181 | 181 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.83 | 9.293 | 285 | 3744.223 | 0.01 |  |  |  | bbI | 2023-11-07
182 | 182 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte |  |  |  | 4034.394 | [5 empty] | 2023-11-07
183 | 183 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.87 | 19.84 | 431 | 4172.714 | 0.019 |  |  |  | bbI | 2023-11-07
184 | 184 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte |  |  |  | 3922.801 | [5 empty] | 2023-11-07
185 | 185 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.85 | 29.131 | 366 | 3270.064 | 0.036 |  |  |  | bbI | 2023-11-07
186 | 186 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte |  |  |  | 4049.156 | [5 empty] | 2023-11-07
187 | 187 | 20231106_HepCl_ReRun_189 | Blank | Blank | 3.87 | 10.259 | 349 | [5 empty] | bb | 2023-11-07
188 | 188 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 3.87 | 3027.379 | 74803 | 3623.284 | 3.342 | 30.5 | 11.49 | -62.3 | bb | 2023-11-07
189 | 189 | 20231106_HepCl_ReRun_191 | Blank | Blank | 3.83 | 5.461 | 215 | [5 empty] | bb | 2023-11-07
190 | 190 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.87 | 38.099 | 1027 | 4646.139 | 0.033 |  |  |  | bbI | 2023-11-07
191 | 191 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte |  |  |  | 4670.868 | [5 empty] | 2023-11-07
192 | 192 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.88 | 34.705 | 698 | 4760.106 | 0.029 |  |  |  | bbI | 2023-11-07
193 | 193 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.87 | 16.086 | 360 | 4087.066 | 0.016 |  |  |  | bbI | 2023-11-07
194 | 194 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.88 | 16.546 | 403 | 4565.343 | 0.014 |  |  |  | bbI | 2023-11-07
195 | 195 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.85 | 20.354 | 479 | 4533.528 | 0.018 |  |  |  | bbI | 2023-11-07
196 | 196 | 20231106_HepCl_ReRun_198 | Blank | Blank | [9 empty] | 2023-11-08
197 | 197 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 3.87 | 2826.421 | 68207 | 2966.934 | 3.811 | 30.5 | 13.2 | -56.7 | bb | 2023-11-08
198 | 198 | 20231106_HepCl_ReRun_200 | Blank | Blank |  |  |  | 0.525 | [5 empty] | 2023-11-08
199 | 199 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte |  |  |  | 4314.995 | [5 empty] | 2023-11-08
200 | 200 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.9 | 8.644 | 217 | 4169.22 | 0.008 |  |  |  | bbI | 2023-11-08
201 | 201 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.88 | 5.067 | 195 | 4343.747 | 0.005 |  |  |  | bbI | 2023-11-08
202 | 202 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte |  |  |  | 3994.895 | [5 empty] | 2023-11-08
203 | 203 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte |  |  |  | 3869.739 | [5 empty] | 2023-11-08
204 | 204 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte |  |  |  | 4094.355 | [5 empty] | 2023-11-08
205 | 205 | 20231106_HepCl_ReRun_207 | Blank | Blank | [9 empty] | 2023-11-08
206 | 206 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.86 | 51.112 | 1346 | 4088.134 | 0.05 |  |  |  | bbI | 2023-11-08
207 | 207 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.82 | 8.796 | 183 | 3656.892 | 0.01 |  |  |  | bbI | 2023-11-08
208 | 208 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.9 | 3.336 | 179 | 4115.619 | 0.003 |  |  |  | bbI | 2023-11-08
209 | 209 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.89 | 5.356 | 204 | 3668.877 | 0.006 |  |  |  | bbI | 2023-11-08
210 | 210 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.9 | 4.782 | 169 | 3204.234 | 0.006 |  |  |  | bbI | 2023-11-08
211 | 211 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.84 | 3.179 | 97 | 3915.398 | 0.003 |  |  |  | bbI | 2023-11-08
212 | 212 | 20231106_HepCl_ReRun_214 | Blank | Blank | 3.87 | 9.578 | 198 | [5 empty] | bb | 2023-11-08
213 | 213 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.84 | 3.194 | 60 | 3898.605 | 0.003 |  |  |  | bbI | 2023-11-08
214 | 214 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.88 | 14.696 | 311 | 4061.239 | 0.014 |  |  |  | bbI | 2023-11-08
215 | 215 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte |  |  |  | 3517.021 | [5 empty] | 2023-11-08
216 | 216 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte |  |  |  | 3202.326 | [5 empty] | 2023-11-08
217 | 217 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte |  |  |  | 3661.565 | [5 empty] | 2023-11-08
218 | 218 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.85 | 3.96 | 193 | 2804.06 | 0.006 |  |  |  | bbI | 2023-11-08
219 | 219 | 20231106_HepCl_ReRun_221 | Blank | Blank | 3.87 | 8.066 | 252 | [5 empty] | bb | 2023-11-08
220 | 220 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 3.87 | 2645.947 | 69209 | 3886.825 | 2.723 | 30.5 | 9.22 | -69.8 | bb | 2023-11-08
221 | 221 | 20231106_HepCl_ReRun_223 | Blank | Blank | 3.87 | 3.376 | 132 | [5 empty] | bb | 2023-11-08
222 | 222 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte |  |  |  | 4173.079 | [5 empty] | 2023-11-08
223 | 223 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.89 | 7.057 | 258 | 3930.435 | 0.007 |  |  |  | bbI | 2023-11-08
224 | 224 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte |  |  |  | 3689.984 | [5 empty] | 2023-11-08
225 | 225 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.9 | 4.366 | 159 | 3838.459 | 0.005 |  |  |  | bbI | 2023-11-08
226 | 226 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.85 | 2.995 | 159 | 3972.163 | 0.003 |  |  |  | bbI | 2023-11-08
227 | 227 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.9 | 0.645 | 36 | 3842.832 | 0.001 |  |  |  | bbI | 2023-11-08
228 | 228 | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.83 | 4.137 | 166 | 4.215 | 3.926 |  | 13.62 |  | bb | 2023-11-08
229 | 229 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.87 | 0.965 | 116 | 3893.189 | 0.001 |  |  |  | bbI | 2023-11-08
230 | 230 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.9 | 5.351 | 219 | 3559.379 | 0.006 |  |  |  | bbI | 2023-11-08
231 | 231 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.84 | 2.901 | 108 | 4010.53 | 0.003 |  |  |  | bbI | 2023-11-08
232 | 232 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.87 | 6.052 | 191 | 4683.644 | 0.005 |  |  |  | bbI | 2023-11-08
233 | 233 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte |  |  |  | 3473.603 | [5 empty] | 2023-11-08
234 | 234 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte |  |  |  | 3760.475 | [5 empty] | 2023-11-08
235 | 235 | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.85 | 11.589 | 261 | 1.939 | 23.907 |  | 86.67 |  | bb | 2023-11-08
236 | 236 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.85 | 3.359 | 200 | 3804.575 | 0.004 |  |  |  | bbI | 2023-11-08
237 | 237 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.85 | 0.614 | 56 | 3650.367 | 0.001 |  |  |  | bbI | 2023-11-08
238 | 238 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.82 | 3.82 | 148 | 4028.423 | 0.004 |  |  |  | bbI | 2023-11-08
239 | 239 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.88 | 8.494 | 187 | 3605.647 | 0.009 |  |  |  | bbI | 2023-11-08
240 | 240 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte |  |  |  | 3280.838 | [5 empty] | 2023-11-08
241 | 241 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte |  |  |  | 3179.154 | [5 empty] | 2023-11-08
242 | 242 | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 0.26 | [5 empty] | 2023-11-08
243 | 243 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 3.86 | 2143.984 | 53680 | 3753.745 | 2.285 | 30.5 | 7.62 | -75 | bb | 2023-11-08
244 | 244 | 20231106_HepCl_ReRun_246 | Blank | Blank | 3.87 | 0.829 | 74 | [5 empty] | bb | 2023-11-08
245 | 245 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte |  |  |  | 4269.086 | [5 empty] | 2023-11-08
246 | 246 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.9 | 0.276 | 37 | 4165.854 | 0 |  |  |  | bbI | 2023-11-08
247 | 247 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.87 | 3.55 | 130 | 4105.743 | 0.003 |  |  |  | bbI | 2023-11-08
248 | 248 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.91 | 5.588 | 229 | 4091.521 | 0.005 |  |  |  | bbI | 2023-11-08
249 | 249 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.84 | 5.883 | 249 | 3699.73 | 0.006 |  |  |  | bbI | 2023-11-08
250 | 250 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.88 | 3.841 | 109 | 3781.379 | 0.004 |  |  |  | bbI | 2023-11-08
251 | 251 | 20231106_HepCl_ReRun_253 | Blank | Blank | 3.84 | 5.148 | 173 | [5 empty] | bb | 2023-11-08
252 | 252 | 20231106_HepCl_ReRun_254 | Blank | Blank | [9 empty] | 2023-11-08
253 | 253 | 20231106_HepCl_ReRun_255 | Blank | Blank | 3.86 | 7.01 | 252 | [5 empty] | bb | 2023-11-08
Compound 2: 6:2 FTS
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20231106_HepCl_ReRun_001 | Blank | Blank | 3.82 | 149.304 | 2967 | [5 empty] | bb | 2023-11-06
2 | 2 | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.81 | 70.268 | 1184 | 2.377 | 118.247 |  | 1745.78 |  | bb | 2023-11-06
3 | 3 | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.83 | 51.466 | 878 | 0.237 | 868.624 |  | 12834.47 |  | bb | 2023-11-06
4 | 4 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.81 | 29.492 | 926 | 8.997 | 13.112 |  | 192.16 |  | bb | 2023-11-06
5 | 5 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.82 | 33.317 | 899 | [5 empty] | bb | 2023-11-06
6 | 6 | 20231106_HepCl_ReRun_006 | Blank | Blank | 3.81 | 46.309 | 811 | 8.596 | 21.549 |  | 316.84 |  | bb | 2023-11-06
7 | 7 | 20231106_HepCl_ReRun_007 | Blank | Blank | 3.81 | 31.905 | 854 | [5 empty] | bb | 2023-11-06
8 | 8 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 3.78 | 67.876 | 1854 | 3119.796 | 0.087 | 0.175 |  |  | bbI | 2023-11-06
9 | 9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 3.78 | 99.02 | 2305 | 3010.686 | 0.132 | 0.281 | 0.34 | 21.5 | bb | 2023-11-06
10 | 10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 3.78 | 95.6 | 2399 | 2911.38 | 0.131 | 0.444 | 0.34 | -23.7 | bb | 2023-11-06
11 | 11 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 3.78 | 141.693 | 3935 | 3209.637 | 0.177 | 0.713 | 1.01 | 41.3 | bb | 2023-11-06
12 | 12 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 3.77 | 160.908 | 4889 | 3047.267 | 0.211 | 1.138 | 1.52 | 33.5 | bb | 2023-11-06
13 | 13 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.78 | 171.963 | 5263 | 3034.11 | 0.227 | 1.825 | 1.75 | -4.2 | bb | 2023-11-06
14 | 14 | 20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.04 | [5 empty] | 2023-11-06
15 | 15 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.78 | 238.272 | 7112 | 2993.711 | 0.318 | 2.919 | 3.1 | 6.3 | bb | 2023-11-06
16 | 16 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.78 | 347.814 | 11862 | 3101.963 | 0.449 | 4.669 | 5.03 | 7.6 | bb | 2023-11-06
17 | 17 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.78 | 584.512 | 18989 | 3138.186 | 0.745 | 7.45 | 9.41 | 26.3 | bb | 2023-11-06
18 | 18 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.78 | 761.471 | 26224 | 3163.276 | 0.963 | 11.875 | 12.63 | 6.3 | bb | 2023-11-06
19 | 19 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.78 | 1328.398 | 45730 | 3229.261 | 1.645 | 19.125 | 22.71 | 18.8 | bb | 2023-11-06
20 | 20 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.77 | 1914.911 | 65185 | 2767.938 | 2.767 | 30.5 | 39.29 | 28.8 | bb | 2023-11-06
21 | 21 | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.82 | 11.18 | 423 | 2.584 | 17.307 |  | 254.14 |  | bb | 2023-11-06
22 | 22 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.78 | 3139.108 | 108007 | 3439.733 | 3.65 | 48.75 | 52.34 | 7.4 | bb | 2023-11-06
23 | 23 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.78 | 4940.688 | 170489 | 3919.615 | 5.042 | 78.125 | 72.91 | -6.7 | bb | 2023-11-06
24 | 24 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.78 | 7579.503 | 259401 | 3888.294 | 7.797 | 125 | 113.62 | -9.1 | bb | 2023-11-06
25 | 25 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 3.78 | 10693.694 | 369588 | 4835.771 | 8.845 | 200 | 129.11 | -35.4 | bb | 2023-11-06
26 | 26 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 3.78 | 18052.461 | 621276 | 5718.93 | 12.626 | 325.625 | 184.98 | -43.2 | bb | 2023-11-06
27 | 27 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.83 | 19.347 | 590 | 1.072 | 72.19 |  | 1065.19 |  | bb | 2023-11-06
28 | 28 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.82 | 17.326 | 404 | [5 empty] | bb | 2023-11-06
29 | 29 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 3.78 | 92.881 | 2263 | 2902.343 | 0.128 | 0.713 | 0.29 | -59.4 | bb | 2023-11-06
30 | 30 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 3.78 | 92.564 | 2777 | 2148.496 | 0.172 | 1.138 | 0.94 | -17 | bb | 2023-11-06
31 | 31 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.78 | 168.791 | 5259 | 2368.634 | 0.285 | 1.825 | 2.61 | 43 | bb | 2023-11-06
32 | 32 | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.82 | 4.34 | 147 | 0.567 | 30.617 |  | 450.84 |  | bb | 2023-11-06
33 | 33 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 3.78 | 88.341 | 2578 | 2884.729 | 0.122 | 0.713 | 0.21 | -70.8 | bb | 2023-11-06
34 | 34 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 3.78 | 114.997 | 4024 | 3311.456 | 0.139 | 1.138 | 0.45 | -60.4 | bb | 2023-11-06
35 | 35 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.78 | 141.014 | 4907 | 2974.036 | 0.19 | 1.825 | 1.2 | -34.2 | bb | 2023-11-06
36 | 36 | 20231106_HepCl_ReRun_036 | Blank | Blank | 3.82 | 14.094 | 247 | [5 empty] | bb | 2023-11-06
37 | 37 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 3.77 | 115.598 | 3475 | 2898.57 | 0.16 | 0.713 | 0.75 | 6 | bb | 2023-11-06
38 | 38 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 3.78 | 120.169 | 3550 | 2369.425 | 0.203 | 1.138 | 1.4 | 22.7 | bb | 2023-11-06
39 | 39 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.78 | 177.583 | 5236 | 2930.851 | 0.242 | 1.825 | 1.98 | 8.4 | bb | 2023-11-06
40 | 40 | 20231106_HepCl_ReRun_040 | Blank | Blank |  |  |  | 2.239 | [5 empty] | 2023-11-06
41 | 41 | 20231106_HepCl_ReRun_041 | QC A | QC | 3.78 | 94.863 | 3071 | 3334.889 | 0.114 | 0.625 | 0.08 | -87.4 | bb | 2023-11-06
42 | 42 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.78 | 191.108 | 5682 | 3127.347 | 0.244 | 2.506 | 2.01 | -19.8 | bb | 2023-11-06
43 | 43 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.78 | 449.471 | 14963 | 3371.759 | 0.533 | 6.269 | 6.28 | 0.1 | bb | 2023-11-06
44 | 44 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.78 | 1519.89 | 52572 | 3225.131 | 1.885 | 25 | 26.25 | 5 | bb | 2023-11-06
45 | 45 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.81 | 16.706 | 473 | [5 empty] | bb | 2023-11-06
46 | 46 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 3.81 | 26.536 | 537 | [5 empty] | bb | 2023-11-06
47 | 47 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 3.78 | 44.494 | 1585 | 2990.745 | 0.06 | 0.175 |  |  | bbI | 2023-11-06
48 | 48 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 3.78 | 89.744 | 2579 | 3253.131 | 0.11 | 0.281 | 0.03 | -90 | bb | 2023-11-06
49 | 49 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 3.78 | 79.804 | 2635 | 2768.087 | 0.115 | 0.444 | 0.1 | -77.1 | bb | 2023-11-06
50 | 50 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 3.78 | 111.645 | 3306 | 3070.196 | 0.145 | 0.713 | 0.55 | -23.2 | bb | 2023-11-06
51 | 51 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 3.78 | 159.707 | 4651 | 2763.465 | 0.231 | 1.138 | 1.81 | 59.4 | bb | 2023-11-06
52 | 52 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.78 | 157.768 | 4759 | 2982.801 | 0.212 | 1.825 | 1.52 | -16.5 | bb | 2023-11-06
53 | 53 | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 1.263 | [5 empty] | 2023-11-06
54 | 54 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.78 | 193.975 | 6727 | 3350.341 | 0.232 | 2.919 | 1.82 | -37.7 | bb | 2023-11-06
55 | 55 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.78 | 335.471 | 11180 | 2541.916 | 0.528 | 4.669 | 6.2 | 32.8 | bb | 2023-11-06
56 | 56 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.78 | 445.528 | 14562 | 3151.321 | 0.566 | 7.45 | 6.75 | -9.3 | bb | 2023-11-06
57 | 57 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.78 | 725.711 | 25169 | 3307.122 | 0.878 | 11.875 | 11.37 | -4.3 | bb | 2023-11-06
58 | 58 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.78 | 1252.312 | 42857 | 3103.845 | 1.614 | 19.125 | 22.25 | 16.3 | bb | 2023-11-06
59 | 59 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.78 | 2141.835 | 73548 | 3365.248 | 2.546 | 30.5 | 36.02 | 18.1 | bb | 2023-11-06
60 | 60 | 20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 1.55 | [5 empty] | 2023-11-06
61 | 61 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.78 | 3087.364 | 107125 | 3673.549 | 3.362 | 48.75 | 48.08 | -1.4 | bb | 2023-11-06
62 | 62 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.78 | 5185.91 | 177930 | 4066.942 | 5.101 | 78.125 | 73.77 | -5.6 | bb | 2023-11-06
63 | 63 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.78 | 8270.189 | 287977 | 3980.426 | 8.311 | 125 | 121.21 | -3 | bb | 2023-11-06
64 | 64 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 3.78 | 8570.126 | 291253 | 4885.239 | 7.017 | 200 | 102.09 | -49 | bb | 2023-11-06
65 | 65 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 3.78 | 20518.283 | 706717 | 6843.28 | 11.993 | 325.625 | 175.63 | -46.1 | bb | 2023-11-07
66 | 66 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.81 | 13.005 | 297 | 0.37 | 140.595 |  | 2076.03 |  | bb | 2023-11-07
67 | 67 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.328 | [5 empty] | 2023-11-07
68 | 68 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 3.78 | 117.315 | 3438 | 3113.354 | 0.151 | 0.713 | 0.62 | -12.3 | bb | 2023-11-07
69 | 69 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 3.78 | 117.511 | 3059 | 2642.521 | 0.178 | 1.138 | 1.03 | -9.8 | bb | 2023-11-07
70 | 70 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.78 | 156.739 | 4706 | 3064.16 | 0.205 | 1.825 | 1.42 | -22.1 | bb | 2023-11-07
71 | 71 | 20231106_HepCl_ReRun_071 | Blank | Blank |  |  |  | 5.703 | [5 empty] | 2023-11-07
72 | 72 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 3.78 | 94.069 | 2660 | 2424.344 | 0.155 | 0.713 | 0.69 | -3 | bb | 2023-11-07
73 | 73 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 3.78 | 162.413 | 5489 | 3320.192 | 0.196 | 1.138 | 1.29 | 13.3 | bb | 2023-11-07
74 | 74 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.78 | 145.769 | 5063 | 3262.376 | 0.179 | 1.825 | 1.04 | -43.1 | bb | 2023-11-07
75 | 75 | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.79 | 0.88 | 95 | 0.322 | 10.932 |  | 159.94 |  | bb | 2023-11-07
76 | 76 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 3.78 | 90.3 | 2912 | 3100.597 | 0.116 | 0.713 | 0.12 | -83.3 | bb | 2023-11-07
77 | 77 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 3.78 | 136.129 | 4459 | 3376.185 | 0.161 | 1.138 | 0.78 | -31.4 | bb | 2023-11-07
78 | 78 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.78 | 168.728 | 5698 | 3313.979 | 0.204 | 1.825 | 1.41 | -22.9 | bb | 2023-11-07
79 | 79 | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 4.597 | [5 empty] | 2023-11-07
80 | 80 | 20231106_HepCl_ReRun_080 | QC A | QC | 3.78 | 122.98 | 3746 | 3603.385 | 0.137 | 0.625 | 0.41 | -33.6 | bb | 2023-11-07
81 | 81 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.78 | 213.334 | 7274 | 2887.97 | 0.295 | 2.506 | 2.76 | 10.3 | bb | 2023-11-07
82 | 82 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.78 | 390.65 | 13088 | 3610.88 | 0.433 | 6.269 | 4.79 | -23.5 | bb | 2023-11-07
83 | 83 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.78 | 1626.875 | 56148 | 3571.287 | 1.822 | 25 | 25.32 | 1.3 | bb | 2023-11-07
84 | 84 | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.83 | 20.354 | 553 | [5 empty] | bb | 2023-11-07
85 | 85 | 20231106_HepCl_ReRun_085 | Blank | Blank | 3.83 | 1.805 | 126 | [5 empty] | bb | 2023-11-07
86 | 86 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 3.78 | 2152.861 | 74098 | 3342.74 | 2.576 | 30.5 | 36.47 | 19.6 | bb | 2023-11-07
87 | 87 | 20231106_HepCl_ReRun_087 | Blank | Blank | 3.82 | 12.764 | 267 | [5 empty] | bb | 2023-11-07
88 | 88 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.78 | 205.262 | 6517 | 4776.272 | 0.172 |  | 0.94 |  | bb | 2023-11-07
89 | 89 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.78 | 292.673 | 9773 | 3195.141 | 0.366 |  | 3.81 |  | bb | 2023-11-07
90 | 90 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.78 | 210.794 | 7030 | 4462.737 | 0.189 |  | 1.19 |  | bb | 2023-11-07
91 | 91 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.78 | 144.765 | 4211 | 4447.937 | 0.13 |  | 0.32 |  | bb | 2023-11-07
92 | 92 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.78 | 99.272 | 2968 | 4036.639 | 0.098 |  |  |  | bbI | 2023-11-07
93 | 93 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.77 | 136.658 | 4376 | 3155.708 | 0.173 |  | 0.96 |  | bb | 2023-11-07
94 | 94 | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.83 | 17.262 | 362 | 1.329 | 51.955 |  | 766.16 |  | bb | 2023-11-07
95 | 95 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.78 | 110.127 | 3275 | 4054.992 | 0.109 |  | 0 |  | bb | 2023-11-07
96 | 96 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.78 | 98.905 | 2884 | 4225.446 | 0.094 |  |  |  | bbI | 2023-11-07
97 | 97 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.77 | 130.767 | 3718 | 4254.767 | 0.123 |  | 0.21 |  | bb | 2023-11-07
98 | 98 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.78 | 137.994 | 4125 | 4188.782 | 0.132 |  | 0.34 |  | bb | 2023-11-07
99 | 99 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.78 | 140.797 | 4509 | 4496.683 | 0.125 |  | 0.25 |  | bb | 2023-11-07
100 | 100 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.78 | 121.203 | 3598 | 4153.806 | 0.117 |  | 0.12 |  | bb | 2023-11-07
101 | 101 | 20231106_HepCl_ReRun_101 | Blank | Blank | 3.82 | 6.098 | 182 | [5 empty] | bb | 2023-11-07
102 | 102 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.78 | 108.098 | 3082 | 3209.497 | 0.135 |  | 0.39 |  | bb | 2023-11-07
103 | 103 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.78 | 139.492 | 4379 | 3877.044 | 0.144 |  | 0.52 |  | bb | 2023-11-07
104 | 104 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.78 | 106.808 | 3014 | 3869.138 | 0.11 |  | 0.03 |  | bb | 2023-11-07
105 | 105 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.77 | 94.632 | 2824 | 3451.207 | 0.11 |  | 0.02 |  | bb | 2023-11-07
106 | 106 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.78 | 83.841 | 2740 | 3384.839 | 0.099 |  |  |  | bbI | 2023-11-07
107 | 107 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.78 | 78.491 | 2813 | 3411.54 | 0.092 |  |  |  | bbI | 2023-11-07
108 | 108 | 20231106_HepCl_ReRun_108 | Blank | Blank | 3.82 | 2.795 | 99 | [5 empty] | bb | 2023-11-07
109 | 109 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 3.78 | 2132.004 | 73766 | 3304.299 | 2.581 | 30.5 | 36.54 | 19.8 | bb | 2023-11-07
110 | 110 | 20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 3.311 | [5 empty] | 2023-11-07
111 | 111 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.78 | 372.564 | 13071 | 3946.084 | 0.378 |  | 3.98 |  | bb | 2023-11-07
112 | 112 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.78 | 110.375 | 3443 | 4015.599 | 0.11 |  | 0.02 |  | bb | 2023-11-07
113 | 113 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.78 | 78.054 | 2592 | 3722.044 | 0.084 |  |  |  | bbI | 2023-11-07
114 | 114 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.78 | 138.96 | 4052 | 3719.442 | 0.149 |  | 0.61 |  | bb | 2023-11-07
115 | 115 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.78 | 126.729 | 4009 | 3898.412 | 0.13 |  | 0.32 |  | bb | 2023-11-07
116 | 116 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.78 | 122.661 | 3841 | 3840.235 | 0.128 |  | 0.29 |  | bb | 2023-11-07
117 | 117 | 20231106_HepCl_ReRun_117 | Blank | Blank | [9 empty] | 2023-11-07
118 | 118 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.78 | 115.056 | 3650 | 3796.952 | 0.121 |  | 0.19 |  | bb | 2023-11-07
119 | 119 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.77 | 98.341 | 3529 | 3988.584 | 0.099 |  |  |  | bbI | 2023-11-07
120 | 120 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.78 | 149.401 | 5058 | 4089.779 | 0.146 |  | 0.56 |  | bb | 2023-11-07
121 | 121 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.78 | 185.739 | 6213 | 4257.981 | 0.174 |  | 0.98 |  | bb | 2023-11-07
122 | 122 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.78 | 297.333 | 9442 | 4107.876 | 0.29 |  | 2.68 |  | bb | 2023-11-07
123 | 123 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.78 | 246.998 | 8342 | 3864.161 | 0.256 |  | 2.18 |  | bb | 2023-11-07
124 | 124 | 20231106_HepCl_ReRun_124 | Blank | Blank | 3.82 | 10.597 | 266 | [5 empty] | bb | 2023-11-07
125 | 125 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.78 | 86.154 | 2791 | 3543.577 | 0.097 |  |  |  | bbI | 2023-11-07
126 | 126 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.78 | 115.902 | 3813 | 3555.943 | 0.13 |  | 0.32 |  | bb | 2023-11-07
127 | 127 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.78 | 137.6 | 4290 | 3988.081 | 0.138 |  | 0.44 |  | bb | 2023-11-07
128 | 128 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.78 | 56.173 | 1957 | 3506.183 | 0.064 |  |  |  | bbI | 2023-11-07
129 | 129 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.77 | 80.953 | 2729 | 2791.078 | 0.116 |  | 0.11 |  | bb | 2023-11-07
130 | 130 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.78 | 62.378 | 2190 | 3117.41 | 0.08 |  |  |  | bbI | 2023-11-07
131 | 131 | 20231106_HepCl_ReRun_131 | Blank | Blank |  |  |  | 1.14 | [5 empty] | 2023-11-07
132 | 132 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 3.78 | 2330.145 | 80404 | 3604.25 | 2.586 | 30.5 | 36.61 | 20 | bb | 2023-11-07
133 | 133 | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.81 | 10.285 | 369 | 0.164 | 250.854 |  | 3705.38 |  | bb | 2023-11-07
134 | 134 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.78 | 189.653 | 5817 | 4056.419 | 0.187 |  | 1.16 |  | bb | 2023-11-07
135 | 135 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.78 | 241.631 | 8170 | 4559.46 | 0.212 |  | 1.53 |  | bb | 2023-11-07
136 | 136 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.78 | 290.765 | 9335 | 4493.075 | 0.259 |  | 2.22 |  | bb | 2023-11-07
137 | 137 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.78 | 190.399 | 6349 | 4114.258 | 0.185 |  | 1.13 |  | bb | 2023-11-07
138 | 138 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.78 | 243.066 | 7101 | 4004.459 | 0.243 |  | 1.99 |  | bb | 2023-11-07
139 | 139 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.78 | 176.193 | 5558 | 3985.281 | 0.177 |  | 1.01 |  | bb | 2023-11-07
140 | 140 | 20231106_HepCl_ReRun_140 | Blank | Blank | 3.81 | 26.678 | 650 | [5 empty] | bb | 2023-11-07
141 | 141 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 3.78 | 2373.712 | 81651 | 3653.36 | 2.599 | 30.5 | 36.8 | 20.7 | bb | 2023-11-07
142 | 142 | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.81 | 20.394 | 384 | 2.523 | 32.333 |  | 476.2 |  | bb | 2023-11-07
143 | 143 | 20231106_HepCl_ReRun_143 | Blank | Blank |  |  |  | 5.162 | [5 empty] | 2023-11-07
144 | 144 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.78 | 6391.303 | 220071 | 5093.202 | 5.019 |  | 72.57 |  | bb | 2023-11-07
145 | 145 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.78 | 5906.037 | 203264 | 4939.262 | 4.783 |  | 69.08 |  | bb | 2023-11-07
146 | 146 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.78 | 6012.225 | 207327 | 5233.792 | 4.595 |  | 66.3 |  | bb | 2023-11-07
147 | 147 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.78 | 4942.229 | 169810 | 4518.646 | 4.375 |  | 63.05 |  | bb | 2023-11-07
148 | 148 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.78 | 4874.581 | 169161 | 4595.066 | 4.243 |  | 61.1 |  | bb | 2023-11-07
149 | 149 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.78 | 3923.402 | 133112 | 4516.958 | 3.474 |  | 49.74 |  | bb | 2023-11-07
150 | 150 | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.81 | 16.621 | 419 | 0.579 | 114.826 |  | 1695.23 |  | bb | 2023-11-07
151 | 151 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.78 | 4543.188 | 154753 | 4562.208 | 3.983 |  | 57.26 |  | bb | 2023-11-07
152 | 152 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.78 | 4894.958 | 166429 | 4277.345 | 4.578 |  | 66.04 |  | bb | 2023-11-07
153 | 153 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.78 | 4958.275 | 168396 | 4028.417 | 4.923 |  | 71.15 |  | bb | 2023-11-07
154 | 154 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.78 | 5418.749 | 187049 | 4201.701 | 5.159 |  | 74.63 |  | bb | 2023-11-07
155 | 155 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.78 | 5020.021 | 174032 | 4522.662 | 4.44 |  | 64.01 |  | bb | 2023-11-07
156 | 156 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.78 | 5217.081 | 179277 | 4864.867 | 4.29 |  | 61.79 |  | bb | 2023-11-07
157 | 157 | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.8 | 12.358 | 293 | [5 empty] | bb | 2023-11-07
158 | 158 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.78 | 4317.092 | 148947 | 4130.82 | 4.18 |  | 60.17 |  | bb | 2023-11-07
159 | 159 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.78 | 5081.018 | 175617 | 4509.017 | 4.507 |  | 65.01 |  | bb | 2023-11-07
160 | 160 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.78 | 4593.416 | 157865 | 4377.103 | 4.198 |  | 60.43 |  | bb | 2023-11-07
161 | 161 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.78 | 3518.875 | 119857 | 3747.828 | 3.756 |  | 53.9 |  | bb | 2023-11-07
162 | 162 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.78 | 3942.875 | 135332 | 4038.739 | 3.905 |  | 56.1 |  | bb | 2023-11-07
163 | 163 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.78 | 3520.303 | 121928 | 3789.513 | 3.716 |  | 53.31 |  | bb | 2023-11-07
164 | 164 | 20231106_HepCl_ReRun_166 | Blank | Blank | 3.78 | 0.364 | 40 | 0.413 | 3.525 |  | 50.49 |  | bb | 2023-11-07
165 | 165 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 3.78 | 1761.065 | 59448 | 3208.256 | 2.196 | 30.5 | 30.84 | 1.1 | bb | 2023-11-07
166 | 166 | 20231106_HepCl_ReRun_168 | Blank | Blank | 3.82 | 10.882 | 381 | 2.772 | 15.703 |  | 230.44 |  | bb | 2023-11-07
167 | 167 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.78 | 4863.822 | 167123 | 4403.589 | 4.418 |  | 63.69 |  | bb | 2023-11-07
168 | 168 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.78 | 5016.161 | 172989 | 4598.22 | 4.364 |  | 62.88 |  | bb | 2023-11-07
169 | 169 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.78 | 5046.894 | 172993 | 4689.85 | 4.305 |  | 62.01 |  | bb | 2023-11-07
170 | 170 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.78 | 4255.095 | 146051 | 4132.267 | 4.119 |  | 59.26 |  | bb | 2023-11-07
171 | 171 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.78 | 3951.575 | 134733 | 4125.632 | 3.831 |  | 55.01 |  | bb | 2023-11-07
172 | 172 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.78 | 4628.658 | 158950 | 4422.781 | 4.186 |  | 60.26 |  | bb | 2023-11-07
173 | 173 | 20231106_HepCl_ReRun_175 | Blank | Blank | 3.82 | 11.069 | 300 | [5 empty] | bb | 2023-11-07
174 | 174 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.78 | 4871.632 | 168775 | 4699.705 | 4.146 |  | 59.67 |  | bb | 2023-11-07
175 | 175 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.78 | 4619.891 | 158711 | 4633.228 | 3.988 |  | 57.34 |  | bb | 2023-11-07
176 | 176 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.78 | 5559.856 | 190760 | 4931.43 | 4.51 |  | 65.04 |  | bb | 2023-11-07
177 | 177 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.78 | 4127.725 | 143210 | 4408.869 | 3.745 |  | 53.74 |  | bb | 2023-11-07
178 | 178 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.78 | 4334.655 | 148483 | 4910.336 | 3.531 |  | 50.58 |  | bb | 2023-11-07
179 | 179 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.78 | 4811.418 | 164882 | 4659.491 | 4.13 |  | 59.43 |  | bb | 2023-11-07
180 | 180 | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.81 | 0.8 | 62 | 1.668 | 1.918 |  | 26.75 |  | bb | 2023-11-07
181 | 181 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.78 | 4283.556 | 146551 | 3744.223 | 4.576 |  | 66.02 |  | bb | 2023-11-07
182 | 182 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.78 | 4557.383 | 156709 | 4034.394 | 4.519 |  | 65.17 |  | bb | 2023-11-07
183 | 183 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.78 | 4135.498 | 140199 | 4172.714 | 3.964 |  | 56.98 |  | bb | 2023-11-07
184 | 184 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.78 | 3757.041 | 129864 | 3922.801 | 3.831 |  | 55.01 |  | bb | 2023-11-07
185 | 185 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.78 | 3527.55 | 121269 | 3270.064 | 4.315 |  | 62.16 |  | bb | 2023-11-07
186 | 186 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.78 | 4092.351 | 141067 | 4049.156 | 4.043 |  | 58.14 |  | bb | 2023-11-07
187 | 187 | 20231106_HepCl_ReRun_189 | Blank | Blank | 3.81 | 19.428 | 557 | [5 empty] | bb | 2023-11-07
188 | 188 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 3.78 | 2071.804 | 70999 | 3623.284 | 2.287 | 30.5 | 32.2 | 5.6 | bb | 2023-11-07
189 | 189 | 20231106_HepCl_ReRun_191 | Blank | Blank | [9 empty] | 2023-11-07
190 | 190 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.78 | 4634.335 | 157508 | 4646.139 | 3.99 |  | 57.36 |  | bb | 2023-11-07
191 | 191 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.78 | 4617.606 | 157457 | 4670.868 | 3.954 |  | 56.83 |  | bb | 2023-11-07
192 | 192 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.78 | 5855.985 | 201916 | 4760.106 | 4.921 |  | 71.12 |  | bb | 2023-11-07
193 | 193 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.78 | 4184.689 | 143212 | 4087.066 | 4.096 |  | 58.92 |  | bb | 2023-11-07
194 | 194 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.78 | 4455.831 | 152479 | 4565.343 | 3.904 |  | 56.09 |  | bb | 2023-11-07
195 | 195 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.78 | 4756.871 | 162191 | 4533.528 | 4.197 |  | 60.42 |  | bb | 2023-11-07
196 | 196 | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.81 | 18.713 | 586 | [5 empty] | bb | 2023-11-08
197 | 197 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 3.78 | 2376.775 | 81275 | 2966.934 | 3.204 | 30.5 | 45.75 | 50 | bb | 2023-11-08
198 | 198 | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.82 | 9.87 | 225 | 0.525 | 75.2 |  | 1109.66 |  | bb | 2023-11-08
199 | 199 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.78 | 65.511 | 2011 | 4314.995 | 0.061 |  |  |  | bbI | 2023-11-08
200 | 200 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.78 | 90.844 | 2571 | 4169.22 | 0.087 |  |  |  | bbI | 2023-11-08
201 | 201 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.78 | 59.138 | 1829 | 4343.747 | 0.054 |  |  |  | bbI | 2023-11-08
202 | 202 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.78 | 42.996 | 1394 | 3994.895 | 0.043 |  |  |  | bbI | 2023-11-08
203 | 203 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.78 | 39.354 | 1325 | 3869.739 | 0.041 |  |  |  | bbI | 2023-11-08
204 | 204 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.78 | 65.363 | 1799 | 4094.355 | 0.064 |  |  |  | bbI | 2023-11-08
205 | 205 | 20231106_HepCl_ReRun_207 | Blank | Blank | 3.81 | 22.558 | 431 | [5 empty] | bb | 2023-11-08
206 | 206 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.78 | 164.218 | 5442 | 4088.134 | 0.161 |  | 0.77 |  | bb | 2023-11-08
207 | 207 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.78 | 66.106 | 1966 | 3656.892 | 0.072 |  |  |  | bbI | 2023-11-08
208 | 208 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.77 | 67.688 | 2189 | 4115.619 | 0.066 |  |  |  | bbI | 2023-11-08
209 | 209 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.78 | 58.988 | 1991 | 3668.877 | 0.064 |  |  |  | bbI | 2023-11-08
210 | 210 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.78 | 114.769 | 3960 | 3204.234 | 0.143 |  | 0.51 |  | bb | 2023-11-08
211 | 211 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.77 | 205.511 | 7108 | 3915.398 | 0.21 |  | 1.5 |  | bb | 2023-11-08
212 | 212 | 20231106_HepCl_ReRun_214 | Blank | Blank | 3.82 | 21.328 | 482 | [5 empty] | bb | 2023-11-08
213 | 213 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.78 | 229.48 | 8090 | 3898.605 | 0.235 |  | 1.88 |  | bb | 2023-11-08
214 | 214 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.78 | 54.196 | 1918 | 4061.239 | 0.053 |  |  |  | bbI | 2023-11-08
215 | 215 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.78 | 73.433 | 2244 | 3517.021 | 0.084 |  |  |  | bbI | 2023-11-08
216 | 216 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.78 | 29.845 | 1106 | 3202.326 | 0.037 |  |  |  | bbI | 2023-11-08
217 | 217 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.78 | 51.159 | 1557 | 3661.565 | 0.056 |  |  |  | bbI | 2023-11-08
218 | 218 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.78 | 58.809 | 1433 | 2804.06 | 0.084 |  |  |  | bbI | 2023-11-08
219 | 219 | 20231106_HepCl_ReRun_221 | Blank | Blank | [9 empty] | 2023-11-08
220 | 220 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 3.78 | 2442.702 | 84129 | 3886.825 | 2.514 | 30.5 | 35.55 | 16.5 | bb | 2023-11-08
221 | 221 | 20231106_HepCl_ReRun_223 | Blank | Blank | 3.82 | 17.31 | 342 | [5 empty] | bb | 2023-11-08
222 | 222 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.78 | 77.001 | 2803 | 4173.079 | 0.074 |  |  |  | bbI | 2023-11-08
223 | 223 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.78 | 41.344 | 1442 | 3930.435 | 0.042 |  |  |  | bbI | 2023-11-08
224 | 224 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.77 | 39.448 | 947 | 3689.984 | 0.043 |  |  |  | bbI | 2023-11-08
225 | 225 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.78 | 74.386 | 1977 | 3838.459 | 0.078 |  |  |  | bbI | 2023-11-08
226 | 226 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.77 | 43.105 | 1522 | 3972.163 | 0.043 |  |  |  | bbI | 2023-11-08
227 | 227 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.78 | 58.642 | 2012 | 3842.832 | 0.061 |  |  |  | bbI | 2023-11-08
228 | 228 | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.81 | 5.699 | 139 | 4.215 | 5.408 |  | 78.32 |  | bb | 2023-11-08
229 | 229 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.77 | 87.68 | 2747 | 3893.189 | 0.09 |  |  |  | bbI | 2023-11-08
230 | 230 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.78 | 85.112 | 3077 | 3559.379 | 0.096 |  |  |  | bbI | 2023-11-08
231 | 231 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.78 | 56.909 | 2038 | 4010.53 | 0.057 |  |  |  | bbI | 2023-11-08
232 | 232 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.78 | 6827.96 | 236140 | 4683.644 | 5.831 |  | 84.57 |  | bb | 2023-11-08
233 | 233 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.78 | 217.342 | 7810 | 3473.603 | 0.25 |  | 2.1 |  | bb | 2023-11-08
234 | 234 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.78 | 2704.803 | 93865 | 3760.475 | 2.877 |  | 40.91 |  | bb | 2023-11-08
235 | 235 | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 1.939 | [5 empty] | 2023-11-08
236 | 236 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.78 | 79.311 | 2437 | 3804.575 | 0.083 |  |  |  | bbI | 2023-11-08
237 | 237 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.78 | 63.602 | 1984 | 3650.367 | 0.07 |  |  |  | bbI | 2023-11-08
238 | 238 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.78 | 57.151 | 1930 | 4028.423 | 0.057 |  |  |  | bbI | 2023-11-08
239 | 239 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.78 | 34.578 | 1251 | 3605.647 | 0.038 |  |  |  | bbI | 2023-11-08
240 | 240 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.78 | 55.165 | 1589 | 3280.838 | 0.067 |  |  |  | bbI | 2023-11-08
241 | 241 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.78 | 42.574 | 1442 | 3179.154 | 0.054 |  |  |  | bbI | 2023-11-08
242 | 242 | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.82 | 13.663 | 294 | 0.26 | 210.2 |  | 3104.62 |  | bb | 2023-11-08
243 | 243 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 3.78 | 2361.131 | 81590 | 3753.745 | 2.516 | 30.5 | 35.58 | 16.6 | bb | 2023-11-08
244 | 244 | 20231106_HepCl_ReRun_246 | Blank | Blank | [9 empty] | 2023-11-08
245 | 245 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.78 | 110.167 | 3794 | 4269.086 | 0.103 |  |  |  | bbI | 2023-11-08
246 | 246 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.78 | 102.365 | 3701 | 4165.854 | 0.098 |  |  |  | bbI | 2023-11-08
247 | 247 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.78 | 144.584 | 5057 | 4105.743 | 0.141 |  | 0.48 |  | bb | 2023-11-08
248 | 248 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.78 | 301.788 | 9552 | 4091.521 | 0.295 |  | 2.76 |  | bb | 2023-11-08
249 | 249 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.78 | 299.284 | 9875 | 3699.73 | 0.324 |  | 3.18 |  | bb | 2023-11-08
250 | 250 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.78 | 306.603 | 9765 | 3781.379 | 0.324 |  | 3.19 |  | bb | 2023-11-08
251 | 251 | 20231106_HepCl_ReRun_253 | Blank | Blank | 3.81 | 8.129 | 294 | [5 empty] | bb | 2023-11-08
252 | 252 | 20231106_HepCl_ReRun_254 | Blank | Blank | 3.81 | 13.784 | 295 | [5 empty] | bb | 2023-11-08
253 | 253 | 20231106_HepCl_ReRun_255 | Blank | Blank | 3.82 | 17.656 | 487 | [5 empty] | bb | 2023-11-08
Compound 3: M2-6:2 FTS
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20231106_HepCl_ReRun_001 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
2 | 2 | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.79 | 2.377 | 87 |  | 2.377 | 4 | 0 | -99.9 | bb | 2023-11-06
3 | 3 | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.81 | 0.237 | 33 |  | 0.237 | 4 | 0 | -100 | bb | 2023-11-06
4 | 4 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.81 | 8.997 | 224 |  | 8.997 | 4 | 0.01 | -99.7 | bb | 2023-11-06
5 | 5 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
6 | 6 | 20231106_HepCl_ReRun_006 | Blank | Blank | 3.82 | 8.596 | 316 |  | 8.596 | 4 | 0.01 | -99.7 | bb | 2023-11-06
7 | 7 | 20231106_HepCl_ReRun_007 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
8 | 8 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 3.78 | 3119.796 | 107592 |  | 3119.796 | 4 | 3.86 | -3.6 | bb | 2023-11-06
9 | 9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 3.78 | 3010.686 | 103804 |  | 3010.686 | 4 | 3.72 | -6.9 | bb | 2023-11-06
10 | 10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 3.77 | 2911.38 | 99233 |  | 2911.38 | 4 | 3.6 | -10 | bb | 2023-11-06
11 | 11 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 3.78 | 3209.637 | 111137 |  | 3209.637 | 4 | 3.97 | -0.8 | bb | 2023-11-06
12 | 12 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 3.77 | 3047.267 | 103322 |  | 3047.267 | 4 | 3.77 | -5.8 | bb | 2023-11-06
13 | 13 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.77 | 3034.11 | 104487 |  | 3034.11 | 4 | 3.75 | -6.2 | bb | 2023-11-06
14 | 14 | 20231106_HepCl_ReRun_014 | Blank | Blank | 3.74 | 3.04 | 110 |  | 3.04 | 4 | 0 | -99.9 | bb | 2023-11-06
15 | 15 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.78 | 2993.711 | 101914 |  | 2993.711 | 4 | 3.7 | -7.5 | bb | 2023-11-06
16 | 16 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.78 | 3101.963 | 107129 |  | 3101.963 | 4 | 3.84 | -4.1 | bb | 2023-11-06
17 | 17 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.78 | 3138.186 | 108340 |  | 3138.186 | 4 | 3.88 | -3 | bb | 2023-11-06
18 | 18 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.77 | 3163.276 | 108367 |  | 3163.276 | 4 | 3.91 | -2.2 | bb | 2023-11-06
19 | 19 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.78 | 3229.261 | 110694 |  | 3229.261 | 4 | 3.99 | -0.2 | bb | 2023-11-06
20 | 20 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.77 | 2767.938 | 94438 |  | 2767.938 | 4 | 3.42 | -14.4 | bb | 2023-11-06
21 | 21 | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.82 | 2.584 | 93 |  | 2.584 | 4 | 0 | -99.9 | bb | 2023-11-06
22 | 22 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.78 | 3439.733 | 118429 |  | 3439.733 | 4 | 4.25 | 6.3 | bb | 2023-11-06
23 | 23 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.78 | 3919.615 | 134613 |  | 3919.615 | 4 | 4.85 | 21.2 | bb | 2023-11-06
24 | 24 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.78 | 3888.294 | 133357 |  | 3888.294 | 4 | 4.81 | 20.2 | bb | 2023-11-06
25 | 25 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 3.78 | 4835.771 | 166839 |  | 4835.771 | 4 | 5.98 | 49.5 | bb | 2023-11-06
26 | 26 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 3.78 | 5718.93 | 196763 |  | 5718.93 | 4 | 7.07 | 76.8 | bb | 2023-11-06
27 | 27 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 3.81 | 1.072 | 39 |  | 1.072 | 4 | 0 | -100 | bb | 2023-11-06
28 | 28 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
29 | 29 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 3.78 | 2902.343 | 99894 |  | 2902.343 | 4 | 3.59 | -10.3 | bb | 2023-11-06
30 | 30 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 3.78 | 2148.496 | 73093 |  | 2148.496 | 4 | 2.66 | -33.6 | bb | 2023-11-06
31 | 31 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.78 | 2368.634 | 80581 |  | 2368.634 | 4 | 2.93 | -26.8 | bb | 2023-11-06
32 | 32 | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.8 | 0.567 | 13 |  | 0.567 | 4 | 0 | -100 | bb | 2023-11-06
33 | 33 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 3.78 | 2884.729 | 98797 |  | 2884.729 | 4 | 3.57 | -10.8 | bb | 2023-11-06
34 | 34 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 3.78 | 3311.456 | 113680 |  | 3311.456 | 4 | 4.09 | 2.4 | bb | 2023-11-06
35 | 35 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.77 | 2974.036 | 102368 |  | 2974.036 | 4 | 3.68 | -8.1 | bb | 2023-11-06
36 | 36 | 20231106_HepCl_ReRun_036 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
37 | 37 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 3.77 | 2898.57 | 98875 |  | 2898.57 | 4 | 3.58 | -10.4 | bb | 2023-11-06
38 | 38 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 3.78 | 2369.425 | 80566 |  | 2369.425 | 4 | 2.93 | -26.8 | bb | 2023-11-06
39 | 39 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.78 | 2930.851 | 99746 |  | 2930.851 | 4 | 3.62 | -9.4 | bb | 2023-11-06
40 | 40 | 20231106_HepCl_ReRun_040 | Blank | Blank | 3.75 | 2.239 | 81 |  | 2.239 | 4 | 0 | -99.9 | bb | 2023-11-06
41 | 41 | 20231106_HepCl_ReRun_041 | QC A | QC | 3.78 | 3334.889 | 113286 |  | 3334.889 | 4 | 4.12 | 3.1 | bb | 2023-11-06
42 | 42 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.78 | 3127.347 | 113458 |  | 3127.347 | 4 | 3.87 | -3.3 | MM | 2023-11-06
43 | 43 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.78 | 3371.759 | 115251 |  | 3371.759 | 4 | 4.17 | 4.2 | bb | 2023-11-06
44 | 44 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.78 | 3225.131 | 110928 |  | 3225.131 | 4 | 3.99 | -0.3 | bb | 2023-11-06
45 | 45 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
46 | 46 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
47 | 47 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 3.78 | 2990.745 | 102310 |  | 2990.745 | 4 | 3.7 | -7.6 | bb | 2023-11-06
48 | 48 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 3.78 | 3253.131 | 112734 |  | 3253.131 | 4 | 4.02 | 0.6 | bb | 2023-11-06
49 | 49 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 3.78 | 2768.087 | 94918 |  | 2768.087 | 4 | 3.42 | -14.4 | bb | 2023-11-06
50 | 50 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 3.78 | 3070.196 | 107053 |  | 3070.196 | 4 | 3.8 | -5.1 | bb | 2023-11-06
51 | 51 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 3.78 | 2763.465 | 93414 |  | 2763.465 | 4 | 3.42 | -14.6 | bb | 2023-11-06
52 | 52 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.78 | 2982.801 | 103645 |  | 2982.801 | 4 | 3.69 | -7.8 | bb | 2023-11-06
53 | 53 | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.79 | 1.263 | 40 |  | 1.263 | 4 | 0 | -100 | bb | 2023-11-06
54 | 54 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.77 | 3350.341 | 114845 |  | 3350.341 | 4 | 4.14 | 3.6 | bb | 2023-11-06
55 | 55 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.78 | 2541.916 | 86105 |  | 2541.916 | 4 | 3.14 | -21.4 | bb | 2023-11-06
56 | 56 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.78 | 3151.321 | 108098 |  | 3151.321 | 4 | 3.9 | -2.6 | bb | 2023-11-06
57 | 57 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.78 | 3307.122 | 114196 |  | 3307.122 | 4 | 4.09 | 2.2 | bb | 2023-11-06
58 | 58 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.78 | 3103.845 | 107523 |  | 3103.845 | 4 | 3.84 | -4.1 | bb | 2023-11-06
59 | 59 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.77 | 3365.248 | 115566 |  | 3365.248 | 4 | 4.16 | 4 | bb | 2023-11-06
60 | 60 | 20231106_HepCl_ReRun_060 | Blank | Blank | 3.76 | 1.55 | 56 |  | 1.55 | 4 | 0 | -100 | bb | 2023-11-06
61 | 61 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.78 | 3673.549 | 126337 |  | 3673.549 | 4 | 4.54 | 13.6 | bb | 2023-11-06
62 | 62 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.78 | 4066.942 | 139930 |  | 4066.942 | 4 | 5.03 | 25.7 | bb | 2023-11-06
63 | 63 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.78 | 3980.426 | 148683 |  | 3980.426 | 4 | 4.92 | 23 | MM | 2023-11-06
64 | 64 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 3.78 | 4885.239 | 167128 |  | 4885.239 | 4 | 6.04 | 51 | bb | 2023-11-06
65 | 65 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 3.78 | 6843.28 | 238203 |  | 6843.28 | 4 | 8.46 | 111.5 | bb | 2023-11-07
66 | 66 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.81 | 0.37 | 26 |  | 0.37 | 4 | 0 | -100 | bb | 2023-11-07
67 | 67 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 3.81 | 1.328 | 48 |  | 1.328 | 4 | 0 | -100 | bb | 2023-11-07
68 | 68 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 3.78 | 3113.354 | 106715 |  | 3113.354 | 4 | 3.85 | -3.8 | bb | 2023-11-07
69 | 69 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 3.78 | 2642.521 | 90862 |  | 2642.521 | 4 | 3.27 | -18.3 | bb | 2023-11-07
70 | 70 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.78 | 3064.16 | 104987 |  | 3064.16 | 4 | 3.79 | -5.3 | bb | 2023-11-07
71 | 71 | 20231106_HepCl_ReRun_071 | Blank | Blank | 3.79 | 5.703 | 133 |  | 5.703 | 4 | 0.01 | -99.8 | bb | 2023-11-07
72 | 72 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 3.78 | 2424.344 | 82779 |  | 2424.344 | 4 | 3 | -25.1 | bb | 2023-11-07
73 | 73 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 3.78 | 3320.192 | 115358 |  | 3320.192 | 4 | 4.11 | 2.6 | bb | 2023-11-07
74 | 74 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.77 | 3262.376 | 111380 |  | 3262.376 | 4 | 4.03 | 0.8 | bb | 2023-11-07
75 | 75 | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.81 | 0.322 | 12 |  | 0.322 | 4 | 0 | -100 | bb | 2023-11-07
76 | 76 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 3.78 | 3100.597 | 106576 |  | 3100.597 | 4 | 3.83 | -4.2 | bb | 2023-11-07
77 | 77 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 3.78 | 3376.185 | 117099 |  | 3376.185 | 4 | 4.17 | 4.4 | bb | 2023-11-07
78 | 78 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.77 | 3313.979 | 112946 |  | 3313.979 | 4 | 4.1 | 2.4 | bb | 2023-11-07
79 | 79 | 20231106_HepCl_ReRun_079 | Blank | Blank | 3.81 | 4.597 | 132 |  | 4.597 | 4 | 0.01 | -99.9 | bb | 2023-11-07
80 | 80 | 20231106_HepCl_ReRun_080 | QC A | QC | 3.78 | 3603.385 | 124910 |  | 3603.385 | 4 | 4.46 | 11.4 | bb | 2023-11-07
81 | 81 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.77 | 2887.97 | 97800 |  | 2887.97 | 4 | 3.57 | -10.7 | bb | 2023-11-07
82 | 82 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.78 | 3610.88 | 124330 |  | 3610.88 | 4 | 4.46 | 11.6 | bb | 2023-11-07
83 | 83 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.78 | 3571.287 | 123058 |  | 3571.287 | 4 | 4.42 | 10.4 | bb | 2023-11-07
84 | 84 | 20231106_HepCl_ReRun_084 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
85 | 85 | 20231106_HepCl_ReRun_085 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
86 | 86 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 3.77 | 3342.74 | 115401 |  | 3342.74 | 4 | 4.13 | 3.3 | bb | 2023-11-07
87 | 87 | 20231106_HepCl_ReRun_087 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
88 | 88 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.78 | 4776.272 | 164687 |  | 4776.272 | 4 | 5.91 | 47.6 | bb | 2023-11-07
89 | 89 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.77 | 3195.141 | 107465 |  | 3195.141 | 4 | 3.95 | -1.2 | bb | 2023-11-07
90 | 90 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.78 | 4462.737 | 153013 |  | 4462.737 | 4 | 5.52 | 37.9 | bb | 2023-11-07
91 | 91 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.77 | 4447.937 | 152114 |  | 4447.937 | 4 | 5.5 | 37.5 | bb | 2023-11-07
92 | 92 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.78 | 4036.639 | 139623 |  | 4036.639 | 4 | 4.99 | 24.8 | bb | 2023-11-07
93 | 93 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.78 | 3155.708 | 105566 |  | 3155.708 | 4 | 3.9 | -2.5 | bb | 2023-11-07
94 | 94 | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.81 | 1.329 | 48 |  | 1.329 | 4 | 0 | -100 | bb | 2023-11-07
95 | 95 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.78 | 4054.992 | 139514 |  | 4054.992 | 4 | 5.01 | 25.3 | bb | 2023-11-07
96 | 96 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.78 | 4225.446 | 146406 |  | 4225.446 | 4 | 5.22 | 30.6 | bb | 2023-11-07
97 | 97 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.78 | 4254.767 | 145612 |  | 4254.767 | 4 | 5.26 | 31.5 | bb | 2023-11-07
98 | 98 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.78 | 4188.782 | 143659 |  | 4188.782 | 4 | 5.18 | 29.5 | bb | 2023-11-07
99 | 99 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.78 | 4496.683 | 155205 |  | 4496.683 | 4 | 5.56 | 39 | bb | 2023-11-07
100 | 100 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.78 | 4153.806 | 139829 |  | 4153.806 | 4 | 5.14 | 28.4 | bb | 2023-11-07
101 | 101 | 20231106_HepCl_ReRun_101 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
102 | 102 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.78 | 3209.497 | 108432 |  | 3209.497 | 4 | 3.97 | -0.8 | bb | 2023-11-07
103 | 103 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.78 | 3877.044 | 131675 |  | 3877.044 | 4 | 4.79 | 19.8 | bb | 2023-11-07
104 | 104 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.78 | 3869.138 | 133521 |  | 3869.138 | 4 | 4.78 | 19.6 | bb | 2023-11-07
105 | 105 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.78 | 3451.207 | 117553 |  | 3451.207 | 4 | 4.27 | 6.7 | bb | 2023-11-07
106 | 106 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.77 | 3384.839 | 115881 |  | 3384.839 | 4 | 4.19 | 4.6 | bb | 2023-11-07
107 | 107 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.78 | 3411.54 | 117593 |  | 3411.54 | 4 | 4.22 | 5.5 | bb | 2023-11-07
108 | 108 | 20231106_HepCl_ReRun_108 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
109 | 109 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 3.78 | 3304.299 | 114603 |  | 3304.299 | 4 | 4.09 | 2.1 | bb | 2023-11-07
110 | 110 | 20231106_HepCl_ReRun_110 | Blank | Blank | 3.81 | 3.311 | 119 |  | 3.311 | 4 | 0 | -99.9 | bb | 2023-11-07
111 | 111 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.78 | 3946.084 | 135784 |  | 3946.084 | 4 | 4.88 | 22 | bb | 2023-11-07
112 | 112 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.78 | 4015.599 | 139237 |  | 4015.599 | 4 | 4.96 | 24.1 | bb | 2023-11-07
113 | 113 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.78 | 3722.044 | 127148 |  | 3722.044 | 4 | 4.6 | 15.1 | bb | 2023-11-07
114 | 114 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.78 | 3719.442 | 129467 |  | 3719.442 | 4 | 4.6 | 15 | bb | 2023-11-07
115 | 115 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.78 | 3898.412 | 134486 |  | 3898.412 | 4 | 4.82 | 20.5 | bb | 2023-11-07
116 | 116 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.78 | 3840.235 | 132810 |  | 3840.235 | 4 | 4.75 | 18.7 | bb | 2023-11-07
117 | 117 | 20231106_HepCl_ReRun_117 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
118 | 118 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.78 | 3796.952 | 130307 |  | 3796.952 | 4 | 4.69 | 17.4 | bb | 2023-11-07
119 | 119 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.77 | 3988.584 | 136537 |  | 3988.584 | 4 | 4.93 | 23.3 | bb | 2023-11-07
120 | 120 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.77 | 4089.779 | 139738 |  | 4089.779 | 4 | 5.06 | 26.4 | bb | 2023-11-07
121 | 121 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.78 | 4257.981 | 146948 |  | 4257.981 | 4 | 5.26 | 31.6 | bb | 2023-11-07
122 | 122 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.77 | 4107.876 | 139990 |  | 4107.876 | 4 | 5.08 | 27 | bb | 2023-11-07
123 | 123 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.78 | 3864.161 | 132419 |  | 3864.161 | 4 | 4.78 | 19.4 | bb | 2023-11-07
124 | 124 | 20231106_HepCl_ReRun_124 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
125 | 125 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.78 | 3543.577 | 120707 |  | 3543.577 | 4 | 4.38 | 9.5 | bb | 2023-11-07
126 | 126 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.78 | 3555.943 | 121173 |  | 3555.943 | 4 | 4.4 | 9.9 | bb | 2023-11-07
127 | 127 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.78 | 3988.081 | 136185 |  | 3988.081 | 4 | 4.93 | 23.3 | bb | 2023-11-07
128 | 128 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.78 | 3506.183 | 119804 |  | 3506.183 | 4 | 4.34 | 8.4 | bb | 2023-11-07
129 | 129 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.77 | 2791.078 | 94932 |  | 2791.078 | 4 | 3.45 | -13.7 | bb | 2023-11-07
130 | 130 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.78 | 3117.41 | 106241 |  | 3117.41 | 4 | 3.85 | -3.6 | bb | 2023-11-07
131 | 131 | 20231106_HepCl_ReRun_131 | Blank | Blank | 3.8 | 1.14 | 41 |  | 1.14 | 4 | 0 | -100 | bb | 2023-11-07
132 | 132 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 3.78 | 3604.25 | 123885 |  | 3604.25 | 4 | 4.46 | 11.4 | bb | 2023-11-07
133 | 133 | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.81 | 0.164 | 6 |  | 0.164 | 4 | 0 | -100 | bb | 2023-11-07
134 | 134 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.78 | 4056.419 | 140110 |  | 4056.419 | 4 | 5.02 | 25.4 | bb | 2023-11-07
135 | 135 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.78 | 4559.46 | 157257 |  | 4559.46 | 4 | 5.64 | 40.9 | bb | 2023-11-07
136 | 136 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.78 | 4493.075 | 154908 |  | 4493.075 | 4 | 5.56 | 38.9 | bb | 2023-11-07
137 | 137 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.78 | 4114.258 | 141391 |  | 4114.258 | 4 | 5.09 | 27.2 | bb | 2023-11-07
138 | 138 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.78 | 4004.459 | 137439 |  | 4004.459 | 4 | 4.95 | 23.8 | bb | 2023-11-07
139 | 139 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.77 | 3985.281 | 135627 |  | 3985.281 | 4 | 4.93 | 23.2 | bb | 2023-11-07
140 | 140 | 20231106_HepCl_ReRun_140 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
141 | 141 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 3.77 | 3653.36 | 124641 |  | 3653.36 | 4 | 4.52 | 12.9 | bb | 2023-11-07
142 | 142 | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.72 | 2.523 | 91 |  | 2.523 | 4 | 0 | -99.9 | bb | 2023-11-07
143 | 143 | 20231106_HepCl_ReRun_143 | Blank | Blank | 3.81 | 5.162 | 216 |  | 5.162 | 4 | 0.01 | -99.8 | bb | 2023-11-07
144 | 144 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.78 | 5093.202 | 174541 |  | 5093.202 | 4 | 6.3 | 57.4 | bb | 2023-11-07
145 | 145 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.78 | 4939.262 | 170266 |  | 4939.262 | 4 | 6.11 | 52.7 | bb | 2023-11-07
146 | 146 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.78 | 5233.792 | 181293 |  | 5233.792 | 4 | 6.47 | 61.8 | bb | 2023-11-07
147 | 147 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.78 | 4518.646 | 155494 |  | 4518.646 | 4 | 5.59 | 39.7 | bb | 2023-11-07
148 | 148 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.78 | 4595.066 | 158691 |  | 4595.066 | 4 | 5.68 | 42 | bb | 2023-11-07
149 | 149 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.78 | 4516.958 | 153628 |  | 4516.958 | 4 | 5.58 | 39.6 | bb | 2023-11-07
150 | 150 | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.79 | 0.579 | 34 |  | 0.579 | 4 | 0 | -100 | bb | 2023-11-07
151 | 151 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.78 | 4562.208 | 155516 |  | 4562.208 | 4 | 5.64 | 41 | bb | 2023-11-07
152 | 152 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.78 | 4277.345 | 144960 |  | 4277.345 | 4 | 5.29 | 32.2 | bb | 2023-11-07
153 | 153 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.78 | 4028.417 | 137600 |  | 4028.417 | 4 | 4.98 | 24.5 | bb | 2023-11-07
154 | 154 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.78 | 4201.701 | 142567 |  | 4201.701 | 4 | 5.2 | 29.9 | bb | 2023-11-07
155 | 155 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.78 | 4522.662 | 156282 |  | 4522.662 | 4 | 5.59 | 39.8 | bb | 2023-11-07
156 | 156 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.78 | 4864.867 | 169224 |  | 4864.867 | 4 | 6.02 | 50.4 | bb | 2023-11-07
157 | 157 | 20231106_HepCl_ReRun_159 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
158 | 158 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.78 | 4130.82 | 140590 |  | 4130.82 | 4 | 5.11 | 27.7 | bb | 2023-11-07
159 | 159 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.78 | 4509.017 | 155305 |  | 4509.017 | 4 | 5.58 | 39.4 | bb | 2023-11-07
160 | 160 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.78 | 4377.103 | 150190 |  | 4377.103 | 4 | 5.41 | 35.3 | bb | 2023-11-07
161 | 161 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.78 | 3747.828 | 127274 |  | 3747.828 | 4 | 4.63 | 15.8 | bb | 2023-11-07
162 | 162 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.78 | 4038.739 | 139785 |  | 4038.739 | 4 | 4.99 | 24.8 | bb | 2023-11-07
163 | 163 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.78 | 3789.513 | 131680 |  | 3789.513 | 4 | 4.69 | 17.1 | bb | 2023-11-07
164 | 164 | 20231106_HepCl_ReRun_166 | Blank | Blank | 3.82 | 0.413 | 15 |  | 0.413 | 4 | 0 | -100 | bb | 2023-11-07
165 | 165 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 3.78 | 3208.256 | 110436 |  | 3208.256 | 4 | 3.97 | -0.8 | bb | 2023-11-07
166 | 166 | 20231106_HepCl_ReRun_168 | Blank | Blank | 3.81 | 2.772 | 100 |  | 2.772 | 4 | 0 | -99.9 | bb | 2023-11-07
167 | 167 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.78 | 4403.589 | 152070 |  | 4403.589 | 4 | 5.44 | 36.1 | bb | 2023-11-07
168 | 168 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.78 | 4598.22 | 158526 |  | 4598.22 | 4 | 5.69 | 42.1 | bb | 2023-11-07
169 | 169 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.78 | 4689.85 | 160258 |  | 4689.85 | 4 | 5.8 | 45 | bb | 2023-11-07
170 | 170 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.78 | 4132.267 | 141081 |  | 4132.267 | 4 | 5.11 | 27.7 | bb | 2023-11-07
171 | 171 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.78 | 4125.632 | 141685 |  | 4125.632 | 4 | 5.1 | 27.5 | bb | 2023-11-07
172 | 172 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.78 | 4422.781 | 153183 |  | 4422.781 | 4 | 5.47 | 36.7 | bb | 2023-11-07
173 | 173 | 20231106_HepCl_ReRun_175 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
174 | 174 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.78 | 4699.705 | 162117 |  | 4699.705 | 4 | 5.81 | 45.3 | bb | 2023-11-07
175 | 175 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.78 | 4633.228 | 160681 |  | 4633.228 | 4 | 5.73 | 43.2 | bb | 2023-11-07
176 | 176 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.78 | 4931.43 | 170328 |  | 4931.43 | 4 | 6.1 | 52.4 | bb | 2023-11-07
177 | 177 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.78 | 4408.869 | 152599 |  | 4408.869 | 4 | 5.45 | 36.3 | bb | 2023-11-07
178 | 178 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.78 | 4910.336 | 169194 |  | 4910.336 | 4 | 6.07 | 51.8 | bb | 2023-11-07
179 | 179 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.78 | 4659.491 | 160745 |  | 4659.491 | 4 | 5.76 | 44 | bb | 2023-11-07
180 | 180 | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.81 | 1.668 | 38 |  | 1.668 | 4 | 0 | -99.9 | bb | 2023-11-07
181 | 181 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.78 | 3744.223 | 127381 |  | 3744.223 | 4 | 4.63 | 15.7 | bb | 2023-11-07
182 | 182 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.78 | 4034.394 | 138969 |  | 4034.394 | 4 | 4.99 | 24.7 | bb | 2023-11-07
183 | 183 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.78 | 4172.714 | 143367 |  | 4172.714 | 4 | 5.16 | 29 | bb | 2023-11-07
184 | 184 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.78 | 3922.801 | 137534 |  | 3922.801 | 4 | 4.85 | 21.3 | bb | 2023-11-07
185 | 185 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.78 | 3270.064 | 113843 |  | 3270.064 | 4 | 4.04 | 1.1 | bb | 2023-11-07
186 | 186 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.78 | 4049.156 | 140243 |  | 4049.156 | 4 | 5.01 | 25.2 | bb | 2023-11-07
187 | 187 | 20231106_HepCl_ReRun_189 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
188 | 188 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 3.78 | 3623.284 | 125981 |  | 3623.284 | 4 | 4.48 | 12 | bb | 2023-11-07
189 | 189 | 20231106_HepCl_ReRun_191 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
190 | 190 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.78 | 4646.139 | 159487 |  | 4646.139 | 4 | 5.74 | 43.6 | bb | 2023-11-07
191 | 191 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.78 | 4670.868 | 159954 |  | 4670.868 | 4 | 5.78 | 44.4 | bb | 2023-11-07
192 | 192 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.78 | 4760.106 | 162646 |  | 4760.106 | 4 | 5.89 | 47.1 | bb | 2023-11-07
193 | 193 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.78 | 4087.066 | 139402 |  | 4087.066 | 4 | 5.05 | 26.3 | bb | 2023-11-07
194 | 194 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.78 | 4565.343 | 156691 |  | 4565.343 | 4 | 5.64 | 41.1 | bb | 2023-11-07
195 | 195 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.78 | 4533.528 | 154364 |  | 4533.528 | 4 | 5.61 | 40.1 | bb | 2023-11-07
196 | 196 | 20231106_HepCl_ReRun_198 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
197 | 197 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 3.78 | 2966.934 | 101754 |  | 2966.934 | 4 | 3.67 | -8.3 | bb | 2023-11-08
198 | 198 | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.81 | 0.525 | 61 |  | 0.525 | 4 | 0 | -100 | bb | 2023-11-08
199 | 199 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.78 | 4314.995 | 149357 |  | 4314.995 | 4 | 5.34 | 33.4 | bb | 2023-11-08
200 | 200 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.78 | 4169.22 | 142828 |  | 4169.22 | 4 | 5.15 | 28.9 | bb | 2023-11-08
201 | 201 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.78 | 4343.747 | 149701 |  | 4343.747 | 4 | 5.37 | 34.3 | bb | 2023-11-08
202 | 202 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.77 | 3994.895 | 137003 |  | 3994.895 | 4 | 4.94 | 23.5 | bb | 2023-11-08
203 | 203 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.78 | 3869.739 | 132127 |  | 3869.739 | 4 | 4.78 | 19.6 | bb | 2023-11-08
204 | 204 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.78 | 4094.355 | 141335 |  | 4094.355 | 4 | 5.06 | 26.6 | bb | 2023-11-08
205 | 205 | 20231106_HepCl_ReRun_207 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
206 | 206 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.78 | 4088.134 | 141052 |  | 4088.134 | 4 | 5.05 | 26.4 | bb | 2023-11-08
207 | 207 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.78 | 3656.892 | 126601 |  | 3656.892 | 4 | 4.52 | 13 | bb | 2023-11-08
208 | 208 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.78 | 4115.619 | 141480 |  | 4115.619 | 4 | 5.09 | 27.2 | bb | 2023-11-08
209 | 209 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.78 | 3668.877 | 125748 |  | 3668.877 | 4 | 4.54 | 13.4 | bb | 2023-11-08
210 | 210 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.78 | 3204.234 | 110651 |  | 3204.234 | 4 | 3.96 | -1 | bb | 2023-11-08
211 | 211 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.78 | 3915.398 | 135372 |  | 3915.398 | 4 | 4.84 | 21 | bb | 2023-11-08
212 | 212 | 20231106_HepCl_ReRun_214 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
213 | 213 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.78 | 3898.605 | 133591 |  | 3898.605 | 4 | 4.82 | 20.5 | bb | 2023-11-08
214 | 214 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.78 | 4061.239 | 140995 |  | 4061.239 | 4 | 5.02 | 25.5 | bb | 2023-11-08
215 | 215 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.78 | 3517.021 | 121705 |  | 3517.021 | 4 | 4.35 | 8.7 | bb | 2023-11-08
216 | 216 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.78 | 3202.326 | 110024 |  | 3202.326 | 4 | 3.96 | -1 | bb | 2023-11-08
217 | 217 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.78 | 3661.565 | 126574 |  | 3661.565 | 4 | 4.53 | 13.2 | bb | 2023-11-08
218 | 218 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.78 | 2804.06 | 95771 |  | 2804.06 | 4 | 3.47 | -13.3 | bb | 2023-11-08
219 | 219 | 20231106_HepCl_ReRun_221 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
220 | 220 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 3.78 | 3886.825 | 134191 |  | 3886.825 | 4 | 4.81 | 20.1 | bb | 2023-11-08
221 | 221 | 20231106_HepCl_ReRun_223 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
222 | 222 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.78 | 4173.079 | 145040 |  | 4173.079 | 4 | 5.16 | 29 | bb | 2023-11-08
223 | 223 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.78 | 3930.435 | 136738 |  | 3930.435 | 4 | 4.86 | 21.5 | bb | 2023-11-08
224 | 224 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.78 | 3689.984 | 126097 |  | 3689.984 | 4 | 4.56 | 14.1 | bb | 2023-11-08
225 | 225 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.78 | 3838.459 | 133137 |  | 3838.459 | 4 | 4.75 | 18.6 | bb | 2023-11-08
226 | 226 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.78 | 3972.163 | 136332 |  | 3972.163 | 4 | 4.91 | 22.8 | bb | 2023-11-08
227 | 227 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.78 | 3842.832 | 133074 |  | 3842.832 | 4 | 4.75 | 18.8 | bb | 2023-11-08
228 | 228 | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.82 | 4.215 | 142 |  | 4.215 | 4 | 0.01 | -99.9 | bb | 2023-11-08
229 | 229 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.78 | 3893.189 | 135899 |  | 3893.189 | 4 | 4.81 | 20.3 | bb | 2023-11-08
230 | 230 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.78 | 3559.379 | 121498 |  | 3559.379 | 4 | 4.4 | 10 | bb | 2023-11-08
231 | 231 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.78 | 4010.53 | 139009 |  | 4010.53 | 4 | 4.96 | 24 | bb | 2023-11-08
232 | 232 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.78 | 4683.644 | 161988 |  | 4683.644 | 4 | 5.79 | 44.8 | bb | 2023-11-08
233 | 233 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.78 | 3473.603 | 119313 |  | 3473.603 | 4 | 4.29 | 7.4 | bb | 2023-11-08
234 | 234 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.78 | 3760.475 | 129538 |  | 3760.475 | 4 | 4.65 | 16.2 | bb | 2023-11-08
235 | 235 | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.74 | 1.939 | 70 |  | 1.939 | 4 | 0 | -99.9 | bb | 2023-11-08
236 | 236 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.77 | 3804.575 | 130640 |  | 3804.575 | 4 | 4.7 | 17.6 | bb | 2023-11-08
237 | 237 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.78 | 3650.367 | 127037 |  | 3650.367 | 4 | 4.51 | 12.8 | bb | 2023-11-08
238 | 238 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.78 | 4028.423 | 139652 |  | 4028.423 | 4 | 4.98 | 24.5 | bb | 2023-11-08
239 | 239 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.78 | 3605.647 | 124797 |  | 3605.647 | 4 | 4.46 | 11.5 | bb | 2023-11-08
240 | 240 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.78 | 3280.838 | 113487 |  | 3280.838 | 4 | 4.06 | 1.4 | bb | 2023-11-08
241 | 241 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.78 | 3179.154 | 110380 |  | 3179.154 | 4 | 3.93 | -1.7 | bb | 2023-11-08
242 | 242 | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.79 | 0.26 | 9 |  | 0.26 | 4 | 0 | -100 | bb | 2023-11-08
243 | 243 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 3.78 | 3753.745 | 129845 |  | 3753.745 | 4 | 4.64 | 16 | bb | 2023-11-08
244 | 244 | 20231106_HepCl_ReRun_246 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
245 | 245 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.78 | 4269.086 | 147221 |  | 4269.086 | 4 | 5.28 | 32 | bb | 2023-11-08
246 | 246 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.78 | 4165.854 | 143861 |  | 4165.854 | 4 | 5.15 | 28.8 | bb | 2023-11-08
247 | 247 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.78 | 4105.743 | 139697 |  | 4105.743 | 4 | 5.08 | 26.9 | bb | 2023-11-08
248 | 248 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.78 | 4091.521 | 141866 |  | 4091.521 | 4 | 5.06 | 26.5 | bb | 2023-11-08
249 | 249 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.78 | 3699.73 | 127092 |  | 3699.73 | 4 | 4.57 | 14.4 | bb | 2023-11-08
250 | 250 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.78 | 3781.379 | 130088 |  | 3781.379 | 4 | 4.68 | 16.9 | bb | 2023-11-08
251 | 251 | 20231106_HepCl_ReRun_253 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
252 | 252 | 20231106_HepCl_ReRun_254 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
253 | 253 | 20231106_HepCl_ReRun_255 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
Compound 4: 9H-PFNA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20231106_HepCl_ReRun_001 | Blank | Blank | 3.86 | 1311.774 | 28378 | [5 empty] | bb | 2023-11-06
2 | 2 | 20231106_HepCl_ReRun_002 | Blank | Blank | 3.86 | 881.125 | 16293 | [5 empty] | bb | 2023-11-06
3 | 3 | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.85 | 721.975 | 14600 | 5.359 | 538.888 |  | 1528.2 |  | bb | 2023-11-06
4 | 4 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.85 | 607.958 | 11669 | 0.638 | 3811.649 |  | 10811.72 |  | bb | 2023-11-06
5 | 5 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 3.86 | 475.916 | 8865 | [5 empty] | bb | 2023-11-06
6 | 6 | 20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 13.111 | [5 empty] | 2023-11-06
7 | 7 | 20231106_HepCl_ReRun_007 | Blank | Blank | 3.85 | 443.105 | 9101 | [5 empty] | bb | 2023-11-06
8 | 8 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 3.81 | 1564.92 | 41924 | 75194.047 | 0.083 | 0.175 |  |  | bbI | 2023-11-06
9 | 9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 3.81 | 2451.184 | 72295 | 75489.938 | 0.13 | 0.281 |  |  | bbI | 2023-11-06
10 | 10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 3.81 | 3388.928 | 104660 | 66257.367 | 0.205 | 0.444 | 0.17 | -62.2 | bb | 2023-11-06
11 | 11 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 3.8 | 4601.388 | 140744 | 73904.594 | 0.249 | 0.713 | 0.29 | -58.8 | bb | 2023-11-06
12 | 12 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 3.8 | 8336.581 | 273803 | 72824.063 | 0.458 | 1.138 | 0.89 | -22.1 | bb | 2023-11-06
13 | 13 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.8 | 11641.195 | 382511 | 68219.305 | 0.683 | 1.825 | 1.52 | -16.5 | bb | 2023-11-06
14 | 14 | 20231106_HepCl_ReRun_014 | Blank | Blank |  |  |  | 3.806 | [5 empty] | 2023-11-06
15 | 15 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.81 | 21097.494 | 703071 | 64042.156 | 1.318 | 2.919 | 3.33 | 13.9 | bb | 2023-11-06
16 | 16 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.81 | 29158.406 | 967248 | 65068.902 | 1.792 | 4.669 | 4.67 | 0.1 | bb | 2023-11-06
17 | 17 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.81 | 48933.699 | 1647090 | 70109.063 | 2.792 | 7.45 | 7.51 | 0.8 | bb | 2023-11-06
18 | 18 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.81 | 79207.258 | 2656081 | 69572.93 | 4.554 | 11.875 | 12.5 | 5.3 | bb | 2023-11-06
19 | 19 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.81 | 161645.313 | 5450768 | 72849.406 | 8.876 | 19.125 | 24.76 | 29.5 | bb | 2023-11-06
20 | 20 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.8 | 217131.609 | 7364952 | 60371.941 | 14.386 | 30.5 | 40.4 | 32.4 | bb | 2023-11-06
21 | 21 | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.85 | 265.891 | 4687 | 7.869 | 135.159 |  | 382.98 |  | bb | 2023-11-06
22 | 22 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.81 | 279932.938 | 9373593 | 64329.785 | 17.406 | 48.75 | 48.96 | 0.4 | bb | 2023-11-06
23 | 23 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.81 | 510067.344 | 17057968 | 72004.906 | 28.335 | 78.125 | 79.96 | 2.4 | bb | 2023-11-06
24 | 24 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.81 | 653154.375 | 22047396 | 67768.188 | 38.552 | 125 | 108.94 | -12.8 | bb | 2023-11-06
25 | 25 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 3.81 | 1079783.75 | 36050296 | 89323.164 | 48.354 | 200 | 136.75 | -31.6 | bb | 2023-11-06
26 | 26 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 3.81 | 1637311 | 54934608 | 75562.414 | 86.673 | 325.625 | 245.44 | -24.6 | bb | 2023-11-06
27 | 27 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | [9 empty] | 2023-11-06
28 | 28 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.86 | 244.642 | 4634 | 9.577 | 102.179 |  | 289.43 |  | bb | 2023-11-06
29 | 29 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 3.81 | 4751.315 | 151755 | 72754.594 | 0.261 | 0.713 | 0.33 | -53.9 | bb | 2023-11-06
30 | 30 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 3.81 | 6234.799 | 201623 | 67752.633 | 0.368 | 1.138 | 0.63 | -44.5 | bb | 2023-11-06
31 | 31 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.81 | 10848.09 | 359772 | 65001.293 | 0.668 | 1.825 | 1.48 | -18.9 | bb | 2023-11-06
32 | 32 | 20231106_HepCl_ReRun_032 | Blank | Blank | 3.85 | 231.762 | 4646 | [5 empty] | bb | 2023-11-06
33 | 33 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 3.81 | 5260.333 | 171299 | 73487.75 | 0.286 | 0.713 | 0.4 | -43.9 | bb | 2023-11-06
34 | 34 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 3.81 | 6547.865 | 214088 | 66896.563 | 0.392 | 1.138 | 0.7 | -38.6 | bb | 2023-11-06
35 | 35 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.8 | 11571.36 | 384229 | 66638.734 | 0.695 | 1.825 | 1.56 | -14.7 | bb | 2023-11-06
36 | 36 | 20231106_HepCl_ReRun_036 | Blank | Blank | 3.85 | 224.475 | 4100 | [5 empty] | bb | 2023-11-06
37 | 37 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 3.8 | 4919.652 | 157957 | 70830.211 | 0.278 | 0.713 | 0.38 | -47.3 | bb | 2023-11-06
38 | 38 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 3.8 | 7019.847 | 228077 | 64723.23 | 0.434 | 1.138 | 0.82 | -28.1 | bb | 2023-11-06
39 | 39 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.81 | 11606.056 | 384908 | 61013.957 | 0.761 | 1.825 | 1.75 | -4.3 | bb | 2023-11-06
40 | 40 | 20231106_HepCl_ReRun_040 | Blank | Blank | 3.85 | 223.767 | 3692 | [5 empty] | bb | 2023-11-06
41 | 41 | 20231106_HepCl_ReRun_041 | QC A | QC | 3.81 | 5815.193 | 189001 | 87338.508 | 0.266 | 0.625 | 0.34 | -45.2 | bb | 2023-11-06
42 | 42 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.81 | 22068.494 | 733400 | 76308.461 | 1.157 | 2.506 | 2.87 | 14.5 | bb | 2023-11-06
43 | 43 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.81 | 51954.094 | 1766232 | 84343.641 | 2.464 | 6.269 | 6.58 | 4.9 | bb | 2023-11-06
44 | 44 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.81 | 182690.922 | 6144534 | 70742.273 | 10.33 | 25 | 28.89 | 15.6 | bb | 2023-11-06
45 | 45 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 3.86 | 183.34 | 3807 | 8.655 | 84.733 |  | 239.94 |  | bb | 2023-11-06
46 | 46 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | [9 empty] | 2023-11-06
47 | 47 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 3.81 | 1474.341 | 41632 | 66380.367 | 0.089 | 0.175 |  |  | bbI | 2023-11-06
48 | 48 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 3.81 | 2342.242 | 73223 | 68137.133 | 0.138 | 0.281 |  |  | bbI | 2023-11-06
49 | 49 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 3.81 | 3022.632 | 95828 | 60965.863 | 0.198 | 0.444 | 0.15 | -66.2 | bb | 2023-11-06
50 | 50 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 3.81 | 5242.089 | 172577 | 67549.781 | 0.31 | 0.713 | 0.47 | -34.3 | bb | 2023-11-06
51 | 51 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 3.81 | 6910.878 | 225487 | 68209.023 | 0.405 | 1.138 | 0.74 | -35.2 | bb | 2023-11-06
52 | 52 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.81 | 11166.135 | 371375 | 64340.457 | 0.694 | 1.825 | 1.56 | -14.7 | bb | 2023-11-06
53 | 53 | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.85 | 234.037 | 4499 | 12.536 | 74.677 |  | 211.42 |  | bb | 2023-11-06
54 | 54 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.8 | 20381.063 | 685582 | 67724.766 | 1.204 | 2.919 | 3 | 2.8 | bb | 2023-11-06
55 | 55 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.81 | 26398.207 | 865225 | 64294.883 | 1.642 | 4.669 | 4.25 | -9.1 | bb | 2023-11-06
56 | 56 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.81 | 46315.258 | 1565880 | 63195.191 | 2.932 | 7.45 | 7.9 | 6.1 | bb | 2023-11-06
57 | 57 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.81 | 75334.328 | 2526905 | 63700.488 | 4.731 | 11.875 | 13.01 | 9.5 | bb | 2023-11-06
58 | 58 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.81 | 152744.953 | 5169925 | 68771.055 | 8.884 | 19.125 | 24.79 | 29.6 | bb | 2023-11-06
59 | 59 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.81 | 215742.5 | 7243975 | 69718.367 | 12.378 | 30.5 | 34.7 | 13.8 | bb | 2023-11-06
60 | 60 | 20231106_HepCl_ReRun_060 | Blank | Blank | 3.85 | 243.211 | 4301 | [5 empty] | bb | 2023-11-06
61 | 61 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.8 | 275087.375 | 9279365 | 69190.055 | 15.903 | 48.75 | 44.7 | -8.3 | bb | 2023-11-06
62 | 62 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.81 | 494636.938 | 16853748 | 75166.273 | 26.322 | 78.125 | 74.25 | -5 | bb | 2023-11-06
63 | 63 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.81 | 687010.063 | 23100348 | 64113.281 | 42.862 | 125 | 121.17 | -3.1 | bb | 2023-11-06
64 | 64 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 3.81 | 1203123.5 | 40178528 | 89929.188 | 53.514 | 200 | 151.39 | -24.3 | bb | 2023-11-06
65 | 65 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 3.81 | 1690942.625 | 56753616 | 73827.055 | 91.616 | 325.625 | 259.47 | -20.3 | bb | 2023-11-07
66 | 66 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 3.86 | 175.411 | 3826 | 0.454 | 1545.471 |  | 4383.47 |  | bb | 2023-11-07
67 | 67 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 3.86 | 250.409 | 4456 | 12.974 | 77.203 |  | 218.58 |  | bb | 2023-11-07
68 | 68 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 3.81 | 5032.04 | 161812 | 70878.07 | 0.284 | 0.713 | 0.39 | -44.9 | bb | 2023-11-07
69 | 69 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 3.81 | 6854.112 | 229917 | 65752.898 | 0.417 | 1.138 | 0.77 | -32.3 | bb | 2023-11-07
70 | 70 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.81 | 11005.288 | 365567 | 63387.512 | 0.694 | 1.825 | 1.56 | -14.7 | bb | 2023-11-07
71 | 71 | 20231106_HepCl_ReRun_071 | Blank | Blank | 3.86 | 178.541 | 3006 | 7.057 | 101.199 |  | 286.65 |  | bb | 2023-11-07
72 | 72 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 3.81 | 4577.015 | 146341 | 68083.242 | 0.269 | 0.713 | 0.35 | -50.9 | bb | 2023-11-07
73 | 73 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 3.81 | 5772.026 | 189798 | 64498.293 | 0.358 | 1.138 | 0.6 | -47 | bb | 2023-11-07
74 | 74 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.81 | 10255.913 | 336728 | 62066.836 | 0.661 | 1.825 | 1.46 | -19.9 | bb | 2023-11-07
75 | 75 | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.86 | 219.784 | 4284 | 12.113 | 72.578 |  | 205.46 |  | bb | 2023-11-07
76 | 76 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 3.81 | 5230.587 | 169974 | 66670.469 | 0.314 | 0.713 | 0.48 | -33 | bb | 2023-11-07
77 | 77 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 3.81 | 5196.095 | 167226 | 63275.16 | 0.328 | 1.138 | 0.52 | -54.4 | bb | 2023-11-07
78 | 78 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.8 | 11621.744 | 390664 | 61289.246 | 0.758 | 1.825 | 1.74 | -4.7 | bb | 2023-11-07
79 | 79 | 20231106_HepCl_ReRun_079 | Blank | Blank | [9 empty] | 2023-11-07
80 | 80 | 20231106_HepCl_ReRun_080 | QC A | QC | 3.81 | 4988.641 | 158149 | 88138.531 | 0.226 | 0.625 | 0.23 | -63.3 | bb | 2023-11-07
81 | 81 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.8 | 19008.291 | 624404 | 84922.633 | 0.895 | 2.506 | 2.13 | -15.1 | bb | 2023-11-07
82 | 82 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.81 | 56246.949 | 1905280 | 79141.828 | 2.843 | 6.269 | 7.65 | 22.1 | bb | 2023-11-07
83 | 83 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.8 | 186631.609 | 6350371 | 68663.195 | 10.872 | 25 | 30.43 | 21.7 | bb | 2023-11-07
84 | 84 | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.86 | 108.338 | 2639 | 2.32 | 186.79 |  | 529.44 |  | bb | 2023-11-07
85 | 85 | 20231106_HepCl_ReRun_085 | Blank | Blank | 3.86 | 243.378 | 5428 | [5 empty] | bb | 2023-11-07
86 | 86 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 3.81 | 163088.453 | 5391557 | 70980.813 | 9.191 | 30.5 | 25.66 | -15.9 | bb | 2023-11-07
87 | 87 | 20231106_HepCl_ReRun_087 | Blank | Blank |  |  |  | 3.535 | [5 empty] | 2023-11-07
88 | 88 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.81 | 514.675 | 8646 | 127111.25 | 0.016 |  |  |  | bbI | 2023-11-07
89 | 89 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.81 | 993.376 | 26476 | 107706.938 | 0.037 |  |  |  | bbI | 2023-11-07
90 | 90 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.81 | 642.249 | 12421 | 113437.414 | 0.023 |  |  |  | bbI | 2023-11-07
91 | 91 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.81 | 408.254 | 5604 | 110987.766 | 0.015 |  |  |  | bbI | 2023-11-07
92 | 92 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.81 | 337.916 | 5499 | 99300.031 | 0.014 |  |  |  | bbI | 2023-11-07
93 | 93 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.81 | 462.735 | 7215 | 105545.609 | 0.018 |  |  |  | bbI | 2023-11-07
94 | 94 | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.86 | 135.857 | 3418 | 4.894 | 111.04 |  | 314.56 |  | bb | 2023-11-07
95 | 95 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.81 | 453.139 | 6478 | 104039.031 | 0.017 |  |  |  | bbI | 2023-11-07
96 | 96 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.86 | 381.047 | 4607 | 119178.984 | 0.013 |  |  |  | bbI | 2023-11-07
97 | 97 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.81 | 396.53 | 5932 | 113962.063 | 0.014 |  |  |  | bbI | 2023-11-07
98 | 98 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.81 | 322.722 | 5508 | 118969.555 | 0.011 |  |  |  | bbI | 2023-11-07
99 | 99 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.85 | 377.622 | 4486 | 122033.266 | 0.012 |  |  |  | bbI | 2023-11-07
100 | 100 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.85 | 390.746 | 5225 | 110901.977 | 0.014 |  |  |  | bbI | 2023-11-07
101 | 101 | 20231106_HepCl_ReRun_101 | Blank | Blank | 3.85 | 166.996 | 3172 | [5 empty] | bb | 2023-11-07
102 | 102 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.81 | 412.521 | 6320 | 94312.25 | 0.017 |  |  |  | bbI | 2023-11-07
103 | 103 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.8 | 408.999 | 5523 | 101210.508 | 0.016 |  |  |  | bbI | 2023-11-07
104 | 104 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.81 | 356.726 | 5025 | 104400.188 | 0.014 |  |  |  | bbI | 2023-11-07
105 | 105 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.81 | 404.812 | 5864 | 84769.352 | 0.019 |  |  |  | bbI | 2023-11-07
106 | 106 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.85 | 333.628 | 4606 | 88387.469 | 0.015 |  |  |  | bbI | 2023-11-07
107 | 107 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.86 | 371.009 | 4792 | 75575.227 | 0.02 |  |  |  | bbI | 2023-11-07
108 | 108 | 20231106_HepCl_ReRun_108 | Blank | Blank |  |  |  | 3.622 | [5 empty] | 2023-11-07
109 | 109 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 3.81 | 203688.5 | 6846477 | 68778 | 11.846 | 30.5 | 33.19 | 8.8 | bb | 2023-11-07
110 | 110 | 20231106_HepCl_ReRun_110 | Blank | Blank | 3.85 | 202.452 | 3781 | 1.239 | 653.598 |  | 1853.58 |  | bb | 2023-11-07
111 | 111 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.81 | 1741.671 | 49424 | 98270.578 | 0.071 |  |  |  | bbI | 2023-11-07
112 | 112 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.85 | 359.241 | 4577 | 98628.742 | 0.015 |  |  |  | bbI | 2023-11-07
113 | 113 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.86 | 327.177 | 5240 | 106868.266 | 0.012 |  |  |  | bbI | 2023-11-07
114 | 114 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.81 | 471.128 | 8876 | 104541.266 | 0.018 |  |  |  | bbI | 2023-11-07
115 | 115 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.81 | 416.801 | 6653 | 112556.063 | 0.015 |  |  |  | bbI | 2023-11-07
116 | 116 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.81 | 276.12 | 4768 | 104859.008 | 0.011 |  |  |  | bbI | 2023-11-07
117 | 117 | 20231106_HepCl_ReRun_117 | Blank | Blank | 3.86 | 219.288 | 4057 | [5 empty] | bb | 2023-11-07
118 | 118 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.81 | 414.222 | 6606 | 105383.109 | 0.016 |  |  |  | bbI | 2023-11-07
119 | 119 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.85 | 410.053 | 5828 | 99879.406 | 0.016 |  |  |  | bbI | 2023-11-07
120 | 120 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.81 | 350.653 | 7143 | 94358.172 | 0.015 |  |  |  | bbI | 2023-11-07
121 | 121 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.8 | 772.868 | 20550 | 116692.453 | 0.026 |  |  |  | bbI | 2023-11-07
122 | 122 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.81 | 947.536 | 23810 | 107368.117 | 0.035 |  |  |  | bbI | 2023-11-07
123 | 123 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.81 | 1016.708 | 27596 | 109868.875 | 0.037 |  |  |  | bbI | 2023-11-07
124 | 124 | 20231106_HepCl_ReRun_124 | Blank | Blank | 3.85 | 189.908 | 3607 | [5 empty] | bb | 2023-11-07
125 | 125 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.81 | 381.625 | 5824 | 102934.406 | 0.015 |  |  |  | bbI | 2023-11-07
126 | 126 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.85 | 415.396 | 5471 | 95880.813 | 0.017 |  |  |  | bbI | 2023-11-07
127 | 127 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.81 | 350.126 | 4576 | 111986.016 | 0.013 |  |  |  | bbI | 2023-11-07
128 | 128 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte |  |  |  | 84980.797 | [5 empty] | 2023-11-07
129 | 129 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.85 | 292.296 | 3730 | 64520.02 | 0.018 |  |  |  | bbI | 2023-11-07
130 | 130 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.85 | 367.758 | 5330 | 77287.742 | 0.019 |  |  |  | bbI | 2023-11-07
131 | 131 | 20231106_HepCl_ReRun_131 | Blank | Blank | 3.85 | 214.173 | 4079 | 14.521 | 58.997 |  | 166.94 |  | bb | 2023-11-07
132 | 132 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 3.81 | 202991.078 | 6881815 | 65707.844 | 12.357 | 30.5 | 34.64 | 13.6 | bb | 2023-11-07
133 | 133 | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.85 | 208.384 | 3838 | 7.49 | 111.287 |  | 315.26 |  | bb | 2023-11-07
134 | 134 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.81 | 457.152 | 6935 | 96807.188 | 0.019 |  |  |  | bbI | 2023-11-07
135 | 135 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.81 | 475.469 | 8669 | 110827.492 | 0.017 |  |  |  | bbI | 2023-11-07
136 | 136 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.81 | 554.069 | 8920 | 106620.406 | 0.021 |  |  |  | bbI | 2023-11-07
137 | 137 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.81 | 469.105 | 6650 | 86723.508 | 0.022 |  |  |  | bbI | 2023-11-07
138 | 138 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte |  |  |  | 97677.023 | [5 empty] | 2023-11-07
139 | 139 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.86 | 673.771 | 8906 | 100329.82 | 0.027 |  |  |  | bbI | 2023-11-07
140 | 140 | 20231106_HepCl_ReRun_140 | Blank | Blank | 3.85 | 363.167 | 8123 | [5 empty] | bb | 2023-11-07
141 | 141 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 3.8 | 205355.906 | 6923133 | 64887.246 | 12.659 | 30.5 | 35.5 | 16.4 | bb | 2023-11-07
142 | 142 | 20231106_HepCl_ReRun_142 | Blank | Blank | 3.86 | 295.344 | 5078 | [5 empty] | bb | 2023-11-07
143 | 143 | 20231106_HepCl_ReRun_143 | Blank | Blank | 3.85 | 306.383 | 5571 | 2.895 | 423.327 |  | 1200.4 |  | bb | 2023-11-07
144 | 144 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.81 | 1968.833 | 59138 | 120155.836 | 0.066 |  |  |  | bbI | 2023-11-07
145 | 145 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.81 | 1659.804 | 49492 | 112531.641 | 0.059 |  |  |  | bbI | 2023-11-07
146 | 146 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.81 | 1071.008 | 26239 | 118805.617 | 0.036 |  |  |  | bbI | 2023-11-07
147 | 147 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.8 | 1369.747 | 38750 | 102610.383 | 0.053 |  |  |  | bbI | 2023-11-07
148 | 148 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.85 | 313.797 | 3798 | 102214.219 | 0.012 |  |  |  | bbI | 2023-11-07
149 | 149 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.81 | 374.854 | 4879 | 99535.336 | 0.015 |  |  |  | bbI | 2023-11-07
150 | 150 | 20231106_HepCl_ReRun_152 | Blank | Blank | 3.86 | 275.404 | 5049 | [5 empty] | bb | 2023-11-07
151 | 151 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.8 | 544.265 | 9791 | 103807.086 | 0.021 |  |  |  | bbI | 2023-11-07
152 | 152 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.8 | 482.408 | 8308 | 99959.609 | 0.019 |  |  |  | bbI | 2023-11-07
153 | 153 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.81 | 1167.067 | 33017 | 100739.109 | 0.046 |  |  |  | bbI | 2023-11-07
154 | 154 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.81 | 260.508 | 9475 | 95184.141 | 0.011 |  |  |  | bbI | 2023-11-07
155 | 155 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.81 | 3652.715 | 113771 | 106147.617 | 0.138 |  |  |  | bbI | 2023-11-07
156 | 156 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.81 | 1195.365 | 32755 | 98089.266 | 0.049 |  |  |  | bbI | 2023-11-07
157 | 157 | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.85 | 188.909 | 3291 | 4.77 | 158.414 |  | 448.95 |  | bb | 2023-11-07
158 | 158 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.81 | 339.214 | 5143 | 85510.195 | 0.016 |  |  |  | bbI | 2023-11-07
159 | 159 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.81 | 492.186 | 8924 | 92018.813 | 0.021 |  |  |  | bbI | 2023-11-07
160 | 160 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.81 | 1344.97 | 39408 | 87029.469 | 0.062 |  |  |  | bbI | 2023-11-07
161 | 161 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.81 | 20960.096 | 702734 | 89270.031 | 0.939 |  | 2.25 |  | bb | 2023-11-07
162 | 162 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.86 | 364.35 | 4413 | 85731.852 | 0.017 |  |  |  | bbI | 2023-11-07
163 | 163 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.8 | 304.308 | 3837 | 82351.266 | 0.015 |  |  |  | bbI | 2023-11-07
164 | 164 | 20231106_HepCl_ReRun_166 | Blank | Blank |  |  |  | 6.376 | [5 empty] | 2023-11-07
165 | 165 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 3.8 | 164249.453 | 5518383 | 62184.316 | 10.565 | 30.5 | 29.56 | -3.1 | bb | 2023-11-07
166 | 166 | 20231106_HepCl_ReRun_168 | Blank | Blank | [9 empty] | 2023-11-07
167 | 167 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.8 | 331.232 | 5083 | 96676.672 | 0.014 |  |  |  | bbI | 2023-11-07
168 | 168 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.8 | 363.727 | 6758 | 106880.938 | 0.014 |  |  |  | bbI | 2023-11-07
169 | 169 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.8 | 1032.07 | 27234 | 109563.195 | 0.038 |  |  |  | bbI | 2023-11-07
170 | 170 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.81 | 433.407 | 8770 | 99443.953 | 0.017 |  |  |  | bbI | 2023-11-07
171 | 171 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.86 | 279.778 | 3623 | 106413.328 | 0.011 |  |  |  | bbI | 2023-11-07
172 | 172 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte |  |  |  | 109031.305 | [5 empty] | 2023-11-07
173 | 173 | 20231106_HepCl_ReRun_175 | Blank | Blank | 3.85 | 215.745 | 4778 | [5 empty] | bb | 2023-11-07
174 | 174 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.81 | 346.74 | 5135 | 100056 | 0.014 |  |  |  | bbI | 2023-11-07
175 | 175 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.8 | 493.649 | 10680 | 89716.656 | 0.022 |  |  |  | bbI | 2023-11-07
176 | 176 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.85 | 335.928 | 4351 | 109559.086 | 0.012 |  |  |  | bbI | 2023-11-07
177 | 177 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.81 | 556.298 | 12640 | 90430.555 | 0.025 |  |  |  | bbI | 2023-11-07
178 | 178 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.81 | 745.601 | 16891 | 102643.82 | 0.029 |  |  |  | bbI | 2023-11-07
179 | 179 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.81 | 662.704 | 14954 | 96604.727 | 0.027 |  |  |  | bbI | 2023-11-07
180 | 180 | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.86 | 188.246 | 4310 | 14.625 | 51.486 |  | 145.63 |  | bb | 2023-11-07
181 | 181 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.8 | 629.296 | 15328 | 105342.867 | 0.024 |  |  |  | bbI | 2023-11-07
182 | 182 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.85 | 294.01 | 4664 | 98521.195 | 0.012 |  |  |  | bbI | 2023-11-07
183 | 183 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.81 | 312.739 | 4898 | 87807.195 | 0.014 |  |  |  | bbI | 2023-11-07
184 | 184 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.86 | 289.661 | 4097 | 84955.523 | 0.014 |  |  |  | bbI | 2023-11-07
185 | 185 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte |  |  |  | 83049.625 | [5 empty] | 2023-11-07
186 | 186 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.85 | 285.751 | 3659 | 75152.219 | 0.015 |  |  |  | bbI | 2023-11-07
187 | 187 | 20231106_HepCl_ReRun_189 | Blank | Blank | [9 empty] | 2023-11-07
188 | 188 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 3.81 | 187763.328 | 6425514 | 65820.523 | 11.411 | 30.5 | 31.95 | 4.8 | bb | 2023-11-07
189 | 189 | 20231106_HepCl_ReRun_191 | Blank | Blank | 3.85 | 107.023 | 2703 | 6.924 | 61.827 |  | 174.97 |  | bb | 2023-11-07
190 | 190 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.81 | 370.412 | 5965 | 102671.5 | 0.014 |  |  |  | bbI | 2023-11-07
191 | 191 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.81 | 368.678 | 5262 | 101879.82 | 0.014 |  |  |  | bbI | 2023-11-07
192 | 192 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.81 | 929.424 | 25504 | 119917.711 | 0.031 |  |  |  | bbI | 2023-11-07
193 | 193 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.81 | 397.327 | 5254 | 97705.281 | 0.016 |  |  |  | bbI | 2023-11-07
194 | 194 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.81 | 355.733 | 5451 | 102602.477 | 0.014 |  |  |  | bbI | 2023-11-07
195 | 195 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.81 | 345.67 | 4848 | 91857.102 | 0.015 |  |  |  | bbI | 2023-11-07
196 | 196 | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.86 | 207.931 | 4316 | 3.35 | 248.276 |  | 703.85 |  | bb | 2023-11-08
197 | 197 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 3.81 | 180207.031 | 6069535 | 51659.656 | 13.953 | 30.5 | 39.17 | 28.4 | bb | 2023-11-08
198 | 198 | 20231106_HepCl_ReRun_200 | Blank | Blank | 3.85 | 181.452 | 3449 | 2.811 | 258.203 |  | 732.01 |  | bb | 2023-11-08
199 | 199 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.81 | 211102.141 | 6984337 | 122806.492 | 6.876 |  | 19.09 |  | bb | 2023-11-08
200 | 200 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.81 | 308503.531 | 10401536 | 105540.742 | 11.692 |  | 32.75 |  | bb | 2023-11-08
201 | 201 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.8 | 356524.719 | 12119818 | 108226.414 | 13.177 |  | 36.97 |  | bb | 2023-11-08
202 | 202 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.81 | 384793.406 | 13078994 | 98816.094 | 15.576 |  | 43.77 |  | bb | 2023-11-08
203 | 203 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.81 | 283668.125 | 9436496 | 114010.602 | 9.952 |  | 27.82 |  | bb | 2023-11-08
204 | 204 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.81 | 270666.656 | 8983324 | 111625.508 | 9.699 |  | 27.1 |  | bb | 2023-11-08
205 | 205 | 20231106_HepCl_ReRun_207 | Blank | Blank | [9 empty] | 2023-11-08
206 | 206 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.8 | 298650.938 | 9971744 | 111956.539 | 10.67 |  | 29.85 |  | bb | 2023-11-08
207 | 207 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.81 | 333344 | 11203420 | 104588.813 | 12.749 |  | 35.75 |  | bb | 2023-11-08
208 | 208 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.8 | 274019.813 | 9070789 | 87999.492 | 12.456 |  | 34.92 |  | bb | 2023-11-08
209 | 209 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.81 | 299645.75 | 10020487 | 97360.305 | 12.311 |  | 34.51 |  | bb | 2023-11-08
210 | 210 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.81 | 235749.125 | 7878546 | 84831.938 | 11.116 |  | 31.12 |  | bb | 2023-11-08
211 | 211 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.8 | 323462.594 | 11000032 | 92662.234 | 13.963 |  | 39.19 |  | bb | 2023-11-08
212 | 212 | 20231106_HepCl_ReRun_214 | Blank | Blank | 3.85 | 12.237 | 909 | [5 empty] | bb | 2023-11-08
213 | 213 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.81 | 283991.438 | 9699914 | 103418.438 | 10.984 |  | 30.75 |  | bb | 2023-11-08
214 | 214 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.81 | 303358.25 | 10312001 | 106341.625 | 11.411 |  | 31.95 |  | bb | 2023-11-08
215 | 215 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.81 | 279995.094 | 9373390 | 92977.766 | 12.046 |  | 33.76 |  | bb | 2023-11-08
216 | 216 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.81 | 210909.141 | 7091518 | 79338.781 | 10.633 |  | 29.75 |  | bb | 2023-11-08
217 | 217 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.81 | 219247.406 | 7365930 | 92511.57 | 9.48 |  | 26.48 |  | bb | 2023-11-08
218 | 218 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.81 | 177725.781 | 5986646 | 85295.422 | 8.335 |  | 23.23 |  | bb | 2023-11-08
219 | 219 | 20231106_HepCl_ReRun_221 | Blank | Blank | [9 empty] | 2023-11-08
220 | 220 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 3.81 | 190585.047 | 6503453 | 59542.715 | 12.803 | 30.5 | 35.91 | 17.7 | bb | 2023-11-08
221 | 221 | 20231106_HepCl_ReRun_223 | Blank | Blank | 3.85 | 209.716 | 4694 | [5 empty] | bb | 2023-11-08
222 | 222 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.81 | 226021.344 | 7559059 | 110140.234 | 8.208 |  | 22.87 |  | bb | 2023-11-08
223 | 223 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.81 | 278933.406 | 9405930 | 108602.188 | 10.274 |  | 28.73 |  | bb | 2023-11-08
224 | 224 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.81 | 276028.344 | 9283592 | 90797.094 | 12.16 |  | 34.08 |  | bb | 2023-11-08
225 | 225 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.81 | 367082 | 12520329 | 111401.727 | 13.18 |  | 36.98 |  | bb | 2023-11-08
226 | 226 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.81 | 353533.938 | 11858321 | 110531.492 | 12.794 |  | 35.88 |  | bb | 2023-11-08
227 | 227 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.81 | 323413.156 | 10911557 | 112650.75 | 11.484 |  | 32.16 |  | bb | 2023-11-08
228 | 228 | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.85 | 165.117 | 3099 | 9.549 | 69.166 |  | 195.78 |  | bb | 2023-11-08
229 | 229 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.81 | 345005.031 | 11796750 | 113099.43 | 12.202 |  | 34.2 |  | bb | 2023-11-08
230 | 230 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.81 | 336113.375 | 11298503 | 91099.672 | 14.758 |  | 41.45 |  | bb | 2023-11-08
231 | 231 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.81 | 369345.031 | 12524720 | 95262.484 | 15.509 |  | 43.58 |  | bb | 2023-11-08
232 | 232 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.81 | 278402.094 | 9377985 | 95167.969 | 11.702 |  | 32.78 |  | bb | 2023-11-08
233 | 233 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.81 | 253162.359 | 8413451 | 101755.492 | 9.952 |  | 27.82 |  | bb | 2023-11-08
234 | 234 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.81 | 245721.672 | 8129133 | 87558.133 | 11.226 |  | 31.43 |  | bb | 2023-11-08
235 | 235 | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.86 | 148.115 | 2568 | 0.334 | 1773.832 |  | 5031.24 |  | bb | 2023-11-08
236 | 236 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.81 | 216459.453 | 6929789 | 118122.609 | 7.33 |  | 20.38 |  | bb | 2023-11-08
237 | 237 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.81 | 338452.781 | 11647518 | 102091.367 | 13.261 |  | 37.2 |  | bb | 2023-11-08
238 | 238 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.81 | 282383.156 | 9499879 | 125547.516 | 8.997 |  | 25.11 |  | bb | 2023-11-08
239 | 239 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.81 | 194233.328 | 6480172 | 96524.773 | 8.049 |  | 22.42 |  | bb | 2023-11-08
240 | 240 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.81 | 237021.109 | 8063289 | 88051.555 | 10.767 |  | 30.13 |  | bb | 2023-11-08
241 | 241 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.81 | 243039.016 | 8258455 | 96476.289 | 10.077 |  | 28.17 |  | bb | 2023-11-08
242 | 242 | 20231106_HepCl_ReRun_244 | Blank | Blank | 3.85 | 184.859 | 3229 | [5 empty] | bb | 2023-11-08
243 | 243 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 3.81 | 152733.531 | 5120461 | 53564.039 | 11.406 | 30.5 | 31.94 | 4.7 | bb | 2023-11-08
244 | 244 | 20231106_HepCl_ReRun_246 | Blank | Blank | 3.86 | 213.496 | 3757 | 1.2 | 711.653 |  | 2018.26 |  | bb | 2023-11-08
245 | 245 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.81 | 352025.063 | 11860594 | 120083.727 | 11.726 |  | 32.85 |  | bb | 2023-11-08
246 | 246 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.81 | 261868.188 | 8873026 | 120913.375 | 8.663 |  | 24.16 |  | bb | 2023-11-08
247 | 247 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.81 | 358415.75 | 12138892 | 120947.359 | 11.854 |  | 33.21 |  | bb | 2023-11-08
248 | 248 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.81 | 284973.594 | 9606363 | 98304.648 | 11.596 |  | 32.48 |  | bb | 2023-11-08
249 | 249 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.81 | 236682.141 | 7911366 | 108873.719 | 8.696 |  | 24.25 |  | bb | 2023-11-08
250 | 250 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.81 | 301589.063 | 10239744 | 104880.836 | 11.502 |  | 32.21 |  | bb | 2023-11-08
251 | 251 | 20231106_HepCl_ReRun_253 | Blank | Blank |  |  |  | 12.938 | [5 empty] | 2023-11-08
252 | 252 | 20231106_HepCl_ReRun_254 | Blank | Blank | 3.86 | 153.336 | 3207 | [5 empty] | bb | 2023-11-08
253 | 253 | 20231106_HepCl_ReRun_255 | Blank | Blank | 3.85 | 181.382 | 3786 | [5 empty] | bb | 2023-11-08
Compound 5: M9PFNA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20231106_HepCl_ReRun_001 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
2 | 2 | 20231106_HepCl_ReRun_002 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
3 | 3 | 20231106_HepCl_ReRun_003 | Blank | Blank | 3.95 | 5.359 | 177 |  | 5.359 | 4 | 0 | -100 | bb | 2023-11-06
4 | 4 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 3.96 | 0.638 | 27 |  | 0.638 | 4 | 0 | -100 | bb | 2023-11-06
5 | 5 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
6 | 6 | 20231106_HepCl_ReRun_006 | Blank | Blank | 3.96 | 13.111 | 178 |  | 13.111 | 4 | 0 | -100 | bb | 2023-11-06
7 | 7 | 20231106_HepCl_ReRun_007 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
8 | 8 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 3.96 | 75194.047 | 2065091 |  | 75194.047 | 4 | 4.54 | 13.4 | bb | 2023-11-06
9 | 9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 3.97 | 75489.938 | 2134003 |  | 75489.938 | 4 | 4.55 | 13.9 | bb | 2023-11-06
10 | 10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 3.96 | 66257.367 | 1850615 |  | 66257.367 | 4 | 4 | -0.1 | bb | 2023-11-06
11 | 11 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 3.97 | 73904.594 | 2064188 |  | 73904.594 | 4 | 4.46 | 11.5 | bb | 2023-11-06
12 | 12 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 3.96 | 72824.063 | 2060878 |  | 72824.063 | 4 | 4.39 | 9.8 | bb | 2023-11-06
13 | 13 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 3.96 | 68219.305 | 1940316 |  | 68219.305 | 4 | 4.12 | 2.9 | bb | 2023-11-06
14 | 14 | 20231106_HepCl_ReRun_014 | Blank | Blank | 3.96 | 3.806 | 133 |  | 3.806 | 4 | 0 | -100 | bb | 2023-11-06
15 | 15 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 3.96 | 64042.156 | 1714197 |  | 64042.156 | 4 | 3.86 | -3.4 | bb | 2023-11-06
16 | 16 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 3.97 | 65068.902 | 1781481 |  | 65068.902 | 4 | 3.93 | -1.9 | bb | 2023-11-06
17 | 17 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 3.96 | 70109.063 | 1960652 |  | 70109.063 | 4 | 4.23 | 5.7 | bb | 2023-11-06
18 | 18 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 3.96 | 69572.93 | 1929098 |  | 69572.93 | 4 | 4.2 | 4.9 | bb | 2023-11-06
19 | 19 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 3.96 | 72849.406 | 2016856 |  | 72849.406 | 4 | 4.39 | 9.9 | bb | 2023-11-06
20 | 20 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 3.96 | 60371.941 | 1653881 |  | 60371.941 | 4 | 3.64 | -8.9 | bb | 2023-11-06
21 | 21 | 20231106_HepCl_ReRun_021 | Blank | Blank | 3.99 | 7.869 | 293 |  | 7.869 | 4 | 0 | -100 | bb | 2023-11-06
22 | 22 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 3.96 | 64329.785 | 1726832 |  | 64329.785 | 4 | 3.88 | -3 | bb | 2023-11-06
23 | 23 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 3.96 | 72004.906 | 1945760 |  | 72004.906 | 4 | 4.34 | 8.6 | bb | 2023-11-06
24 | 24 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 3.97 | 67768.188 | 1918844 |  | 67768.188 | 4 | 4.09 | 2.2 | bb | 2023-11-06
25 | 25 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 3.97 | 89323.164 | 2487380 |  | 89323.164 | 4 | 5.39 | 34.7 | bb | 2023-11-06
26 | 26 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 3.97 | 75562.414 | 2130806 |  | 75562.414 | 4 | 4.56 | 14 | bb | 2023-11-06
27 | 27 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
28 | 28 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 3.98 | 9.577 | 349 |  | 9.577 | 4 | 0 | -100 | bb | 2023-11-06
29 | 29 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 3.96 | 72754.594 | 1997669 |  | 72754.594 | 4 | 4.39 | 9.7 | bb | 2023-11-06
30 | 30 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 3.97 | 67752.633 | 1891724 |  | 67752.633 | 4 | 4.09 | 2.2 | bb | 2023-11-06
31 | 31 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 3.96 | 65001.293 | 1804041 |  | 65001.293 | 4 | 3.92 | -2 | bb | 2023-11-06
32 | 32 | 20231106_HepCl_ReRun_032 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
33 | 33 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 3.96 | 73487.75 | 2025898 |  | 73487.75 | 4 | 4.43 | 10.8 | bb | 2023-11-06
34 | 34 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 3.96 | 66896.563 | 1878349 |  | 66896.563 | 4 | 4.04 | 0.9 | bb | 2023-11-06
35 | 35 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 3.96 | 66638.734 | 1854201 |  | 66638.734 | 4 | 4.02 | 0.5 | bb | 2023-11-06
36 | 36 | 20231106_HepCl_ReRun_036 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
37 | 37 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 3.96 | 70830.211 | 1970689 |  | 70830.211 | 4 | 4.27 | 6.8 | bb | 2023-11-06
38 | 38 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 3.96 | 64723.23 | 1829755 |  | 64723.23 | 4 | 3.9 | -2.4 | bb | 2023-11-06
39 | 39 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 3.96 | 61013.957 | 1674807 |  | 61013.957 | 4 | 3.68 | -8 | bb | 2023-11-06
40 | 40 | 20231106_HepCl_ReRun_040 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
41 | 41 | 20231106_HepCl_ReRun_041 | QC A | QC | 3.96 | 87338.508 | 2479133 |  | 87338.508 | 4 | 5.27 | 31.7 | bb | 2023-11-06
42 | 42 | 20231106_HepCl_ReRun_042 | QC B | QC | 3.96 | 76308.461 | 2025701 |  | 76308.461 | 4 | 4.6 | 15.1 | bb | 2023-11-06
43 | 43 | 20231106_HepCl_ReRun_043 | QC C | QC | 3.96 | 84343.641 | 2347158 |  | 84343.641 | 4 | 5.09 | 27.2 | bb | 2023-11-06
44 | 44 | 20231106_HepCl_ReRun_044 | QC D | QC | 3.96 | 70742.273 | 1886520 |  | 70742.273 | 4 | 4.27 | 6.7 | bb | 2023-11-06
45 | 45 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 4 | 8.655 | 211 |  | 8.655 | 4 | 0 | -100 | bb | 2023-11-06
46 | 46 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
47 | 47 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 3.96 | 66380.367 | 1787191 |  | 66380.367 | 4 | 4 | 0.1 | bb | 2023-11-06
48 | 48 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 3.96 | 68137.133 | 1884396 |  | 68137.133 | 4 | 4.11 | 2.8 | bb | 2023-11-06
49 | 49 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 3.96 | 60965.863 | 1687426 |  | 60965.863 | 4 | 3.68 | -8.1 | bb | 2023-11-06
50 | 50 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 3.96 | 67549.781 | 1870626 |  | 67549.781 | 4 | 4.08 | 1.9 | bb | 2023-11-06
51 | 51 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 3.96 | 68209.023 | 1921450 |  | 68209.023 | 4 | 4.11 | 2.9 | bb | 2023-11-06
52 | 52 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 3.96 | 64340.457 | 1828969 |  | 64340.457 | 4 | 3.88 | -3 | bb | 2023-11-06
53 | 53 | 20231106_HepCl_ReRun_053 | Blank | Blank | 3.97 | 12.536 | 315 |  | 12.536 | 4 | 0 | -100 | bb | 2023-11-06
54 | 54 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 3.96 | 67724.766 | 1893346 |  | 67724.766 | 4 | 4.09 | 2.1 | bb | 2023-11-06
55 | 55 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 3.96 | 64294.883 | 1769355 |  | 64294.883 | 4 | 3.88 | -3 | bb | 2023-11-06
56 | 56 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 3.96 | 63195.191 | 1738904 |  | 63195.191 | 4 | 3.81 | -4.7 | bb | 2023-11-06
57 | 57 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 3.96 | 63700.488 | 1730554 |  | 63700.488 | 4 | 3.84 | -3.9 | bb | 2023-11-06
58 | 58 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 3.96 | 68771.055 | 1866345 |  | 68771.055 | 4 | 4.15 | 3.7 | bb | 2023-11-06
59 | 59 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 3.96 | 69718.367 | 1978759 |  | 69718.367 | 4 | 4.21 | 5.1 | bb | 2023-11-06
60 | 60 | 20231106_HepCl_ReRun_060 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
61 | 61 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 3.96 | 69190.055 | 1947788 |  | 69190.055 | 4 | 4.17 | 4.4 | bb | 2023-11-06
62 | 62 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 3.96 | 75166.273 | 2079953 |  | 75166.273 | 4 | 4.53 | 13.4 | bb | 2023-11-06
63 | 63 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 3.96 | 64113.281 | 1795978 |  | 64113.281 | 4 | 3.87 | -3.3 | bb | 2023-11-06
64 | 64 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 3.96 | 89929.188 | 2487582 |  | 89929.188 | 4 | 5.43 | 35.6 | bb | 2023-11-06
65 | 65 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 3.97 | 73827.055 | 2077619 |  | 73827.055 | 4 | 4.45 | 11.3 | bb | 2023-11-07
66 | 66 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4 | 0.454 | 37 |  | 0.454 | 4 | 0 | -100 | bb | 2023-11-07
67 | 67 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 3.97 | 12.974 | 336 |  | 12.974 | 4 | 0 | -100 | bb | 2023-11-07
68 | 68 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 3.96 | 70878.07 | 1984627 |  | 70878.07 | 4 | 4.28 | 6.9 | bb | 2023-11-07
69 | 69 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 3.97 | 65752.898 | 1872039 |  | 65752.898 | 4 | 3.97 | -0.8 | bb | 2023-11-07
70 | 70 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 3.96 | 63387.512 | 1742378 |  | 63387.512 | 4 | 3.82 | -4.4 | bb | 2023-11-07
71 | 71 | 20231106_HepCl_ReRun_071 | Blank | Blank | 3.97 | 7.057 | 248 |  | 7.057 | 4 | 0 | -100 | bb | 2023-11-07
72 | 72 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 3.96 | 68083.242 | 1884270 |  | 68083.242 | 4 | 4.11 | 2.7 | bb | 2023-11-07
73 | 73 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 3.96 | 64498.293 | 1769525 |  | 64498.293 | 4 | 3.89 | -2.7 | bb | 2023-11-07
74 | 74 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 3.96 | 62066.836 | 1728402 |  | 62066.836 | 4 | 3.74 | -6.4 | bb | 2023-11-07
75 | 75 | 20231106_HepCl_ReRun_075 | Blank | Blank | 3.97 | 12.113 | 359 |  | 12.113 | 4 | 0 | -100 | bb | 2023-11-07
76 | 76 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 3.96 | 66670.469 | 1849073 |  | 66670.469 | 4 | 4.02 | 0.6 | bb | 2023-11-07
77 | 77 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 3.97 | 63275.16 | 1767799 |  | 63275.16 | 4 | 3.82 | -4.6 | bb | 2023-11-07
78 | 78 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 3.96 | 61289.246 | 1692146 |  | 61289.246 | 4 | 3.7 | -7.6 | bb | 2023-11-07
79 | 79 | 20231106_HepCl_ReRun_079 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
80 | 80 | 20231106_HepCl_ReRun_080 | QC A | QC | 3.96 | 88138.531 | 2488217 |  | 88138.531 | 4 | 5.32 | 32.9 | bb | 2023-11-07
81 | 81 | 20231106_HepCl_ReRun_081 | QC B | QC | 3.96 | 84922.633 | 2354939 |  | 84922.633 | 4 | 5.12 | 28.1 | bb | 2023-11-07
82 | 82 | 20231106_HepCl_ReRun_082 | QC C | QC | 3.96 | 79141.828 | 2191526 |  | 79141.828 | 4 | 4.77 | 19.4 | bb | 2023-11-07
83 | 83 | 20231106_HepCl_ReRun_083 | QC D | QC | 3.96 | 68663.195 | 1862578 |  | 68663.195 | 4 | 4.14 | 3.6 | bb | 2023-11-07
84 | 84 | 20231106_HepCl_ReRun_084 | Blank | Blank | 3.93 | 2.32 | 130 |  | 2.32 | 4 | 0 | -100 | bb | 2023-11-07
85 | 85 | 20231106_HepCl_ReRun_085 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
86 | 86 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 3.96 | 70980.813 | 1997170 |  | 70980.813 | 4 | 4.28 | 7.1 | bb | 2023-11-07
87 | 87 | 20231106_HepCl_ReRun_087 | Blank | Blank | 3.91 | 3.535 | 119 |  | 3.535 | 4 | 0 | -100 | bb | 2023-11-07
88 | 88 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 3.97 | 127111.25 | 3554961 |  | 127111.25 | 4 | 7.67 | 91.7 | bb | 2023-11-07
89 | 89 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 3.96 | 107706.938 | 3042009 |  | 107706.938 | 4 | 6.5 | 62.4 | bb | 2023-11-07
90 | 90 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 3.96 | 113437.414 | 3224395 |  | 113437.414 | 4 | 6.84 | 71.1 | bb | 2023-11-07
91 | 91 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 3.96 | 110987.766 | 3056744 |  | 110987.766 | 4 | 6.7 | 67.4 | bb | 2023-11-07
92 | 92 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 3.96 | 99300.031 | 2693969 |  | 99300.031 | 4 | 5.99 | 49.8 | bb | 2023-11-07
93 | 93 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 3.96 | 105545.609 | 2893700 |  | 105545.609 | 4 | 6.37 | 59.2 | bb | 2023-11-07
94 | 94 | 20231106_HepCl_ReRun_094 | Blank | Blank | 3.97 | 4.894 | 185 |  | 4.894 | 4 | 0 | -100 | bb | 2023-11-07
95 | 95 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 3.96 | 104039.031 | 2808497 |  | 104039.031 | 4 | 6.28 | 56.9 | bb | 2023-11-07
96 | 96 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 3.96 | 119178.984 | 3371710 |  | 119178.984 | 4 | 7.19 | 79.7 | bb | 2023-11-07
97 | 97 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 3.96 | 113962.063 | 3231676 |  | 113962.063 | 4 | 6.87 | 71.9 | bb | 2023-11-07
98 | 98 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 3.96 | 118969.555 | 3318540 |  | 118969.555 | 4 | 7.18 | 79.4 | bb | 2023-11-07
99 | 99 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 3.96 | 122033.266 | 3442648 |  | 122033.266 | 4 | 7.36 | 84 | bb | 2023-11-07
100 | 100 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 3.96 | 110901.977 | 3062770 |  | 110901.977 | 4 | 6.69 | 67.3 | bb | 2023-11-07
101 | 101 | 20231106_HepCl_ReRun_101 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
102 | 102 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 3.96 | 94312.25 | 2578495 |  | 94312.25 | 4 | 5.69 | 42.2 | bb | 2023-11-07
103 | 103 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 3.96 | 101210.508 | 2857656 |  | 101210.508 | 4 | 6.11 | 52.6 | bb | 2023-11-07
104 | 104 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 3.96 | 104400.188 | 2856369 |  | 104400.188 | 4 | 6.3 | 57.5 | bb | 2023-11-07
105 | 105 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 3.96 | 84769.352 | 2328900 |  | 84769.352 | 4 | 5.11 | 27.8 | bb | 2023-11-07
106 | 106 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 3.96 | 88387.469 | 2483812 |  | 88387.469 | 4 | 5.33 | 33.3 | bb | 2023-11-07
107 | 107 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 3.96 | 75575.227 | 2013143 |  | 75575.227 | 4 | 4.56 | 14 | bb | 2023-11-07
108 | 108 | 20231106_HepCl_ReRun_108 | Blank | Blank | 3.97 | 3.622 | 139 |  | 3.622 | 4 | 0 | -100 | bb | 2023-11-07
109 | 109 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 3.96 | 68778 | 1876567 |  | 68778 | 4 | 4.15 | 3.7 | bb | 2023-11-07
110 | 110 | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.01 | 1.239 | 97 |  | 1.239 | 4 | 0 | -100 | bb | 2023-11-07
111 | 111 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 3.96 | 98270.578 | 2691723 |  | 98270.578 | 4 | 5.93 | 48.2 | bb | 2023-11-07
112 | 112 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 3.96 | 98628.742 | 2774183 |  | 98628.742 | 4 | 5.95 | 48.7 | bb | 2023-11-07
113 | 113 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 3.96 | 106868.266 | 3033461 |  | 106868.266 | 4 | 6.45 | 61.2 | bb | 2023-11-07
114 | 114 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 3.96 | 104541.266 | 2937157 |  | 104541.266 | 4 | 6.31 | 57.7 | bb | 2023-11-07
115 | 115 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 3.96 | 112556.063 | 3156701 |  | 112556.063 | 4 | 6.79 | 69.8 | bb | 2023-11-07
116 | 116 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 3.96 | 104859.008 | 2910590 |  | 104859.008 | 4 | 6.33 | 58.1 | bb | 2023-11-07
117 | 117 | 20231106_HepCl_ReRun_117 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
118 | 118 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 3.96 | 105383.109 | 2935355 |  | 105383.109 | 4 | 6.36 | 58.9 | bb | 2023-11-07
119 | 119 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 3.96 | 99879.406 | 2791654 |  | 99879.406 | 4 | 6.03 | 50.6 | bb | 2023-11-07
120 | 120 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 3.96 | 94358.172 | 2525923 |  | 94358.172 | 4 | 5.69 | 42.3 | bb | 2023-11-07
121 | 121 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 3.96 | 116692.453 | 3338579 |  | 116692.453 | 4 | 7.04 | 76 | bb | 2023-11-07
122 | 122 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 3.96 | 107368.117 | 3028922 |  | 107368.117 | 4 | 6.48 | 61.9 | bb | 2023-11-07
123 | 123 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 3.96 | 109868.875 | 3072305 |  | 109868.875 | 4 | 6.63 | 65.7 | bb | 2023-11-07
124 | 124 | 20231106_HepCl_ReRun_124 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
125 | 125 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 3.96 | 102934.406 | 2925716 |  | 102934.406 | 4 | 6.21 | 55.2 | bb | 2023-11-07
126 | 126 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 3.96 | 95880.813 | 2743191 |  | 95880.813 | 4 | 5.78 | 44.6 | bb | 2023-11-07
127 | 127 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 3.96 | 111986.016 | 3190599 |  | 111986.016 | 4 | 6.76 | 68.9 | bb | 2023-11-07
128 | 128 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 3.96 | 84980.797 | 2345692 |  | 84980.797 | 4 | 5.13 | 28.2 | bb | 2023-11-07
129 | 129 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 3.96 | 64520.02 | 1765655 |  | 64520.02 | 4 | 3.89 | -2.7 | bb | 2023-11-07
130 | 130 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 3.96 | 77287.742 | 2053044 |  | 77287.742 | 4 | 4.66 | 16.6 | bb | 2023-11-07
131 | 131 | 20231106_HepCl_ReRun_131 | Blank | Blank | 3.92 | 14.521 | 393 |  | 14.521 | 4 | 0 | -100 | bb | 2023-11-07
132 | 132 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 3.96 | 65707.844 | 1818297 |  | 65707.844 | 4 | 3.96 | -0.9 | bb | 2023-11-07
133 | 133 | 20231106_HepCl_ReRun_133 | Blank | Blank | 3.97 | 7.49 | 223 |  | 7.49 | 4 | 0 | -100 | bb | 2023-11-07
134 | 134 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 3.97 | 96807.188 | 2668441 |  | 96807.188 | 4 | 5.84 | 46 | bb | 2023-11-07
135 | 135 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 3.96 | 110827.492 | 3028426 |  | 110827.492 | 4 | 6.69 | 67.1 | bb | 2023-11-07
136 | 136 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 3.96 | 106620.406 | 3040845 |  | 106620.406 | 4 | 6.43 | 60.8 | bb | 2023-11-07
137 | 137 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 3.96 | 86723.508 | 2258546 |  | 86723.508 | 4 | 5.23 | 30.8 | bb | 2023-11-07
138 | 138 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 3.96 | 97677.023 | 2650361 |  | 97677.023 | 4 | 5.89 | 47.3 | bb | 2023-11-07
139 | 139 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 3.96 | 100329.82 | 2834738 |  | 100329.82 | 4 | 6.05 | 51.3 | bb | 2023-11-07
140 | 140 | 20231106_HepCl_ReRun_140 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
141 | 141 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 3.96 | 64887.246 | 1825407 |  | 64887.246 | 4 | 3.91 | -2.1 | bb | 2023-11-07
142 | 142 | 20231106_HepCl_ReRun_142 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
143 | 143 | 20231106_HepCl_ReRun_143 | Blank | Blank | 3.92 | 2.895 | 130 |  | 2.895 | 4 | 0 | -100 | bb | 2023-11-07
144 | 144 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 3.96 | 120155.836 | 3332802 |  | 120155.836 | 4 | 7.25 | 81.2 | bb | 2023-11-07
145 | 145 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 3.96 | 112531.641 | 3182520 |  | 112531.641 | 4 | 6.79 | 69.7 | bb | 2023-11-07
146 | 146 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 3.96 | 118805.617 | 3380132 |  | 118805.617 | 4 | 7.17 | 79.2 | bb | 2023-11-07
147 | 147 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 3.96 | 102610.383 | 2811674 |  | 102610.383 | 4 | 6.19 | 54.8 | bb | 2023-11-07
148 | 148 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 3.96 | 102214.219 | 2882559 |  | 102214.219 | 4 | 6.17 | 54.2 | bb | 2023-11-07
149 | 149 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 3.96 | 99535.336 | 2741889 |  | 99535.336 | 4 | 6 | 50.1 | bb | 2023-11-07
150 | 150 | 20231106_HepCl_ReRun_152 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
151 | 151 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 3.96 | 103807.086 | 2960972 |  | 103807.086 | 4 | 6.26 | 56.6 | bb | 2023-11-07
152 | 152 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 3.96 | 99959.609 | 2879778 |  | 99959.609 | 4 | 6.03 | 50.8 | bb | 2023-11-07
153 | 153 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 3.96 | 100739.109 | 2878863 |  | 100739.109 | 4 | 6.08 | 51.9 | bb | 2023-11-07
154 | 154 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 3.96 | 95184.141 | 2623040 |  | 95184.141 | 4 | 5.74 | 43.6 | bb | 2023-11-07
155 | 155 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 3.96 | 106147.617 | 3012649 |  | 106147.617 | 4 | 6.4 | 60.1 | bb | 2023-11-07
156 | 156 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 3.96 | 98089.266 | 2678359 |  | 98089.266 | 4 | 5.92 | 47.9 | bb | 2023-11-07
157 | 157 | 20231106_HepCl_ReRun_159 | Blank | Blank | 3.95 | 4.77 | 219 |  | 4.77 | 4 | 0 | -100 | bb | 2023-11-07
158 | 158 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 3.96 | 85510.195 | 2336976 |  | 85510.195 | 4 | 5.16 | 29 | bb | 2023-11-07
159 | 159 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 3.96 | 92018.813 | 2410122 |  | 92018.813 | 4 | 5.55 | 38.8 | bb | 2023-11-07
160 | 160 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 3.97 | 87029.469 | 2377545 |  | 87029.469 | 4 | 5.25 | 31.3 | bb | 2023-11-07
161 | 161 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 3.96 | 89270.031 | 2472065 |  | 89270.031 | 4 | 5.39 | 34.6 | bb | 2023-11-07
162 | 162 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 3.96 | 85731.852 | 2440697 |  | 85731.852 | 4 | 5.17 | 29.3 | bb | 2023-11-07
163 | 163 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 3.96 | 82351.266 | 2320337 |  | 82351.266 | 4 | 4.97 | 24.2 | bb | 2023-11-07
164 | 164 | 20231106_HepCl_ReRun_166 | Blank | Blank | 3.96 | 6.376 | 199 |  | 6.376 | 4 | 0 | -100 | bb | 2023-11-07
165 | 165 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 3.96 | 62184.316 | 1723499 |  | 62184.316 | 4 | 3.75 | -6.2 | bb | 2023-11-07
166 | 166 | 20231106_HepCl_ReRun_168 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
167 | 167 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 3.96 | 96676.672 | 2624358 |  | 96676.672 | 4 | 5.83 | 45.8 | bb | 2023-11-07
168 | 168 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 3.96 | 106880.938 | 3036519 |  | 106880.938 | 4 | 6.45 | 61.2 | bb | 2023-11-07
169 | 169 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 3.96 | 109563.195 | 3144449 |  | 109563.195 | 4 | 6.61 | 65.2 | bb | 2023-11-07
170 | 170 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 3.97 | 99443.953 | 2784842 |  | 99443.953 | 4 | 6 | 50 | bb | 2023-11-07
171 | 171 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 3.96 | 106413.328 | 3078344 |  | 106413.328 | 4 | 6.42 | 60.5 | bb | 2023-11-07
172 | 172 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 3.97 | 109031.305 | 3128109 |  | 109031.305 | 4 | 6.58 | 64.4 | bb | 2023-11-07
173 | 173 | 20231106_HepCl_ReRun_175 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
174 | 174 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 3.96 | 100056 | 2828201 |  | 100056 | 4 | 6.04 | 50.9 | bb | 2023-11-07
175 | 175 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 3.96 | 89716.656 | 2419546 |  | 89716.656 | 4 | 5.41 | 35.3 | bb | 2023-11-07
176 | 176 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 3.97 | 109559.086 | 2959963 |  | 109559.086 | 4 | 6.61 | 65.2 | bb | 2023-11-07
177 | 177 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 3.96 | 90430.555 | 2440787 |  | 90430.555 | 4 | 5.46 | 36.4 | bb | 2023-11-07
178 | 178 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 3.97 | 102643.82 | 2885297 |  | 102643.82 | 4 | 6.19 | 54.8 | bb | 2023-11-07
179 | 179 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 3.96 | 96604.727 | 2642798 |  | 96604.727 | 4 | 5.83 | 45.7 | bb | 2023-11-07
180 | 180 | 20231106_HepCl_ReRun_182 | Blank | Blank | 3.96 | 14.625 | 344 |  | 14.625 | 4 | 0 | -100 | bb | 2023-11-07
181 | 181 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 3.96 | 105342.867 | 3033331 |  | 105342.867 | 4 | 6.35 | 58.9 | bb | 2023-11-07
182 | 182 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 3.97 | 98521.195 | 2867227 |  | 98521.195 | 4 | 5.94 | 48.6 | bb | 2023-11-07
183 | 183 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 3.97 | 87807.195 | 2354495 |  | 87807.195 | 4 | 5.3 | 32.4 | bb | 2023-11-07
184 | 184 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 3.97 | 84955.523 | 2334746 |  | 84955.523 | 4 | 5.13 | 28.1 | bb | 2023-11-07
185 | 185 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 3.96 | 83049.625 | 2385833 |  | 83049.625 | 4 | 5.01 | 25.3 | bb | 2023-11-07
186 | 186 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 3.97 | 75152.219 | 2040597 |  | 75152.219 | 4 | 4.53 | 13.3 | bb | 2023-11-07
187 | 187 | 20231106_HepCl_ReRun_189 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
188 | 188 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 3.96 | 65820.523 | 1851785 |  | 65820.523 | 4 | 3.97 | -0.7 | bb | 2023-11-07
189 | 189 | 20231106_HepCl_ReRun_191 | Blank | Blank | 3.92 | 6.924 | 263 |  | 6.924 | 4 | 0 | -100 | bb | 2023-11-07
190 | 190 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 3.96 | 102671.5 | 2832084 |  | 102671.5 | 4 | 6.19 | 54.8 | bb | 2023-11-07
191 | 191 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 3.96 | 101879.82 | 2750035 |  | 101879.82 | 4 | 6.15 | 53.7 | bb | 2023-11-07
192 | 192 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 3.97 | 119917.711 | 3374345 |  | 119917.711 | 4 | 7.23 | 80.9 | bb | 2023-11-07
193 | 193 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 3.96 | 97705.281 | 2789053 |  | 97705.281 | 4 | 5.89 | 47.4 | bb | 2023-11-07
194 | 194 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 3.96 | 102602.477 | 2908390 |  | 102602.477 | 4 | 6.19 | 54.7 | bb | 2023-11-07
195 | 195 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 3.96 | 91857.102 | 2594836 |  | 91857.102 | 4 | 5.54 | 38.5 | bb | 2023-11-07
196 | 196 | 20231106_HepCl_ReRun_198 | Blank | Blank | 3.92 | 3.35 | 148 |  | 3.35 | 4 | 0 | -100 | bb | 2023-11-08
197 | 197 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 3.97 | 51659.656 | 1392823 |  | 51659.656 | 4 | 3.12 | -22.1 | bb | 2023-11-08
198 | 198 | 20231106_HepCl_ReRun_200 | Blank | Blank | 4 | 2.811 | 117 |  | 2.811 | 4 | 0 | -100 | bb | 2023-11-08
199 | 199 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 3.97 | 122806.492 | 3541354 |  | 122806.492 | 4 | 7.41 | 85.2 | bb | 2023-11-08
200 | 200 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 3.96 | 105540.742 | 2950542 |  | 105540.742 | 4 | 6.37 | 59.2 | bb | 2023-11-08
201 | 201 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 3.96 | 108226.414 | 2970949 |  | 108226.414 | 4 | 6.53 | 63.2 | bb | 2023-11-08
202 | 202 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 3.96 | 98816.094 | 2714665 |  | 98816.094 | 4 | 5.96 | 49 | bb | 2023-11-08
203 | 203 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 3.96 | 114010.602 | 3199621 |  | 114010.602 | 4 | 6.88 | 71.9 | bb | 2023-11-08
204 | 204 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 3.97 | 111625.508 | 3109118 |  | 111625.508 | 4 | 6.73 | 68.3 | bb | 2023-11-08
205 | 205 | 20231106_HepCl_ReRun_207 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
206 | 206 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 3.96 | 111956.539 | 3172459 |  | 111956.539 | 4 | 6.75 | 68.8 | bb | 2023-11-08
207 | 207 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 3.96 | 104588.813 | 2995732 |  | 104588.813 | 4 | 6.31 | 57.7 | bb | 2023-11-08
208 | 208 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 3.96 | 87999.492 | 2312692 |  | 87999.492 | 4 | 5.31 | 32.7 | bb | 2023-11-08
209 | 209 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 3.97 | 97360.305 | 2799244 |  | 97360.305 | 4 | 5.87 | 46.8 | bb | 2023-11-08
210 | 210 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 3.96 | 84831.938 | 2370620 |  | 84831.938 | 4 | 5.12 | 27.9 | bb | 2023-11-08
211 | 211 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 3.96 | 92662.234 | 2609777 |  | 92662.234 | 4 | 5.59 | 39.7 | bb | 2023-11-08
212 | 212 | 20231106_HepCl_ReRun_214 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
213 | 213 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 3.96 | 103418.438 | 3004115 |  | 103418.438 | 4 | 6.24 | 56 | bb | 2023-11-08
214 | 214 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 3.97 | 106341.625 | 3036697 |  | 106341.625 | 4 | 6.42 | 60.4 | bb | 2023-11-08
215 | 215 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 3.97 | 92977.766 | 2623362 |  | 92977.766 | 4 | 5.61 | 40.2 | bb | 2023-11-08
216 | 216 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 3.97 | 79338.781 | 2209920 |  | 79338.781 | 4 | 4.79 | 19.7 | bb | 2023-11-08
217 | 217 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 3.97 | 92511.57 | 2666977 |  | 92511.57 | 4 | 5.58 | 39.5 | bb | 2023-11-08
218 | 218 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 3.97 | 85295.422 | 2390381 |  | 85295.422 | 4 | 5.15 | 28.6 | bb | 2023-11-08
219 | 219 | 20231106_HepCl_ReRun_221 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
220 | 220 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 3.96 | 59542.715 | 1668513 |  | 59542.715 | 4 | 3.59 | -10.2 | bb | 2023-11-08
221 | 221 | 20231106_HepCl_ReRun_223 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
222 | 222 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 3.96 | 110140.234 | 3188337 |  | 110140.234 | 4 | 6.64 | 66.1 | bb | 2023-11-08
223 | 223 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 3.97 | 108602.188 | 3147145 |  | 108602.188 | 4 | 6.55 | 63.8 | bb | 2023-11-08
224 | 224 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 3.96 | 90797.094 | 2563937 |  | 90797.094 | 4 | 5.48 | 36.9 | bb | 2023-11-08
225 | 225 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 3.97 | 111401.727 | 3155190 |  | 111401.727 | 4 | 6.72 | 68 | bb | 2023-11-08
226 | 226 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 3.96 | 110531.492 | 3088762 |  | 110531.492 | 4 | 6.67 | 66.7 | bb | 2023-11-08
227 | 227 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 3.97 | 112650.75 | 3159743 |  | 112650.75 | 4 | 6.8 | 69.9 | bb | 2023-11-08
228 | 228 | 20231106_HepCl_ReRun_230 | Blank | Blank | 3.96 | 9.549 | 291 |  | 9.549 | 4 | 0 | -100 | bb | 2023-11-08
229 | 229 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 3.97 | 113099.43 | 3265188 |  | 113099.43 | 4 | 6.82 | 70.6 | bb | 2023-11-08
230 | 230 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 3.96 | 91099.672 | 2589675 |  | 91099.672 | 4 | 5.5 | 37.4 | bb | 2023-11-08
231 | 231 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 3.97 | 95262.484 | 2696060 |  | 95262.484 | 4 | 5.75 | 43.7 | bb | 2023-11-08
232 | 232 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 3.96 | 95167.969 | 2738875 |  | 95167.969 | 4 | 5.74 | 43.5 | bb | 2023-11-08
233 | 233 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 3.97 | 101755.492 | 2993356 |  | 101755.492 | 4 | 6.14 | 53.5 | bb | 2023-11-08
234 | 234 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 3.97 | 87558.133 | 2520853 |  | 87558.133 | 4 | 5.28 | 32.1 | bb | 2023-11-08
235 | 235 | 20231106_HepCl_ReRun_237 | Blank | Blank | 3.98 | 0.334 | 38 |  | 0.334 | 4 | 0 | -100 | bb | 2023-11-08
236 | 236 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 3.97 | 118122.609 | 3489290 |  | 118122.609 | 4 | 7.13 | 78.1 | bb | 2023-11-08
237 | 237 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 3.97 | 102091.367 | 2883097 |  | 102091.367 | 4 | 6.16 | 54 | bb | 2023-11-08
238 | 238 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 3.97 | 125547.516 | 3727889 |  | 125547.516 | 4 | 7.57 | 89.3 | bb | 2023-11-08
239 | 239 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 3.97 | 96524.773 | 2796460 |  | 96524.773 | 4 | 5.82 | 45.6 | bb | 2023-11-08
240 | 240 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 3.97 | 88051.555 | 2541873 |  | 88051.555 | 4 | 5.31 | 32.8 | bb | 2023-11-08
241 | 241 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 3.97 | 96476.289 | 2806186 |  | 96476.289 | 4 | 5.82 | 45.5 | bb | 2023-11-08
242 | 242 | 20231106_HepCl_ReRun_244 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
243 | 243 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 3.97 | 53564.039 | 1450588 |  | 53564.039 | 4 | 3.23 | -19.2 | bb | 2023-11-08
244 | 244 | 20231106_HepCl_ReRun_246 | Blank | Blank | 3.96 | 1.2 | 64 |  | 1.2 | 4 | 0 | -100 | bb | 2023-11-08
245 | 245 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 3.97 | 120083.727 | 3464551 |  | 120083.727 | 4 | 7.24 | 81.1 | bb | 2023-11-08
246 | 246 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 3.97 | 120913.375 | 3529813 |  | 120913.375 | 4 | 7.29 | 82.4 | bb | 2023-11-08
247 | 247 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 3.97 | 120947.359 | 3381191 |  | 120947.359 | 4 | 7.3 | 82.4 | bb | 2023-11-08
248 | 248 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 3.97 | 98304.648 | 2728535 |  | 98304.648 | 4 | 5.93 | 48.3 | bb | 2023-11-08
249 | 249 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 3.97 | 108873.719 | 3186634 |  | 108873.719 | 4 | 6.57 | 64.2 | bb | 2023-11-08
250 | 250 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 3.97 | 104880.836 | 3105835 |  | 104880.836 | 4 | 6.33 | 58.2 | bb | 2023-11-08
251 | 251 | 20231106_HepCl_ReRun_253 | Blank | Blank | 3.96 | 12.938 | 311 |  | 12.938 | 4 | 0 | -100 | bb | 2023-11-08
252 | 252 | 20231106_HepCl_ReRun_254 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
253 | 253 | 20231106_HepCl_ReRun_255 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
Compound 6: 8:2 FTS
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20231106_HepCl_ReRun_001 | Blank | Blank |  |  |  | 2.685 | [5 empty] | 2023-11-06
2 | 2 | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.06 | 1.228 | 71 | 3.534 | 1.39 |  | 7.33 |  | bb | 2023-11-06
3 | 3 | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.04 | 4.574 | 189 | [5 empty] | bb | 2023-11-06
4 | 4 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.08 | 1.594 | 79 | 0.86 | 7.414 |  | 43.47 |  | bb | 2023-11-06
5 | 5 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.06 | 0.326 | 19 | 0.572 | 2.28 |  | 12.67 |  | bb | 2023-11-06
6 | 6 | 20231106_HepCl_ReRun_006 | Blank | Blank |  |  |  | 0.473 | [5 empty] | 2023-11-06
7 | 7 | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.05 | 0.796 | 47 | 1.479 | 2.153 |  | 11.91 |  | bb | 2023-11-06
8 | 8 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 4.05 | 72.711 | 2860 | 1404.893 | 0.207 | 0.175 | 0.23 | 33.9 | bb | 2023-11-06
9 | 9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 4.05 | 68.343 | 2531 | 1292.746 | 0.211 | 0.281 | 0.26 | -7.2 | bb | 2023-11-06
10 | 10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 4.05 | 57.947 | 2438 | 1261.171 | 0.184 | 0.444 | 0.09 | -78.6 | bb | 2023-11-06
11 | 11 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 4.05 | 158.026 | 5515 | 1380.42 | 0.458 | 0.713 | 1.74 | 144.1 | bb | 2023-11-06
12 | 12 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 4.05 | 145.233 | 5514 | 1318.556 | 0.441 | 1.138 | 1.64 | 43.8 | bb | 2023-11-06
13 | 13 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.05 | 183.985 | 6873 | 1285.358 | 0.573 | 1.825 | 2.43 | 33 | bb | 2023-11-06
14 | 14 | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.06 | 0.207 | 12 | 0.151 | 5.483 |  | 31.89 |  | bb | 2023-11-06
15 | 15 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.05 | 221.348 | 8806 | 1359.29 | 0.651 | 2.919 | 2.9 | -0.6 | bb | 2023-11-06
16 | 16 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.05 | 362.115 | 14411 | 1397.361 | 1.037 | 4.669 | 5.21 | 11.6 | bb | 2023-11-06
17 | 17 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.05 | 517.63 | 19667 | 1364.509 | 1.517 | 7.45 | 8.1 | 8.7 | bb | 2023-11-06
18 | 18 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.05 | 764.35 | 28612 | 1367.205 | 2.236 | 11.875 | 12.41 | 4.5 | bb | 2023-11-06
19 | 19 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.05 | 1354.45 | 51659 | 1491.731 | 3.632 | 19.125 | 20.78 | 8.7 | bb | 2023-11-06
20 | 20 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.05 | 2082.749 | 78265 | 1313.037 | 6.345 | 30.5 | 37.06 | 21.5 | bb | 2023-11-06
21 | 21 | 20231106_HepCl_ReRun_021 | Blank | Blank |  |  |  | 3.142 | [5 empty] | 2023-11-06
22 | 22 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.05 | 2907.99 | 115954 | 1391.243 | 8.361 | 48.75 | 49.15 | 0.8 | bb | 2023-11-06
23 | 23 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.05 | 4992.377 | 192789 | 1497.735 | 13.333 | 78.125 | 78.98 | 1.1 | bb | 2023-11-06
24 | 24 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.05 | 7282.961 | 267543 | 1422.099 | 20.485 | 125 | 121.89 | -2.5 | bb | 2023-11-06
25 | 25 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 4.05 | 11365.859 | 404654 | 1729.245 | 26.291 | 200 | 156.72 | -21.6 | bb | 2023-11-06
26 | 26 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 4.05 | 18668.201 | 676678 | 1482.366 | 50.374 | 325.625 | 301.21 | -7.5 | bb | 2023-11-06
27 | 27 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank |  |  |  | 0.598 | [5 empty] | 2023-11-06
28 | 28 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.08 | 0.433 | 25 | [5 empty] | bb | 2023-11-06
29 | 29 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 4.05 | 143.369 | 5623 | 1378.086 | 0.416 | 0.713 | 1.49 | 109 | bb | 2023-11-06
30 | 30 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 4.05 | 121.842 | 4186 | 1331.639 | 0.366 | 1.138 | 1.19 | 4.4 | bb | 2023-11-06
31 | 31 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.05 | 134.609 | 4719 | 1266.854 | 0.425 | 1.825 | 1.54 | -15.5 | bb | 2023-11-06
32 | 32 | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 0.372 | [5 empty] | 2023-11-06
33 | 33 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 4.05 | 150.419 | 6273 | 1402.833 | 0.429 | 0.713 | 1.57 | 119.7 | bb | 2023-11-06
34 | 34 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 4.05 | 113.359 | 4706 | 1393.114 | 0.325 | 1.138 | 0.94 | -16.9 | bb | 2023-11-06
35 | 35 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.05 | 139.414 | 5172 | 1260.427 | 0.442 | 1.825 | 1.65 | -9.8 | bb | 2023-11-06
36 | 36 | 20231106_HepCl_ReRun_036 | Blank | Blank | 4.06 | 1.437 | 84 | 0.326 | 17.632 |  | 104.77 |  | bb | 2023-11-06
37 | 37 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 4.05 | 144.658 | 5667 | 1405.047 | 0.412 | 0.713 | 1.46 | 105.3 | bb | 2023-11-06
38 | 38 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 4.05 | 129.741 | 5334 | 1291.964 | 0.402 | 1.138 | 1.4 | 23.3 | bb | 2023-11-06
39 | 39 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.05 | 140.109 | 4945 | 1310.875 | 0.428 | 1.825 | 1.56 | -14.7 | bb | 2023-11-06
40 | 40 | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.05 | 0.854 | 50 | [5 empty] | bb | 2023-11-06
41 | 41 | 20231106_HepCl_ReRun_041 | QC A | QC | 4.05 | 83.107 | 3237 | 1704.962 | 0.195 | 0.625 | 0.16 | -74.1 | bb | 2023-11-06
42 | 42 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.05 | 231.524 | 9385 | 1697.607 | 0.546 | 2.506 | 2.27 | -9.6 | bb | 2023-11-06
43 | 43 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.05 | 564.021 | 21717 | 1749.339 | 1.29 | 6.269 | 6.73 | 7.4 | bb | 2023-11-06
44 | 44 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.05 | 1745.07 | 65258 | 1516.869 | 4.602 | 25 | 26.6 | 6.4 | bb | 2023-11-06
45 | 45 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 0.278 | [5 empty] | 2023-11-06
46 | 46 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.02 | 0.155 | 15 | 0.087 | 7.126 |  | 41.75 |  | bb | 2023-11-06
47 | 47 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 4.05 | 62.805 | 2424 | 1396.361 | 0.18 | 0.175 | 0.07 | -59.1 | bb | 2023-11-06
48 | 48 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 4.05 | 45.842 | 1834 | 1232.084 | 0.149 | 0.281 |  |  | bbI | 2023-11-06
49 | 49 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 4.05 | 50.07 | 2160 | 1101.421 | 0.182 | 0.444 | 0.08 | -81.3 | bb | 2023-11-06
50 | 50 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 4.05 | 147.42 | 5938 | 1415.325 | 0.417 | 0.713 | 1.49 | 109.4 | bb | 2023-11-06
51 | 51 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 4.05 | 123.641 | 5052 | 1218 | 0.406 | 1.138 | 1.43 | 25.6 | bb | 2023-11-06
52 | 52 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.05 | 146.812 | 5980 | 1179.859 | 0.498 | 1.825 | 1.98 | 8.4 | bb | 2023-11-06
53 | 53 | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 0.376 | [5 empty] | 2023-11-06
54 | 54 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.05 | 193.381 | 7888 | 1307.154 | 0.592 | 2.919 | 2.54 | -12.9 | bb | 2023-11-06
55 | 55 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.05 | 279.907 | 10543 | 1383.121 | 0.809 | 4.669 | 3.85 | -17.6 | bb | 2023-11-06
56 | 56 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.05 | 502.933 | 19943 | 1332.81 | 1.509 | 7.45 | 8.05 | 8 | bb | 2023-11-06
57 | 57 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.05 | 682.924 | 25853 | 1340.69 | 2.038 | 11.875 | 11.22 | -5.5 | bb | 2023-11-06
58 | 58 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.05 | 1244.492 | 46411 | 1394.592 | 3.569 | 19.125 | 20.41 | 6.7 | bb | 2023-11-06
59 | 59 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.05 | 2029.725 | 77604 | 1381.883 | 5.875 | 30.5 | 34.24 | 12.3 | bb | 2023-11-06
60 | 60 | 20231106_HepCl_ReRun_060 | Blank | Blank |  |  |  | 0.214 | [5 empty] | 2023-11-06
61 | 61 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.05 | 2793.315 | 105829 | 1344.772 | 8.309 | 48.75 | 48.84 | 0.2 | bb | 2023-11-06
62 | 62 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.05 | 4922.576 | 187116 | 1523.123 | 12.928 | 78.125 | 76.55 | -2 | bb | 2023-11-06
63 | 63 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.05 | 6916.329 | 258533 | 1427.197 | 19.384 | 125 | 115.29 | -7.8 | bb | 2023-11-06
64 | 64 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 4.05 | 11634.63 | 427689 | 1844.922 | 25.225 | 200 | 150.33 | -24.8 | bb | 2023-11-06
65 | 65 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 4.05 | 18651.377 | 669332 | 1484.605 | 50.253 | 325.625 | 300.48 | -7.7 | bb | 2023-11-07
66 | 66 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 3.149 | [5 empty] | 2023-11-07
67 | 67 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.05 | 0.073 | 18 | 0.348 | 0.839 |  | 4.03 |  | bb | 2023-11-07
68 | 68 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 4.05 | 124.174 | 4828 | 1319.151 | 0.377 | 0.713 | 1.25 | 75.6 | bb | 2023-11-07
69 | 69 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 4.05 | 102.804 | 3685 | 1336.031 | 0.308 | 1.138 | 0.84 | -26.3 | bb | 2023-11-07
70 | 70 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.05 | 120.27 | 4640 | 1220.534 | 0.394 | 1.825 | 1.36 | -25.6 | bb | 2023-11-07
71 | 71 | 20231106_HepCl_ReRun_071 | Blank | Blank | [9 empty] | 2023-11-07
72 | 72 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 4.05 | 135.904 | 5456 | 1228.853 | 0.442 | 0.713 | 1.65 | 131.1 | bb | 2023-11-07
73 | 73 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 4.05 | 102.66 | 4075 | 1286.003 | 0.319 | 1.138 | 0.91 | -20.2 | bb | 2023-11-07
74 | 74 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.05 | 143.867 | 6016 | 1181.049 | 0.487 | 1.825 | 1.92 | 5 | bb | 2023-11-07
75 | 75 | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.06 | 0.203 | 12 | [5 empty] | bb | 2023-11-07
76 | 76 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 4.05 | 113.447 | 4183 | 1246.689 | 0.364 | 0.713 | 1.18 | 65.1 | bb | 2023-11-07
77 | 77 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 4.05 | 100.041 | 3971 | 1098.067 | 0.364 | 1.138 | 1.18 | 3.6 | bb | 2023-11-07
78 | 78 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.05 | 129.094 | 5605 | 1218.418 | 0.424 | 1.825 | 1.53 | -15.9 | bb | 2023-11-07
79 | 79 | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 0.232 | [5 empty] | 2023-11-07
80 | 80 | 20231106_HepCl_ReRun_080 | QC A | QC | 4.05 | 74.95 | 2549 | 1712.896 | 0.175 | 0.625 | 0.04 | -93.2 | bb | 2023-11-07
81 | 81 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.05 | 222.404 | 8528 | 1702.264 | 0.523 | 2.506 | 2.13 | -15.1 | bb | 2023-11-07
82 | 82 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.05 | 532.346 | 19977 | 1712.841 | 1.243 | 6.269 | 6.45 | 2.9 | bb | 2023-11-07
83 | 83 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.05 | 1766.013 | 67718 | 1490.915 | 4.738 | 25 | 27.42 | 9.7 | bb | 2023-11-07
84 | 84 | 20231106_HepCl_ReRun_084 | Blank | Blank |  |  |  | 0.156 | [5 empty] | 2023-11-07
85 | 85 | 20231106_HepCl_ReRun_085 | Blank | Blank | [9 empty] | 2023-11-07
86 | 86 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 4.05 | 1957.167 | 73926 | 1311.876 | 5.968 | 30.5 | 34.79 | 14.1 | bb | 2023-11-07
87 | 87 | 20231106_HepCl_ReRun_087 | Blank | Blank | [9 empty] | 2023-11-07
88 | 88 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.05 | 455.547 | 17598 | 2610.134 | 0.698 |  | 3.18 |  | bb | 2023-11-07
89 | 89 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.05 | 404.303 | 15376 | 2373.515 | 0.681 |  | 3.08 |  | bb | 2023-11-07
90 | 90 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.05 | 382.016 | 15123 | 2397.312 | 0.637 |  | 2.82 |  | bb | 2023-11-07
91 | 91 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.05 | 153.025 | 6504 | 2359.211 | 0.259 |  | 0.55 |  | bb | 2023-11-07
92 | 92 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.05 | 408.419 | 15643 | 2233.017 | 0.732 |  | 3.38 |  | bb | 2023-11-07
93 | 93 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 164.305 | 6189 | 2234.758 | 0.294 |  | 0.76 |  | bb | 2023-11-07
94 | 94 | 20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 1.338 | [5 empty] | 2023-11-07
95 | 95 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.05 | 142.847 | 6171 | 2519.11 | 0.227 |  | 0.35 |  | bb | 2023-11-07
96 | 96 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.05 | 149.053 | 5522 | 2439.5 | 0.244 |  | 0.46 |  | bb | 2023-11-07
97 | 97 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.05 | 261.987 | 10292 | 2235.031 | 0.469 |  | 1.81 |  | bb | 2023-11-07
98 | 98 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.05 | 218.786 | 8347 | 2343.231 | 0.373 |  | 1.23 |  | bb | 2023-11-07
99 | 99 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.05 | 139.54 | 5755 | 2400.613 | 0.233 |  | 0.39 |  | bb | 2023-11-07
100 | 100 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.05 | 143.163 | 5696 | 2513.042 | 0.228 |  | 0.36 |  | bb | 2023-11-07
101 | 101 | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.06 | 1.23 | 72 | [5 empty] | bb | 2023-11-07
102 | 102 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.05 | 231.241 | 9183 | 2126.96 | 0.435 |  | 1.6 |  | bb | 2023-11-07
103 | 103 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.05 | 357.515 | 13825 | 2142.216 | 0.668 |  | 3 |  | bb | 2023-11-07
104 | 104 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.05 | 167.013 | 6154 | 2372.858 | 0.282 |  | 0.68 |  | bb | 2023-11-07
105 | 105 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.05 | 315.459 | 11990 | 1960.445 | 0.644 |  | 2.85 |  | bb | 2023-11-07
106 | 106 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.05 | 136.141 | 5084 | 1928.79 | 0.282 |  | 0.69 |  | bb | 2023-11-07
107 | 107 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.05 | 111.834 | 4367 | 1861.049 | 0.24 |  | 0.43 |  | bb | 2023-11-07
108 | 108 | 20231106_HepCl_ReRun_108 | Blank | Blank |  |  |  | 0.289 | [5 empty] | 2023-11-07
109 | 109 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 4.05 | 1863.931 | 71024 | 1289.396 | 5.782 | 30.5 | 33.68 | 10.4 | bb | 2023-11-07
110 | 110 | 20231106_HepCl_ReRun_110 | Blank | Blank |  |  |  | 0.254 | [5 empty] | 2023-11-07
111 | 111 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.05 | 1087.664 | 41574 | 2277.866 | 1.91 |  | 10.45 |  | bb | 2023-11-07
112 | 112 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.05 | 109.222 | 3935 | 1961.049 | 0.223 |  | 0.33 |  | bb | 2023-11-07
113 | 113 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.05 | 134.736 | 5715 | 2366.501 | 0.228 |  | 0.36 |  | bb | 2023-11-07
114 | 114 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.05 | 200.381 | 7915 | 2257.612 | 0.355 |  | 1.12 |  | bb | 2023-11-07
115 | 115 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 271.229 | 10772 | 2389.278 | 0.454 |  | 1.72 |  | bb | 2023-11-07
116 | 116 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.05 | 147.119 | 5598 | 2282.505 | 0.258 |  | 0.54 |  | bb | 2023-11-07
117 | 117 | 20231106_HepCl_ReRun_117 | Blank | Blank |  |  |  | 0.204 | [5 empty] | 2023-11-07
118 | 118 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.05 | 167.121 | 6732 | 2296.952 | 0.291 |  | 0.74 |  | bb | 2023-11-07
119 | 119 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.05 | 147.658 | 5726 | 2078.18 | 0.284 |  | 0.7 |  | bb | 2023-11-07
120 | 120 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.05 | 219.107 | 8462 | 2319.462 | 0.378 |  | 1.26 |  | bb | 2023-11-07
121 | 121 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.05 | 691.918 | 26391 | 2365.194 | 1.17 |  | 6.01 |  | bb | 2023-11-07
122 | 122 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.05 | 669.348 | 25473 | 2378.302 | 1.126 |  | 5.75 |  | bb | 2023-11-07
123 | 123 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.05 | 469.471 | 17560 | 2149.042 | 0.874 |  | 4.23 |  | bb | 2023-11-07
124 | 124 | 20231106_HepCl_ReRun_124 | Blank | Blank |  |  |  | 0.326 | [5 empty] | 2023-11-07
125 | 125 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.05 | 152.881 | 6501 | 2177.948 | 0.281 |  | 0.68 |  | bb | 2023-11-07
126 | 126 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.05 | 356.513 | 14002 | 1966.663 | 0.725 |  | 3.34 |  | bb | 2023-11-07
127 | 127 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.05 | 221.256 | 8132 | 2330.105 | 0.38 |  | 1.27 |  | bb | 2023-11-07
128 | 128 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.05 | 118.698 | 4868 | 1788.846 | 0.265 |  | 0.58 |  | bb | 2023-11-07
129 | 129 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.05 | 89.453 | 3749 | 1405.372 | 0.255 |  | 0.52 |  | bb | 2023-11-07
130 | 130 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.05 | 130.351 | 5363 | 1826.592 | 0.285 |  | 0.7 |  | bb | 2023-11-07
131 | 131 | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.05 | 0.165 | 30 | 0.336 | 1.964 |  | 10.78 |  | bb | 2023-11-07
132 | 132 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 4.05 | 1845.258 | 70940 | 1187.432 | 6.216 | 30.5 | 36.28 | 19 | bb | 2023-11-07
133 | 133 | 20231106_HepCl_ReRun_133 | Blank | Blank | [9 empty] | 2023-11-07
134 | 134 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.05 | 207.095 | 7876 | 2059.017 | 0.402 |  | 1.41 |  | bb | 2023-11-07
135 | 135 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.05 | 353.006 | 13268 | 2482.213 | 0.569 |  | 2.41 |  | bb | 2023-11-07
136 | 136 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 232.939 | 8411 | 2293.903 | 0.406 |  | 1.43 |  | bb | 2023-11-07
137 | 137 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.05 | 220.973 | 8680 | 2212.084 | 0.4 |  | 1.39 |  | bb | 2023-11-07
138 | 138 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.05 | 205.306 | 8231 | 2292.081 | 0.358 |  | 1.14 |  | bb | 2023-11-07
139 | 139 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.05 | 193.423 | 7058 | 2233.479 | 0.346 |  | 1.07 |  | bb | 2023-11-07
140 | 140 | 20231106_HepCl_ReRun_140 | Blank | Blank |  |  |  | 0.47 | [5 empty] | 2023-11-07
141 | 141 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 4.05 | 1715.24 | 65787 | 1154.56 | 5.942 | 30.5 | 34.64 | 13.6 | bb | 2023-11-07
142 | 142 | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.02 | 0.165 | 18 | 0.21 | 3.143 |  | 17.85 |  | bb | 2023-11-07
143 | 143 | 20231106_HepCl_ReRun_143 | Blank | Blank | [9 empty] | 2023-11-07
144 | 144 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.05 | 4630.355 | 168087 | 2534.763 | 7.307 |  | 42.83 |  | bb | 2023-11-07
145 | 145 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.05 | 4343.149 | 162971 | 2424.725 | 7.165 |  | 41.98 |  | bb | 2023-11-07
146 | 146 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.05 | 4442.118 | 170191 | 2412.854 | 7.364 |  | 43.17 |  | bb | 2023-11-07
147 | 147 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.05 | 3689.881 | 140619 | 2272.698 | 6.494 |  | 37.95 |  | bb | 2023-11-07
148 | 148 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.05 | 3383.043 | 128491 | 2150.592 | 6.292 |  | 36.74 |  | bb | 2023-11-07
149 | 149 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.05 | 3382.356 | 128897 | 2036.234 | 6.644 |  | 38.85 |  | bb | 2023-11-07
150 | 150 | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.05 | 0.254 | 15 | [5 empty] | bb | 2023-11-07
151 | 151 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.05 | 3519.728 | 133013 | 2160.931 | 6.515 |  | 38.08 |  | bb | 2023-11-07
152 | 152 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.05 | 3685.783 | 137360 | 2170.902 | 6.791 |  | 39.74 |  | bb | 2023-11-07
153 | 153 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.05 | 3722.396 | 140864 | 2075.98 | 7.172 |  | 42.02 |  | bb | 2023-11-07
154 | 154 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.05 | 3417.414 | 128123 | 2070.364 | 6.603 |  | 38.6 |  | bb | 2023-11-07
155 | 155 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.05 | 3877.35 | 146603 | 2153.034 | 7.204 |  | 42.21 |  | bb | 2023-11-07
156 | 156 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.05 | 4665.503 | 182306 | 2243.382 | 8.319 |  | 48.9 |  | bb | 2023-11-07
157 | 157 | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.08 | 0.303 | 18 | [5 empty] | bb | 2023-11-07
158 | 158 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.05 | 3521.64 | 135865 | 2093.272 | 6.729 |  | 39.37 |  | bb | 2023-11-07
159 | 159 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.05 | 3449.08 | 131631 | 2165.184 | 6.372 |  | 37.22 |  | bb | 2023-11-07
160 | 160 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.05 | 3022.028 | 109144 | 2151.335 | 5.619 |  | 32.7 |  | bb | 2023-11-07
161 | 161 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.05 | 2831.834 | 110285 | 1989.004 | 5.695 |  | 33.16 |  | bb | 2023-11-07
162 | 162 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.05 | 2885.134 | 111975 | 1868.877 | 6.175 |  | 36.04 |  | bb | 2023-11-07
163 | 163 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.05 | 2826.412 | 111737 | 1927.043 | 5.867 |  | 34.19 |  | bb | 2023-11-07
164 | 164 | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.05 | 0.883 | 52 | 0.343 | 10.297 |  | 60.77 |  | bb | 2023-11-07
165 | 165 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 4.05 | 1524.464 | 59642 | 1024.203 | 5.954 | 30.5 | 34.71 | 13.8 | bb | 2023-11-07
166 | 166 | 20231106_HepCl_ReRun_168 | Blank | Blank | [9 empty] | 2023-11-07
167 | 167 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.05 | 3666.114 | 142234 | 2335.834 | 6.278 |  | 36.66 |  | bb | 2023-11-07
168 | 168 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.05 | 3464.063 | 129174 | 2227.993 | 6.219 |  | 36.3 |  | bb | 2023-11-07
169 | 169 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.05 | 4187.168 | 160122 | 2176.578 | 7.695 |  | 45.16 |  | bb | 2023-11-07
170 | 170 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.05 | 3326.948 | 127478 | 2189.358 | 6.078 |  | 35.46 |  | bb | 2023-11-07
171 | 171 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.05 | 3317.284 | 126490 | 2193.385 | 6.05 |  | 35.29 |  | bb | 2023-11-07
172 | 172 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.05 | 3435.218 | 132714 | 2207.604 | 6.224 |  | 36.33 |  | bb | 2023-11-07
173 | 173 | 20231106_HepCl_ReRun_175 | Blank | Blank | [9 empty] | 2023-11-07
174 | 174 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.05 | 3227.542 | 125488 | 2186.832 | 5.904 |  | 34.41 |  | bb | 2023-11-07
175 | 175 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.05 | 3479.264 | 131384 | 2072.819 | 6.714 |  | 39.27 |  | bb | 2023-11-07
176 | 176 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.05 | 3257.239 | 117830 | 2269.267 | 5.741 |  | 33.44 |  | bb | 2023-11-07
177 | 177 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.05 | 3758.575 | 140254 | 2124.282 | 7.077 |  | 41.45 |  | bb | 2023-11-07
178 | 178 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.05 | 3681.078 | 139784 | 2198.032 | 6.699 |  | 39.18 |  | bb | 2023-11-07
179 | 179 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.05 | 3579.895 | 132544 | 2234.033 | 6.41 |  | 37.45 |  | bb | 2023-11-07
180 | 180 | 20231106_HepCl_ReRun_182 | Blank | Blank | 4.02 | 0.146 | 9 | 0.546 | 1.07 |  | 5.41 |  | bb | 2023-11-07
181 | 181 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.05 | 3289.579 | 127687 | 2107.887 | 6.242 |  | 36.44 |  | bb | 2023-11-07
182 | 182 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.05 | 2823.59 | 105276 | 2019.512 | 5.593 |  | 32.54 |  | bb | 2023-11-07
183 | 183 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.05 | 3351.329 | 126030 | 2174.987 | 6.163 |  | 35.97 |  | bb | 2023-11-07
184 | 184 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.05 | 2774.165 | 104724 | 1885.64 | 5.885 |  | 34.3 |  | bb | 2023-11-07
185 | 185 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.05 | 2413.183 | 95545 | 1554.347 | 6.21 |  | 36.25 |  | bb | 2023-11-07
186 | 186 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.05 | 2684.16 | 102251 | 1868.361 | 5.747 |  | 33.47 |  | bb | 2023-11-07
187 | 187 | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.09 | 0.168 | 9 | 0.191 | 3.518 |  | 20.1 |  | bb | 2023-11-07
188 | 188 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 4.05 | 1400.518 | 54563 | 996.986 | 5.619 | 30.5 | 32.7 | 7.2 | bb | 2023-11-07
189 | 189 | 20231106_HepCl_ReRun_191 | Blank | Blank | [9 empty] | 2023-11-07
190 | 190 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.05 | 3195.274 | 116858 | 2224.932 | 5.744 |  | 33.46 |  | bb | 2023-11-07
191 | 191 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.05 | 3959.748 | 152075 | 2450.528 | 6.464 |  | 37.77 |  | bb | 2023-11-07
192 | 192 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.05 | 4245.902 | 161078 | 2422.206 | 7.012 |  | 41.06 |  | bb | 2023-11-07
193 | 193 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.05 | 3195.347 | 124660 | 2046.279 | 6.246 |  | 36.47 |  | bb | 2023-11-07
194 | 194 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.05 | 3478.008 | 137993 | 2164.459 | 6.427 |  | 37.55 |  | bb | 2023-11-07
195 | 195 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.05 | 3191.466 | 119308 | 2047.973 | 6.233 |  | 36.39 |  | bb | 2023-11-07
196 | 196 | 20231106_HepCl_ReRun_198 | Blank | Blank | [9 empty] | 2023-11-08
197 | 197 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 4.05 | 1391.144 | 54588 | 1022.509 | 5.442 | 30.5 | 31.64 | 3.7 | bb | 2023-11-08
198 | 198 | 20231106_HepCl_ReRun_200 | Blank | Blank |  |  |  | 0.632 | [5 empty] | 2023-11-08
199 | 199 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.05 | 444.989 | 16010 | 2325.502 | 0.765 |  | 3.58 |  | bb | 2023-11-08
200 | 200 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.05 | 169.976 | 6440 | 2318.157 | 0.293 |  | 0.75 |  | bb | 2023-11-08
201 | 201 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.05 | 77.56 | 3248 | 2446.042 | 0.127 |  |  |  | bbI | 2023-11-08
202 | 202 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.05 | 65.056 | 3165 | 2371.027 | 0.11 |  |  |  | bbI | 2023-11-08
203 | 203 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.05 | 59.29 | 2368 | 2403.863 | 0.099 |  |  |  | bbI | 2023-11-08
204 | 204 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.05 | 51.517 | 2109 | 2416.499 | 0.085 |  |  |  | bbI | 2023-11-08
205 | 205 | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.05 | 0.506 | 30 | 1.208 | 1.675 |  | 9.04 |  | bb | 2023-11-08
206 | 206 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.05 | 134.341 | 5716 | 2279.891 | 0.236 |  | 0.41 |  | bb | 2023-11-08
207 | 207 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.05 | 120.986 | 4204 | 2295.882 | 0.211 |  | 0.26 |  | bb | 2023-11-08
208 | 208 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.05 | 66.306 | 2725 | 2274.982 | 0.117 |  |  |  | bbI | 2023-11-08
209 | 209 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.05 | 85.625 | 3468 | 1978.307 | 0.173 |  | 0.03 |  | bb | 2023-11-08
210 | 210 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.05 | 98.728 | 3906 | 1924.4 | 0.205 |  | 0.22 |  | bb | 2023-11-08
211 | 211 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.05 | 162.785 | 6639 | 1966.173 | 0.331 |  | 0.98 |  | bb | 2023-11-08
212 | 212 | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.04 | 0.323 | 19 | [5 empty] | bb | 2023-11-08
213 | 213 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.05 | 71.583 | 3014 | 2244.717 | 0.128 |  |  |  | bbI | 2023-11-08
214 | 214 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.05 | 45.632 | 2059 | 2253.734 | 0.081 |  |  |  | bbI | 2023-11-08
215 | 215 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.05 | 68.436 | 2773 | 2225.482 | 0.123 |  |  |  | bbI | 2023-11-08
216 | 216 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.05 | 68.165 | 2481 | 1927.321 | 0.141 |  |  |  | bbI | 2023-11-08
217 | 217 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.05 | 68.579 | 2800 | 1956.9 | 0.14 |  |  |  | bbI | 2023-11-08
218 | 218 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.05 | 44.038 | 1395 | 1911.096 | 0.092 |  |  |  | bbI | 2023-11-08
219 | 219 | 20231106_HepCl_ReRun_221 | Blank | Blank |  |  |  | 0.465 | [5 empty] | 2023-11-08
220 | 220 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 4.05 | 1477.327 | 58320 | 1040.056 | 5.682 | 30.5 | 33.08 | 8.5 | bb | 2023-11-08
221 | 221 | 20231106_HepCl_ReRun_223 | Blank | Blank | [9 empty] | 2023-11-08
222 | 222 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.05 | 86.69 | 3498 | 2183.897 | 0.159 |  |  |  | bbI | 2023-11-08
223 | 223 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.05 | 47.479 | 2095 | 2102.246 | 0.09 |  |  |  | bbI | 2023-11-08
224 | 224 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.05 | 49.284 | 2089 | 2034.973 | 0.097 |  |  |  | bbI | 2023-11-08
225 | 225 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.05 | 49.724 | 2144 | 2392.802 | 0.083 |  |  |  | bbI | 2023-11-08
226 | 226 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.05 | 54.339 | 2333 | 2360.479 | 0.092 |  |  |  | bbI | 2023-11-08
227 | 227 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.05 | 71.881 | 2863 | 2468 | 0.117 |  |  |  | bbI | 2023-11-08
228 | 228 | 20231106_HepCl_ReRun_230 | Blank | Blank |  |  |  | 0.182 | [5 empty] | 2023-11-08
229 | 229 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.05 | 64.293 | 2544 | 2278.214 | 0.113 |  |  |  | bbI | 2023-11-08
230 | 230 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.05 | 73.579 | 3336 | 2180.041 | 0.135 |  |  |  | bbI | 2023-11-08
231 | 231 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.05 | 48.467 | 1868 | 2026.516 | 0.096 |  |  |  | bbI | 2023-11-08
232 | 232 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.05 | 4290.784 | 170039 | 1951.082 | 8.797 |  | 51.77 |  | bb | 2023-11-08
233 | 233 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.05 | 187.424 | 7368 | 1931.021 | 0.388 |  | 1.32 |  | bb | 2023-11-08
234 | 234 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.05 | 1550.709 | 60551 | 1544.025 | 4.017 |  | 23.09 |  | bb | 2023-11-08
235 | 235 | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 0.132 | [5 empty] | 2023-11-08
236 | 236 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.05 | 63.795 | 2704 | 2285.38 | 0.112 |  |  |  | bbI | 2023-11-08
237 | 237 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.05 | 58.385 | 2570 | 2275.792 | 0.103 |  |  |  | bbI | 2023-11-08
238 | 238 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.05 | 60.019 | 2065 | 2367.659 | 0.101 |  |  |  | bbI | 2023-11-08
239 | 239 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.05 | 43.619 | 1630 | 2073.876 | 0.084 |  |  |  | bbI | 2023-11-08
240 | 240 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.05 | 49.147 | 2150 | 1820.366 | 0.108 |  |  |  | bbI | 2023-11-08
241 | 241 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.05 | 49.617 | 1955 | 1823.648 | 0.109 |  |  |  | bbI | 2023-11-08
242 | 242 | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 0.507 | [5 empty] | 2023-11-08
243 | 243 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 4.05 | 1451.808 | 56361 | 952.141 | 6.099 | 30.5 | 35.58 | 16.7 | bb | 2023-11-08
244 | 244 | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.03 | 0.616 | 36 | 0.149 | 16.537 |  | 98.2 |  | bb | 2023-11-08
245 | 245 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.05 | 189.058 | 7842 | 2702.162 | 0.28 |  | 0.67 |  | bb | 2023-11-08
246 | 246 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.05 | 105.691 | 4814 | 2421.035 | 0.175 |  | 0.04 |  | bb | 2023-11-08
247 | 247 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.05 | 149.801 | 6328 | 2636.327 | 0.227 |  | 0.36 |  | bb | 2023-11-08
248 | 248 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.05 | 311.631 | 12419 | 2181.149 | 0.571 |  | 2.42 |  | bb | 2023-11-08
249 | 249 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.05 | 312.361 | 11999 | 2069.581 | 0.604 |  | 2.61 |  | bb | 2023-11-08
250 | 250 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.05 | 359.079 | 14109 | 2232.48 | 0.643 |  | 2.85 |  | bb | 2023-11-08
251 | 251 | 20231106_HepCl_ReRun_253 | Blank | Blank |  |  |  | 0.066 | [5 empty] | 2023-11-08
252 | 252 | 20231106_HepCl_ReRun_254 | Blank | Blank | [9 empty] | 2023-11-08
253 | 253 | 20231106_HepCl_ReRun_255 | Blank | Blank |  |  |  | 0.647 | [5 empty] | 2023-11-08
Compound 7: M2-8:2 FTS
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.07 | 2.685 | 157 |  | 2.685 | 4 | 0.01 | -99.8 | bb | 2023-11-06
2 | 2 | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.04 | 3.534 | 118 |  | 3.534 | 4 | 0.01 | -99.7 | bb | 2023-11-06
3 | 3 | 20231106_HepCl_ReRun_003 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
4 | 4 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.05 | 0.86 | 50 |  | 0.86 | 4 | 0 | -99.9 | bb | 2023-11-06
5 | 5 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.06 | 0.572 | 55 |  | 0.572 | 4 | 0 | -100 | bb | 2023-11-06
6 | 6 | 20231106_HepCl_ReRun_006 | Blank | Blank | 4.04 | 0.473 | 28 |  | 0.473 | 4 | 0 | -100 | bb | 2023-11-06
7 | 7 | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.09 | 1.479 | 80 |  | 1.479 | 4 | 0 | -99.9 | bb | 2023-11-06
8 | 8 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 4.05 | 1404.893 | 51799 |  | 1404.893 | 4 | 4.18 | 4.5 | bb | 2023-11-06
9 | 9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 4.05 | 1292.746 | 46823 |  | 1292.746 | 4 | 3.85 | -3.9 | bb | 2023-11-06
10 | 10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 4.05 | 1261.171 | 45058 |  | 1261.171 | 4 | 3.75 | -6.2 | bb | 2023-11-06
11 | 11 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 4.05 | 1380.42 | 51580 |  | 1380.42 | 4 | 4.11 | 2.7 | bb | 2023-11-06
12 | 12 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 4.05 | 1318.556 | 49054 |  | 1318.556 | 4 | 3.92 | -1.9 | bb | 2023-11-06
13 | 13 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.05 | 1285.358 | 48475 |  | 1285.358 | 4 | 3.82 | -4.4 | bb | 2023-11-06
14 | 14 | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.01 | 0.151 | 15 |  | 0.151 | 4 | 0 | -100 | bb | 2023-11-06
15 | 15 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.05 | 1359.29 | 50339 |  | 1359.29 | 4 | 4.04 | 1.1 | bb | 2023-11-06
16 | 16 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.05 | 1397.361 | 52204 |  | 1397.361 | 4 | 4.16 | 3.9 | bb | 2023-11-06
17 | 17 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.05 | 1364.509 | 50516 |  | 1364.509 | 4 | 4.06 | 1.5 | bb | 2023-11-06
18 | 18 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.05 | 1367.205 | 51313 |  | 1367.205 | 4 | 4.07 | 1.7 | bb | 2023-11-06
19 | 19 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.05 | 1491.731 | 55680 |  | 1491.731 | 4 | 4.44 | 10.9 | bb | 2023-11-06
20 | 20 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.05 | 1313.037 | 49232 |  | 1313.037 | 4 | 3.91 | -2.4 | bb | 2023-11-06
21 | 21 | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.01 | 3.142 | 118 |  | 3.142 | 4 | 0.01 | -99.8 | bb | 2023-11-06
22 | 22 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.05 | 1391.243 | 51807 |  | 1391.243 | 4 | 4.14 | 3.5 | bb | 2023-11-06
23 | 23 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.05 | 1497.735 | 54749 |  | 1497.735 | 4 | 4.46 | 11.4 | bb | 2023-11-06
24 | 24 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.05 | 1422.099 | 52406 |  | 1422.099 | 4 | 4.23 | 5.8 | bb | 2023-11-06
25 | 25 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 4.05 | 1729.245 | 60671 |  | 1729.245 | 4 | 5.14 | 28.6 | bb | 2023-11-06
26 | 26 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 4.05 | 1482.366 | 54721 |  | 1482.366 | 4 | 4.41 | 10.2 | bb | 2023-11-06
27 | 27 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.07 | 0.598 | 35 |  | 0.598 | 4 | 0 | -100 | bb | 2023-11-06
28 | 28 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
29 | 29 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 4.05 | 1378.086 | 49348 |  | 1378.086 | 4 | 4.1 | 2.5 | bb | 2023-11-06
30 | 30 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 4.05 | 1331.639 | 48209 |  | 1331.639 | 4 | 3.96 | -1 | bb | 2023-11-06
31 | 31 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.05 | 1266.854 | 47688 |  | 1266.854 | 4 | 3.77 | -5.8 | bb | 2023-11-06
32 | 32 | 20231106_HepCl_ReRun_032 | Blank | Blank | 4.01 | 0.372 | 22 |  | 0.372 | 4 | 0 | -100 | bb | 2023-11-06
33 | 33 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 4.05 | 1402.833 | 55016 |  | 1402.833 | 4 | 4.17 | 4.3 | bb | 2023-11-06
34 | 34 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 4.05 | 1393.114 | 51826 |  | 1393.114 | 4 | 4.14 | 3.6 | bb | 2023-11-06
35 | 35 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.05 | 1260.427 | 46864 |  | 1260.427 | 4 | 3.75 | -6.3 | bb | 2023-11-06
36 | 36 | 20231106_HepCl_ReRun_036 | Blank | Blank | 4.07 | 0.326 | 19 |  | 0.326 | 4 | 0 | -100 | bb | 2023-11-06
37 | 37 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 4.05 | 1405.047 | 55126 |  | 1405.047 | 4 | 4.18 | 4.5 | bb | 2023-11-06
38 | 38 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 4.05 | 1291.964 | 49703 |  | 1291.964 | 4 | 3.84 | -3.9 | bb | 2023-11-06
39 | 39 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.05 | 1310.875 | 49191 |  | 1310.875 | 4 | 3.9 | -2.5 | bb | 2023-11-06
40 | 40 | 20231106_HepCl_ReRun_040 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
41 | 41 | 20231106_HepCl_ReRun_041 | QC A | QC | 4.05 | 1704.962 | 62062 |  | 1704.962 | 4 | 5.07 | 26.8 | bb | 2023-11-06
42 | 42 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.05 | 1697.607 | 63218 |  | 1697.607 | 4 | 5.05 | 26.2 | bb | 2023-11-06
43 | 43 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.05 | 1749.339 | 68129 |  | 1749.339 | 4 | 5.2 | 30.1 | bb | 2023-11-06
44 | 44 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.05 | 1516.869 | 56991 |  | 1516.869 | 4 | 4.51 | 12.8 | bb | 2023-11-06
45 | 45 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 4.1 | 0.278 | 12 |  | 0.278 | 4 | 0 | -100 | bb | 2023-11-06
46 | 46 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.04 | 0.087 | 5 |  | 0.087 | 4 | 0 | -100 | bb | 2023-11-06
47 | 47 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 4.05 | 1396.361 | 52081 |  | 1396.361 | 4 | 4.15 | 3.8 | bb | 2023-11-06
48 | 48 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 4.05 | 1232.084 | 47762 |  | 1232.084 | 4 | 3.66 | -8.4 | bb | 2023-11-06
49 | 49 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 4.05 | 1101.421 | 39699 |  | 1101.421 | 4 | 3.28 | -18.1 | bb | 2023-11-06
50 | 50 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 4.05 | 1415.325 | 53894 |  | 1415.325 | 4 | 4.21 | 5.2 | bb | 2023-11-06
51 | 51 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 4.05 | 1218 | 45808 |  | 1218 | 4 | 3.62 | -9.4 | bb | 2023-11-06
52 | 52 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.05 | 1179.859 | 41926 |  | 1179.859 | 4 | 3.51 | -12.3 | bb | 2023-11-06
53 | 53 | 20231106_HepCl_ReRun_053 | Blank | Blank | 4.09 | 0.376 | 22 |  | 0.376 | 4 | 0 | -100 | bb | 2023-11-06
54 | 54 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.05 | 1307.154 | 48444 |  | 1307.154 | 4 | 3.89 | -2.8 | bb | 2023-11-06
55 | 55 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.05 | 1383.121 | 50458 |  | 1383.121 | 4 | 4.11 | 2.9 | bb | 2023-11-06
56 | 56 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.05 | 1332.81 | 48860 |  | 1332.81 | 4 | 3.96 | -0.9 | bb | 2023-11-06
57 | 57 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.05 | 1340.69 | 49527 |  | 1340.69 | 4 | 3.99 | -0.3 | bb | 2023-11-06
58 | 58 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.05 | 1394.592 | 50530 |  | 1394.592 | 4 | 4.15 | 3.7 | bb | 2023-11-06
59 | 59 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.05 | 1381.883 | 52583 |  | 1381.883 | 4 | 4.11 | 2.8 | bb | 2023-11-06
60 | 60 | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.05 | 0.214 | 13 |  | 0.214 | 4 | 0 | -100 | bb | 2023-11-06
61 | 61 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.05 | 1344.772 | 48392 |  | 1344.772 | 4 | 4 | 0 | bb | 2023-11-06
62 | 62 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.05 | 1523.123 | 57253 |  | 1523.123 | 4 | 4.53 | 13.3 | bb | 2023-11-06
63 | 63 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.05 | 1427.197 | 52880 |  | 1427.197 | 4 | 4.25 | 6.1 | bb | 2023-11-06
64 | 64 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 4.05 | 1844.922 | 67300 |  | 1844.922 | 4 | 5.49 | 37.2 | bb | 2023-11-06
65 | 65 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 4.05 | 1484.605 | 50568 |  | 1484.605 | 4 | 4.42 | 10.4 | bb | 2023-11-07
66 | 66 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4.02 | 3.149 | 105 |  | 3.149 | 4 | 0.01 | -99.8 | bb | 2023-11-07
67 | 67 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.1 | 0.348 | 15 |  | 0.348 | 4 | 0 | -100 | bb | 2023-11-07
68 | 68 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 4.05 | 1319.151 | 49822 |  | 1319.151 | 4 | 3.92 | -1.9 | bb | 2023-11-07
69 | 69 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 4.05 | 1336.031 | 51239 |  | 1336.031 | 4 | 3.97 | -0.6 | bb | 2023-11-07
70 | 70 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.05 | 1220.534 | 47238 |  | 1220.534 | 4 | 3.63 | -9.2 | bb | 2023-11-07
71 | 71 | 20231106_HepCl_ReRun_071 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
72 | 72 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 4.05 | 1228.853 | 46834 |  | 1228.853 | 4 | 3.66 | -8.6 | bb | 2023-11-07
73 | 73 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 4.05 | 1286.003 | 49431 |  | 1286.003 | 4 | 3.83 | -4.4 | bb | 2023-11-07
74 | 74 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.05 | 1181.049 | 43690 |  | 1181.049 | 4 | 3.51 | -12.2 | bb | 2023-11-07
75 | 75 | 20231106_HepCl_ReRun_075 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
76 | 76 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 4.05 | 1246.689 | 48897 |  | 1246.689 | 4 | 3.71 | -7.3 | bb | 2023-11-07
77 | 77 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 4.05 | 1098.067 | 41974 |  | 1098.067 | 4 | 3.27 | -18.3 | bb | 2023-11-07
78 | 78 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.05 | 1218.418 | 46861 |  | 1218.418 | 4 | 3.62 | -9.4 | bb | 2023-11-07
79 | 79 | 20231106_HepCl_ReRun_079 | Blank | Blank | 4.05 | 0.232 | 36 |  | 0.232 | 4 | 0 | -100 | bb | 2023-11-07
80 | 80 | 20231106_HepCl_ReRun_080 | QC A | QC | 4.05 | 1712.896 | 62704 |  | 1712.896 | 4 | 5.1 | 27.4 | bb | 2023-11-07
81 | 81 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.05 | 1702.264 | 67418 |  | 1702.264 | 4 | 5.06 | 26.6 | bb | 2023-11-07
82 | 82 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.05 | 1712.841 | 65873 |  | 1712.841 | 4 | 5.09 | 27.4 | bb | 2023-11-07
83 | 83 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.05 | 1490.915 | 57447 |  | 1490.915 | 4 | 4.43 | 10.9 | bb | 2023-11-07
84 | 84 | 20231106_HepCl_ReRun_084 | Blank | Blank | 4.05 | 0.156 | 9 |  | 0.156 | 4 | 0 | -100 | bb | 2023-11-07
85 | 85 | 20231106_HepCl_ReRun_085 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
86 | 86 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 4.05 | 1311.876 | 49396 |  | 1311.876 | 4 | 3.9 | -2.4 | bb | 2023-11-07
87 | 87 | 20231106_HepCl_ReRun_087 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
88 | 88 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.05 | 2610.134 | 101241 |  | 2610.134 | 4 | 7.76 | 94.1 | bb | 2023-11-07
89 | 89 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.05 | 2373.515 | 89417 |  | 2373.515 | 4 | 7.06 | 76.5 | bb | 2023-11-07
90 | 90 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.05 | 2397.312 | 91363 |  | 2397.312 | 4 | 7.13 | 78.3 | bb | 2023-11-07
91 | 91 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.05 | 2359.211 | 88771 |  | 2359.211 | 4 | 7.02 | 75.4 | bb | 2023-11-07
92 | 92 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.05 | 2233.017 | 83972 |  | 2233.017 | 4 | 6.64 | 66.1 | bb | 2023-11-07
93 | 93 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 2234.758 | 81479 |  | 2234.758 | 4 | 6.65 | 66.2 | bb | 2023-11-07
94 | 94 | 20231106_HepCl_ReRun_094 | Blank | Blank | 4.1 | 1.338 | 61 |  | 1.338 | 4 | 0 | -99.9 | bb | 2023-11-07
95 | 95 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.05 | 2519.11 | 93264 |  | 2519.11 | 4 | 7.49 | 87.3 | bb | 2023-11-07
96 | 96 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.05 | 2439.5 | 90275 |  | 2439.5 | 4 | 7.26 | 81.4 | bb | 2023-11-07
97 | 97 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.05 | 2235.031 | 83647 |  | 2235.031 | 4 | 6.65 | 66.2 | bb | 2023-11-07
98 | 98 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.05 | 2343.231 | 87432 |  | 2343.231 | 4 | 6.97 | 74.3 | bb | 2023-11-07
99 | 99 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.05 | 2400.613 | 90344 |  | 2400.613 | 4 | 7.14 | 78.5 | bb | 2023-11-07
100 | 100 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.05 | 2513.042 | 92824 |  | 2513.042 | 4 | 7.48 | 86.9 | bb | 2023-11-07
101 | 101 | 20231106_HepCl_ReRun_101 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
102 | 102 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.05 | 2126.96 | 81491 |  | 2126.96 | 4 | 6.33 | 58.2 | bb | 2023-11-07
103 | 103 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.05 | 2142.216 | 78827 |  | 2142.216 | 4 | 6.37 | 59.3 | bb | 2023-11-07
104 | 104 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.05 | 2372.858 | 87556 |  | 2372.858 | 4 | 7.06 | 76.5 | bb | 2023-11-07
105 | 105 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.05 | 1960.445 | 72293 |  | 1960.445 | 4 | 5.83 | 45.8 | bb | 2023-11-07
106 | 106 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.05 | 1928.79 | 72627 |  | 1928.79 | 4 | 5.74 | 43.4 | bb | 2023-11-07
107 | 107 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.05 | 1861.049 | 71569 |  | 1861.049 | 4 | 5.54 | 38.4 | bb | 2023-11-07
108 | 108 | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.1 | 0.289 | 13 |  | 0.289 | 4 | 0 | -100 | bb | 2023-11-07
109 | 109 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 4.05 | 1289.396 | 48638 |  | 1289.396 | 4 | 3.84 | -4.1 | bb | 2023-11-07
110 | 110 | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.04 | 0.254 | 45 |  | 0.254 | 4 | 0 | -100 | bb | 2023-11-07
111 | 111 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.05 | 2277.866 | 87482 |  | 2277.866 | 4 | 6.78 | 69.4 | bb | 2023-11-07
112 | 112 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.05 | 1961.049 | 74969 |  | 1961.049 | 4 | 5.83 | 45.8 | bb | 2023-11-07
113 | 113 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.05 | 2366.501 | 90032 |  | 2366.501 | 4 | 7.04 | 76 | bb | 2023-11-07
114 | 114 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.05 | 2257.612 | 86402 |  | 2257.612 | 4 | 6.72 | 67.9 | bb | 2023-11-07
115 | 115 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 2389.278 | 89739 |  | 2389.278 | 4 | 7.11 | 77.7 | bb | 2023-11-07
116 | 116 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.05 | 2282.505 | 85777 |  | 2282.505 | 4 | 6.79 | 69.7 | bb | 2023-11-07
117 | 117 | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.06 | 0.204 | 16 |  | 0.204 | 4 | 0 | -100 | bb | 2023-11-07
118 | 118 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.05 | 2296.952 | 84946 |  | 2296.952 | 4 | 6.83 | 70.8 | bb | 2023-11-07
119 | 119 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.05 | 2078.18 | 78341 |  | 2078.18 | 4 | 6.18 | 54.5 | bb | 2023-11-07
120 | 120 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.05 | 2319.462 | 87625 |  | 2319.462 | 4 | 6.9 | 72.5 | bb | 2023-11-07
121 | 121 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.05 | 2365.194 | 86293 |  | 2365.194 | 4 | 7.04 | 75.9 | bb | 2023-11-07
122 | 122 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.05 | 2378.302 | 91198 |  | 2378.302 | 4 | 7.07 | 76.9 | bb | 2023-11-07
123 | 123 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.05 | 2149.042 | 80393 |  | 2149.042 | 4 | 6.39 | 59.8 | bb | 2023-11-07
124 | 124 | 20231106_HepCl_ReRun_124 | Blank | Blank | 4.09 | 0.326 | 19 |  | 0.326 | 4 | 0 | -100 | bb | 2023-11-07
125 | 125 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.05 | 2177.948 | 82550 |  | 2177.948 | 4 | 6.48 | 62 | bb | 2023-11-07
126 | 126 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.05 | 1966.663 | 74380 |  | 1966.663 | 4 | 5.85 | 46.2 | bb | 2023-11-07
127 | 127 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.05 | 2330.105 | 87302 |  | 2330.105 | 4 | 6.93 | 73.3 | bb | 2023-11-07
128 | 128 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.05 | 1788.846 | 68641 |  | 1788.846 | 4 | 5.32 | 33 | bb | 2023-11-07
129 | 129 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.05 | 1405.372 | 53963 |  | 1405.372 | 4 | 4.18 | 4.5 | bb | 2023-11-07
130 | 130 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.05 | 1826.592 | 70966 |  | 1826.592 | 4 | 5.43 | 35.8 | bb | 2023-11-07
131 | 131 | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.09 | 0.336 | 18 |  | 0.336 | 4 | 0 | -100 | bb | 2023-11-07
132 | 132 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 4.05 | 1187.432 | 45510 |  | 1187.432 | 4 | 3.53 | -11.7 | bb | 2023-11-07
133 | 133 | 20231106_HepCl_ReRun_133 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
134 | 134 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.05 | 2059.017 | 77275 |  | 2059.017 | 4 | 6.12 | 53.1 | bb | 2023-11-07
135 | 135 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.05 | 2482.213 | 93606 |  | 2482.213 | 4 | 7.38 | 84.6 | bb | 2023-11-07
136 | 136 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 2293.903 | 85682 |  | 2293.903 | 4 | 6.82 | 70.6 | bb | 2023-11-07
137 | 137 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.05 | 2212.084 | 83409 |  | 2212.084 | 4 | 6.58 | 64.5 | bb | 2023-11-07
138 | 138 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.05 | 2292.081 | 84135 |  | 2292.081 | 4 | 6.82 | 70.4 | bb | 2023-11-07
139 | 139 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.05 | 2233.479 | 87834 |  | 2233.479 | 4 | 6.64 | 66.1 | bb | 2023-11-07
140 | 140 | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.06 | 0.47 | 28 |  | 0.47 | 4 | 0 | -100 | bb | 2023-11-07
141 | 141 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 4.05 | 1154.56 | 44383 |  | 1154.56 | 4 | 3.43 | -14.1 | bb | 2023-11-07
142 | 142 | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.09 | 0.21 | 11 |  | 0.21 | 4 | 0 | -100 | bb | 2023-11-07
143 | 143 | 20231106_HepCl_ReRun_143 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
144 | 144 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.05 | 2534.763 | 92974 |  | 2534.763 | 4 | 7.54 | 88.5 | bb | 2023-11-07
145 | 145 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.05 | 2424.725 | 87860 |  | 2424.725 | 4 | 7.21 | 80.3 | bb | 2023-11-07
146 | 146 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.05 | 2412.854 | 89939 |  | 2412.854 | 4 | 7.18 | 79.4 | bb | 2023-11-07
147 | 147 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.05 | 2272.698 | 84725 |  | 2272.698 | 4 | 6.76 | 69 | bb | 2023-11-07
148 | 148 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.05 | 2150.592 | 81038 |  | 2150.592 | 4 | 6.4 | 59.9 | bb | 2023-11-07
149 | 149 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.05 | 2036.234 | 75303 |  | 2036.234 | 4 | 6.06 | 51.4 | bb | 2023-11-07
150 | 150 | 20231106_HepCl_ReRun_152 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
151 | 151 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.05 | 2160.931 | 80933 |  | 2160.931 | 4 | 6.43 | 60.7 | bb | 2023-11-07
152 | 152 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.05 | 2170.902 | 80815 |  | 2170.902 | 4 | 6.46 | 61.4 | bb | 2023-11-07
153 | 153 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.05 | 2075.98 | 76015 |  | 2075.98 | 4 | 6.18 | 54.4 | bb | 2023-11-07
154 | 154 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.05 | 2070.364 | 77569 |  | 2070.364 | 4 | 6.16 | 54 | bb | 2023-11-07
155 | 155 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.05 | 2153.034 | 81167 |  | 2153.034 | 4 | 6.4 | 60.1 | bb | 2023-11-07
156 | 156 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.05 | 2243.382 | 84356 |  | 2243.382 | 4 | 6.67 | 66.8 | bb | 2023-11-07
157 | 157 | 20231106_HepCl_ReRun_159 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
158 | 158 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.05 | 2093.272 | 76894 |  | 2093.272 | 4 | 6.23 | 55.7 | bb | 2023-11-07
159 | 159 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.05 | 2165.184 | 82844 |  | 2165.184 | 4 | 6.44 | 61 | bb | 2023-11-07
160 | 160 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.05 | 2151.335 | 79762 |  | 2151.335 | 4 | 6.4 | 60 | bb | 2023-11-07
161 | 161 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.05 | 1989.004 | 74074 |  | 1989.004 | 4 | 5.92 | 47.9 | bb | 2023-11-07
162 | 162 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.05 | 1868.877 | 72443 |  | 1868.877 | 4 | 5.56 | 39 | bb | 2023-11-07
163 | 163 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.05 | 1927.043 | 72505 |  | 1927.043 | 4 | 5.73 | 43.3 | bb | 2023-11-07
164 | 164 | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.06 | 0.343 | 20 |  | 0.343 | 4 | 0 | -100 | bb | 2023-11-07
165 | 165 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 4.05 | 1024.203 | 38095 |  | 1024.203 | 4 | 3.05 | -23.8 | bb | 2023-11-07
166 | 166 | 20231106_HepCl_ReRun_168 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
167 | 167 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.05 | 2335.834 | 89409 |  | 2335.834 | 4 | 6.95 | 73.7 | bb | 2023-11-07
168 | 168 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.05 | 2227.993 | 82124 |  | 2227.993 | 4 | 6.63 | 65.7 | bb | 2023-11-07
169 | 169 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.05 | 2176.578 | 83746 |  | 2176.578 | 4 | 6.47 | 61.9 | bb | 2023-11-07
170 | 170 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.05 | 2189.358 | 78834 |  | 2189.358 | 4 | 6.51 | 62.8 | bb | 2023-11-07
171 | 171 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.05 | 2193.385 | 83567 |  | 2193.385 | 4 | 6.52 | 63.1 | bb | 2023-11-07
172 | 172 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.05 | 2207.604 | 80918 |  | 2207.604 | 4 | 6.57 | 64.2 | bb | 2023-11-07
173 | 173 | 20231106_HepCl_ReRun_175 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
174 | 174 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.05 | 2186.832 | 81592 |  | 2186.832 | 4 | 6.5 | 62.6 | bb | 2023-11-07
175 | 175 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.05 | 2072.819 | 77608 |  | 2072.819 | 4 | 6.17 | 54.1 | bb | 2023-11-07
176 | 176 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.05 | 2269.267 | 85774 |  | 2269.267 | 4 | 6.75 | 68.8 | bb | 2023-11-07
177 | 177 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.05 | 2124.282 | 80191 |  | 2124.282 | 4 | 6.32 | 58 | bb | 2023-11-07
178 | 178 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.05 | 2198.032 | 83709 |  | 2198.032 | 4 | 6.54 | 63.5 | bb | 2023-11-07
179 | 179 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.05 | 2234.033 | 83528 |  | 2234.033 | 4 | 6.65 | 66.1 | bb | 2023-11-07
180 | 180 | 20231106_HepCl_ReRun_182 | Blank | Blank | 4.09 | 0.546 | 33 |  | 0.546 | 4 | 0 | -100 | bb | 2023-11-07
181 | 181 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.05 | 2107.887 | 80161 |  | 2107.887 | 4 | 6.27 | 56.8 | bb | 2023-11-07
182 | 182 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.05 | 2019.512 | 74692 |  | 2019.512 | 4 | 6.01 | 50.2 | bb | 2023-11-07
183 | 183 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.05 | 2174.987 | 82892 |  | 2174.987 | 4 | 6.47 | 61.7 | bb | 2023-11-07
184 | 184 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.05 | 1885.64 | 69423 |  | 1885.64 | 4 | 5.61 | 40.2 | bb | 2023-11-07
185 | 185 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.05 | 1554.347 | 59618 |  | 1554.347 | 4 | 4.62 | 15.6 | bb | 2023-11-07
186 | 186 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.05 | 1868.361 | 70814 |  | 1868.361 | 4 | 5.56 | 38.9 | bb | 2023-11-07
187 | 187 | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.06 | 0.191 | 11 |  | 0.191 | 4 | 0 | -100 | bb | 2023-11-07
188 | 188 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 4.05 | 996.986 | 38354 |  | 996.986 | 4 | 2.97 | -25.9 | bb | 2023-11-07
189 | 189 | 20231106_HepCl_ReRun_191 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
190 | 190 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.05 | 2224.932 | 81925 |  | 2224.932 | 4 | 6.62 | 65.5 | bb | 2023-11-07
191 | 191 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.05 | 2450.528 | 94042 |  | 2450.528 | 4 | 7.29 | 82.2 | bb | 2023-11-07
192 | 192 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.05 | 2422.206 | 92834 |  | 2422.206 | 4 | 7.21 | 80.1 | bb | 2023-11-07
193 | 193 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.05 | 2046.279 | 77503 |  | 2046.279 | 4 | 6.09 | 52.2 | bb | 2023-11-07
194 | 194 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.05 | 2164.459 | 83044 |  | 2164.459 | 4 | 6.44 | 61 | bb | 2023-11-07
195 | 195 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.05 | 2047.973 | 77307 |  | 2047.973 | 4 | 6.09 | 52.3 | bb | 2023-11-07
196 | 196 | 20231106_HepCl_ReRun_198 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
197 | 197 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 4.05 | 1022.509 | 39451 |  | 1022.509 | 4 | 3.04 | -24 | bb | 2023-11-08
198 | 198 | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.04 | 0.632 | 37 |  | 0.632 | 4 | 0 | -100 | bb | 2023-11-08
199 | 199 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.05 | 2325.502 | 83837 |  | 2325.502 | 4 | 6.92 | 72.9 | bb | 2023-11-08
200 | 200 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.05 | 2318.157 | 89454 |  | 2318.157 | 4 | 6.9 | 72.4 | bb | 2023-11-08
201 | 201 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.05 | 2446.042 | 96028 |  | 2446.042 | 4 | 7.28 | 81.9 | bb | 2023-11-08
202 | 202 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.05 | 2371.027 | 93883 |  | 2371.027 | 4 | 7.05 | 76.3 | bb | 2023-11-08
203 | 203 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.05 | 2403.863 | 91589 |  | 2403.863 | 4 | 7.15 | 78.8 | bb | 2023-11-08
204 | 204 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.05 | 2416.499 | 92311 |  | 2416.499 | 4 | 7.19 | 79.7 | bb | 2023-11-08
205 | 205 | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.08 | 1.208 | 32 |  | 1.208 | 4 | 0 | -99.9 | bb | 2023-11-08
206 | 206 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.05 | 2279.891 | 87427 |  | 2279.891 | 4 | 6.78 | 69.5 | bb | 2023-11-08
207 | 207 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.05 | 2295.882 | 91146 |  | 2295.882 | 4 | 6.83 | 70.7 | bb | 2023-11-08
208 | 208 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.05 | 2274.982 | 87382 |  | 2274.982 | 4 | 6.77 | 69.2 | bb | 2023-11-08
209 | 209 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.05 | 1978.307 | 76393 |  | 1978.307 | 4 | 5.88 | 47.1 | bb | 2023-11-08
210 | 210 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.05 | 1924.4 | 73973 |  | 1924.4 | 4 | 5.72 | 43.1 | bb | 2023-11-08
211 | 211 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.05 | 1966.173 | 76220 |  | 1966.173 | 4 | 5.85 | 46.2 | bb | 2023-11-08
212 | 212 | 20231106_HepCl_ReRun_214 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
213 | 213 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.05 | 2244.717 | 87807 |  | 2244.717 | 4 | 6.68 | 66.9 | bb | 2023-11-08
214 | 214 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.05 | 2253.734 | 90122 |  | 2253.734 | 4 | 6.7 | 67.6 | bb | 2023-11-08
215 | 215 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.05 | 2225.482 | 85507 |  | 2225.482 | 4 | 6.62 | 65.5 | bb | 2023-11-08
216 | 216 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.05 | 1927.321 | 76090 |  | 1927.321 | 4 | 5.73 | 43.3 | bb | 2023-11-08
217 | 217 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.05 | 1956.9 | 75761 |  | 1956.9 | 4 | 5.82 | 45.5 | bb | 2023-11-08
218 | 218 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.05 | 1911.096 | 76727 |  | 1911.096 | 4 | 5.68 | 42.1 | bb | 2023-11-08
219 | 219 | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.09 | 0.465 | 26 |  | 0.465 | 4 | 0 | -100 | bb | 2023-11-08
220 | 220 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 4.05 | 1040.056 | 41836 |  | 1040.056 | 4 | 3.09 | -22.7 | bb | 2023-11-08
221 | 221 | 20231106_HepCl_ReRun_223 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
222 | 222 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.05 | 2183.897 | 87172 |  | 2183.897 | 4 | 6.5 | 62.4 | bb | 2023-11-08
223 | 223 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.05 | 2102.246 | 79293 |  | 2102.246 | 4 | 6.25 | 56.3 | bb | 2023-11-08
224 | 224 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.05 | 2034.973 | 79528 |  | 2034.973 | 4 | 6.05 | 51.3 | bb | 2023-11-08
225 | 225 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.05 | 2392.802 | 92804 |  | 2392.802 | 4 | 7.12 | 77.9 | bb | 2023-11-08
226 | 226 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.05 | 2360.479 | 94249 |  | 2360.479 | 4 | 7.02 | 75.5 | bb | 2023-11-08
227 | 227 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.05 | 2468 | 99081 |  | 2468 | 4 | 7.34 | 83.5 | bb | 2023-11-08
228 | 228 | 20231106_HepCl_ReRun_230 | Blank | Blank | 4.04 | 0.182 | 11 |  | 0.182 | 4 | 0 | -100 | bb | 2023-11-08
229 | 229 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.05 | 2278.214 | 91468 |  | 2278.214 | 4 | 6.78 | 69.4 | bb | 2023-11-08
230 | 230 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.05 | 2180.041 | 83876 |  | 2180.041 | 4 | 6.48 | 62.1 | bb | 2023-11-08
231 | 231 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.05 | 2026.516 | 78913 |  | 2026.516 | 4 | 6.03 | 50.7 | bb | 2023-11-08
232 | 232 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.05 | 1951.082 | 75634 |  | 1951.082 | 4 | 5.8 | 45.1 | bb | 2023-11-08
233 | 233 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.05 | 1931.021 | 76802 |  | 1931.021 | 4 | 5.74 | 43.6 | bb | 2023-11-08
234 | 234 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.05 | 1544.025 | 60954 |  | 1544.025 | 4 | 4.59 | 14.8 | bb | 2023-11-08
235 | 235 | 20231106_HepCl_ReRun_237 | Blank | Blank | 4.04 | 0.132 | 21 |  | 0.132 | 4 | 0 | -100 | bb | 2023-11-08
236 | 236 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.05 | 2285.38 | 88611 |  | 2285.38 | 4 | 6.8 | 70 | bb | 2023-11-08
237 | 237 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.05 | 2275.792 | 87791 |  | 2275.792 | 4 | 6.77 | 69.2 | bb | 2023-11-08
238 | 238 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.05 | 2367.659 | 92560 |  | 2367.659 | 4 | 7.04 | 76.1 | bb | 2023-11-08
239 | 239 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.05 | 2073.876 | 84779 |  | 2073.876 | 4 | 6.17 | 54.2 | bb | 2023-11-08
240 | 240 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.05 | 1820.366 | 71864 |  | 1820.366 | 4 | 5.41 | 35.4 | bb | 2023-11-08
241 | 241 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.05 | 1823.648 | 69186 |  | 1823.648 | 4 | 5.42 | 35.6 | bb | 2023-11-08
242 | 242 | 20231106_HepCl_ReRun_244 | Blank | Blank | 4.09 | 0.507 | 30 |  | 0.507 | 4 | 0 | -100 | bb | 2023-11-08
243 | 243 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 4.05 | 952.141 | 36160 |  | 952.141 | 4 | 2.83 | -29.2 | bb | 2023-11-08
244 | 244 | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.04 | 0.149 | 9 |  | 0.149 | 4 | 0 | -100 | bb | 2023-11-08
245 | 245 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.05 | 2702.162 | 105538 |  | 2702.162 | 4 | 8.04 | 100.9 | bb | 2023-11-08
246 | 246 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.05 | 2421.035 | 92806 |  | 2421.035 | 4 | 7.2 | 80 | bb | 2023-11-08
247 | 247 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.05 | 2636.327 | 101850 |  | 2636.327 | 4 | 7.84 | 96 | bb | 2023-11-08
248 | 248 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.05 | 2181.149 | 82489 |  | 2181.149 | 4 | 6.49 | 62.2 | bb | 2023-11-08
249 | 249 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.05 | 2069.581 | 81333 |  | 2069.581 | 4 | 6.16 | 53.9 | bb | 2023-11-08
250 | 250 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.05 | 2232.48 | 87859 |  | 2232.48 | 4 | 6.64 | 66 | bb | 2023-11-08
251 | 251 | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.04 | 0.066 | 4 |  | 0.066 | 4 | 0 | -100 | bb | 2023-11-08
252 | 252 | 20231106_HepCl_ReRun_254 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-08
253 | 253 | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.06 | 0.647 | 38 |  | 0.647 | 4 | 0 | -100 | bb | 2023-11-08
Compound 8: 11H-PFUnDA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.04 | 5.602 | 132 | [5 empty] | bb | 2023-11-06
2 | 2 | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.05 | 4.978 | 128 | [5 empty] | bb | 2023-11-06
3 | 3 | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.07 | 0.235 | 18 | 6.982 | 0.135 |  | 1.62 |  | bb | 2023-11-06
4 | 4 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.07 | 1.874 | 69 | 5.073 | 1.478 |  | 24.05 |  | bb | 2023-11-06
5 | 5 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.01 | 0.655 | 65 | 3.127 | 0.838 |  | 13.36 |  | bb | 2023-11-06
6 | 6 | 20231106_HepCl_ReRun_006 | Blank | Blank | [9 empty] | 2023-11-06
7 | 7 | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.03 | 2.4 | 141 | 7.306 | 1.314 |  | 21.31 |  | bb | 2023-11-06
8 | 8 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 4.07 | 141.617 | 5583 | 49528.125 | 0.011 | 0.175 |  |  | bbI | 2023-11-06
9 | 9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 4.07 | 255.263 | 10129 | 47045.578 | 0.022 | 0.281 |  |  | bbI | 2023-11-06
10 | 10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 4.06 | 284.665 | 9000 | 40876.117 | 0.028 | 0.444 |  |  | bbI | 2023-11-06
11 | 11 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 4.06 | 528.606 | 20359 | 45863.699 | 0.046 | 0.713 | 0.14 | -80.5 | bb | 2023-11-06
12 | 12 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 4.06 | 792.861 | 30425 | 43603.629 | 0.073 | 1.138 | 0.58 | -48.7 | bb | 2023-11-06
13 | 13 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.06 | 1062.268 | 36328 | 40008.617 | 0.106 | 1.825 | 1.14 | -37.4 | bb | 2023-11-06
14 | 14 | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.05 | 2.265 | 131 | 0.639 | 14.178 |  | 236.14 |  | bb | 2023-11-06
15 | 15 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.06 | 1882.456 | 72625 | 42977.938 | 0.175 | 2.919 | 2.3 | -21.4 | bb | 2023-11-06
16 | 16 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.06 | 3141.384 | 120411 | 44416.688 | 0.283 | 4.669 | 4.09 | -12.3 | bb | 2023-11-06
17 | 17 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.07 | 4509.686 | 158142 | 41387.938 | 0.436 | 7.45 | 6.65 | -10.8 | bb | 2023-11-06
18 | 18 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.06 | 7758.193 | 294767 | 43287.93 | 0.717 | 11.875 | 11.34 | -4.5 | bb | 2023-11-06
19 | 19 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.07 | 13921.18 | 525530 | 43650.934 | 1.276 | 19.125 | 20.67 | 8.1 | bb | 2023-11-06
20 | 20 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.06 | 18356.199 | 719052 | 43166.441 | 1.701 | 30.5 | 27.78 | -8.9 | bb | 2023-11-06
21 | 21 | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.09 | 6.331 | 205 | [5 empty] | bb | 2023-11-06
22 | 22 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.06 | 26155.953 | 903097 | 44292.91 | 2.362 | 48.75 | 38.82 | -20.4 | bb | 2023-11-06
23 | 23 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.06 | 50974.664 | 1864530 | 50112.5 | 4.069 | 78.125 | 67.32 | -13.8 | bb | 2023-11-06
24 | 24 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.07 | 71695.906 | 2683730 | 38008.148 | 7.545 | 125 | 125.37 | 0.3 | MM | 2023-11-06
25 | 25 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 4.06 | 131145.125 | 4748852 | 52133.781 | 10.062 | 200 | 167.41 | -16.3 | bb | 2023-11-06
26 | 26 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 4.07 | 190048.672 | 6581417 | 43163.68 | 17.612 | 325.625 | 293.48 | -9.9 | bb | 2023-11-06
27 | 27 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.07 | 4.099 | 215 | 16.466 | 0.996 |  | 16 |  | bb | 2023-11-06
28 | 28 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.03 | 3.303 | 150 | 4.24 | 3.116 |  | 51.41 |  | bb | 2023-11-06
29 | 29 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 4.07 | 442.919 | 18043 | 38003.336 | 0.047 | 0.713 | 0.15 | -79.3 | bb | 2023-11-06
30 | 30 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 4.07 | 702.902 | 25764 | 35276.152 | 0.08 | 1.138 | 0.7 | -38.4 | bb | 2023-11-06
31 | 31 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.07 | 1106.86 | 42580 | 33664.371 | 0.132 | 1.825 | 1.57 | -14.2 | bb | 2023-11-06
32 | 32 | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 4.114 | [5 empty] | 2023-11-06
33 | 33 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 4.07 | 430.506 | 15772 | 37383.543 | 0.046 | 0.713 | 0.14 | -80.6 | bb | 2023-11-06
34 | 34 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 4.07 | 699.594 | 26547 | 33432.625 | 0.084 | 1.138 | 0.77 | -32.6 | bb | 2023-11-06
35 | 35 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.06 | 1030.858 | 39085 | 31861.572 | 0.129 | 1.825 | 1.53 | -16.1 | bb | 2023-11-06
36 | 36 | 20231106_HepCl_ReRun_036 | Blank | Blank |  |  |  | 7.725 | [5 empty] | 2023-11-06
37 | 37 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 4.06 | 532.507 | 21317 | 34334.141 | 0.062 | 0.713 | 0.41 | -43.1 | bb | 2023-11-06
38 | 38 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 4.06 | 689.775 | 23705 | 32710.328 | 0.084 | 1.138 | 0.78 | -31.6 | bb | 2023-11-06
39 | 39 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.07 | 1048.295 | 38562 | 30194.342 | 0.139 | 1.825 | 1.69 | -7.5 | bb | 2023-11-06
40 | 40 | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.07 | 5.768 | 153 | 5.681 | 4.061 |  | 67.19 |  | bb | 2023-11-06
41 | 41 | 20231106_HepCl_ReRun_041 | QC A | QC | 4.07 | 508.329 | 19978 | 49355.238 | 0.041 | 0.625 | 0.06 | -90.8 | bb | 2023-11-06
42 | 42 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.06 | 1949.468 | 71825 | 51642.734 | 0.151 | 2.506 | 1.89 | -24.6 | bb | 2023-11-06
43 | 43 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.06 | 5233.741 | 199717 | 53094.781 | 0.394 | 6.269 | 5.95 | -5 | bb | 2023-11-06
44 | 44 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.07 | 18617.66 | 686901 | 46975.297 | 1.585 | 25 | 25.84 | 3.4 | bb | 2023-11-06
45 | 45 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 9.694 | [5 empty] | 2023-11-06
46 | 46 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.04 | 3.32 | 197 | 1.228 | 10.814 |  | 179.97 |  | bb | 2023-11-06
47 | 47 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 4.06 | 131.287 | 5100 | 31262.33 | 0.017 | 0.175 |  |  | bbI | 2023-11-06
48 | 48 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 4.07 | 206.615 | 7775 | 31310.842 | 0.026 | 0.281 |  |  | bbI | 2023-11-06
49 | 49 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 4.06 | 234.735 | 8332 | 28202.996 | 0.033 | 0.444 |  |  | bbI | 2023-11-06
50 | 50 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 4.06 | 428.645 | 14499 | 29024.73 | 0.059 | 0.713 | 0.36 | -50.1 | bb | 2023-11-06
51 | 51 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 4.07 | 632.781 | 22856 | 28126.109 | 0.09 | 1.138 | 0.87 | -23.3 | bb | 2023-11-06
52 | 52 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.06 | 1061.033 | 39434 | 27641.32 | 0.154 | 1.825 | 1.93 | 5.9 | bb | 2023-11-06
53 | 53 | 20231106_HepCl_ReRun_053 | Blank | Blank |  |  |  | 3.035 | [5 empty] | 2023-11-06
54 | 54 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.06 | 1728.86 | 67243 | 30703.615 | 0.225 | 2.919 | 3.13 | 7.3 | bb | 2023-11-06
55 | 55 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.06 | 2842.049 | 102842 | 30452.596 | 0.373 | 4.669 | 5.6 | 20 | bb | 2023-11-06
56 | 56 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.06 | 4361.453 | 168909 | 29372.057 | 0.594 | 7.45 | 9.29 | 24.7 | bb | 2023-11-06
57 | 57 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.07 | 6099.027 | 213476 | 28786.756 | 0.847 | 11.875 | 13.52 | 13.9 | bb | 2023-11-06
58 | 58 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.06 | 12133.325 | 451289 | 31747.754 | 1.529 | 19.125 | 24.9 | 30.2 | bb | 2023-11-06
59 | 59 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.06 | 19046.547 | 718133 | 31395.316 | 2.427 | 30.5 | 39.89 | 30.8 | bb | 2023-11-06
60 | 60 | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.06 | 11.194 | 371 | 2.979 | 15.031 |  | 250.38 |  | bb | 2023-11-06
61 | 61 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.06 | 28121.049 | 1057346 | 34492.879 | 3.261 | 48.75 | 53.83 | 10.4 | bb | 2023-11-06
62 | 62 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.06 | 49863.824 | 1864208 | 41397.715 | 4.818 | 78.125 | 79.83 | 2.2 | bb | 2023-11-06
63 | 63 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.06 | 70819.211 | 2617717 | 38112.547 | 7.433 | 125 | 123.49 | -1.2 | bb | 2023-11-06
64 | 64 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 4.06 | 139249.141 | 5104794 | 57215.652 | 9.735 | 200 | 161.94 | -19 | bb | 2023-11-06
65 | 65 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 4.07 | 204025.5 | 7220709 | 38809.316 | 21.029 | 325.625 | 350.54 | 7.7 | bb | 2023-11-07
66 | 66 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank |  |  |  | 6.959 | [5 empty] | 2023-11-07
67 | 67 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.03 | 3.021 | 145 | 1.469 | 8.226 |  | 136.74 |  | bb | 2023-11-07
68 | 68 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 4.06 | 422.18 | 15499 | 26842.035 | 0.063 | 0.713 | 0.42 | -41.1 | bb | 2023-11-07
69 | 69 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 4.07 | 686.449 | 26862 | 25787.34 | 0.106 | 1.138 | 1.15 | 0.9 | bb | 2023-11-07
70 | 70 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.06 | 1028.637 | 39775 | 26664.004 | 0.154 | 1.825 | 1.95 | 6.6 | bb | 2023-11-07
71 | 71 | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.04 | 0.246 | 35 | 13.529 | 0.073 |  | 0.58 |  | bb | 2023-11-07
72 | 72 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 4.07 | 372.902 | 13666 | 25236.123 | 0.059 | 0.713 | 0.36 | -50 | bb | 2023-11-07
73 | 73 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 4.06 | 648.499 | 25291 | 25884.043 | 0.1 | 1.138 | 1.04 | -8.3 | bb | 2023-11-07
74 | 74 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.06 | 967.527 | 35145 | 27006.176 | 0.143 | 1.825 | 1.76 | -3.4 | bb | 2023-11-07
75 | 75 | 20231106_HepCl_ReRun_075 | Blank | Blank |  |  |  | 8.646 | [5 empty] | 2023-11-07
76 | 76 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 4.06 | 401.949 | 14818 | 24837.654 | 0.065 | 0.713 | 0.45 | -36.8 | bb | 2023-11-07
77 | 77 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 4.07 | 627.016 | 24357 | 23961.09 | 0.105 | 1.138 | 1.12 | -1.8 | bb | 2023-11-07
78 | 78 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.06 | 993.691 | 37872 | 24890.996 | 0.16 | 1.825 | 2.04 | 11.6 | bb | 2023-11-07
79 | 79 | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 10.156 | [5 empty] | 2023-11-07
80 | 80 | 20231106_HepCl_ReRun_080 | QC A | QC | 4.07 | 502.869 | 19829 | 36659.156 | 0.055 | 0.625 | 0.29 | -54.3 | bb | 2023-11-07
81 | 81 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.06 | 2080.625 | 81819 | 37154.508 | 0.224 | 2.506 | 3.11 | 24.1 | bb | 2023-11-07
82 | 82 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.06 | 4570.148 | 205965 | 41341.777 | 0.442 | 6.269 | 6.75 | 7.7 | MM | 2023-11-07
83 | 83 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.06 | 15100.404 | 616116 | 32931.387 | 1.834 | 25 | 30 | 20 | MM | 2023-11-07
84 | 84 | 20231106_HepCl_ReRun_084 | Blank | Blank | 4.06 | 4.429 | 252 | 4.803 | 3.689 |  | 60.97 |  | bb | 2023-11-07
85 | 85 | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.06 | 5.79 | 191 | 0.242 | 95.702 |  | 1597.58 |  | bb | 2023-11-07
86 | 86 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 4.06 | 18915.379 | 692422 | 25853.348 | 2.927 | 30.5 | 48.24 | 58.2 | bb | 2023-11-07
87 | 87 | 20231106_HepCl_ReRun_087 | Blank | Blank | 4.09 | 9.932 | 215 | [5 empty] | bb | 2023-11-07
88 | 88 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte |  |  |  | 98002.758 | [5 empty] | 2023-11-07
89 | 89 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte |  |  |  | 88635.797 | [5 empty] | 2023-11-07
90 | 90 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte |  |  |  | 89190.406 | [5 empty] | 2023-11-07
91 | 91 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte |  |  |  | 89103.305 | [5 empty] | 2023-11-07
92 | 92 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.06 | 18.349 | 722 | 83877.867 | 0.001 |  |  |  | bbI | 2023-11-07
93 | 93 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.05 | 7.83 | 256 | 86799.516 | 0 |  |  |  | bbI | 2023-11-07
94 | 94 | 20231106_HepCl_ReRun_094 | Blank | Blank | 4.08 | 2.282 | 84 | 0.119 | 76.706 |  | 1280.34 |  | bb | 2023-11-07
95 | 95 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.07 | 3.867 | 251 | 93797.602 | 0 |  |  |  | bbI | 2023-11-07
96 | 96 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.07 | 16.547 | 987 | 95630.289 | 0.001 |  |  |  | bbI | 2023-11-07
97 | 97 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.06 | 8.845 | 502 | 88288.648 | 0 |  |  |  | bbI | 2023-11-07
98 | 98 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.07 | 7.013 | 311 | 88574.773 | 0 |  |  |  | bbI | 2023-11-07
99 | 99 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.07 | 9.717 | 489 | 95312.484 | 0 |  |  |  | bbI | 2023-11-07
100 | 100 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte |  |  |  | 94841.609 | [5 empty] | 2023-11-07
101 | 101 | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.04 | 13.821 | 319 | 1.22 | 45.315 |  | 756.11 |  | bb | 2023-11-07
102 | 102 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.06 | 11.272 | 433 | 76076.914 | 0.001 |  |  |  | bbI | 2023-11-07
103 | 103 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.07 | 7.606 | 361 | 76258 | 0 |  |  |  | bbI | 2023-11-07
104 | 104 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.06 | 8.773 | 389 | 80176.891 | 0 |  |  |  | bbI | 2023-11-07
105 | 105 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.06 | 0.389 | 70 | 65362.707 | 0 |  |  |  | bbI | 2023-11-07
106 | 106 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.06 | 4.735 | 316 | 66185.563 | 0 |  |  |  | bbI | 2023-11-07
107 | 107 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.06 | 18.365 | 455 | 64907.961 | 0.001 |  |  |  | bbI | 2023-11-07
108 | 108 | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.07 | 2.447 | 167 | 7.197 | 1.36 |  | 22.08 |  | bb | 2023-11-07
109 | 109 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 4.06 | 18737.344 | 706695 | 23824.189 | 3.146 | 30.5 | 51.91 | 70.2 | bb | 2023-11-07
110 | 110 | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.06 | 3.336 | 134 | 1.133 | 11.778 |  | 196.05 |  | bb | 2023-11-07
111 | 111 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.06 | 64.871 | 2246 | 78984.391 | 0.003 |  |  |  | bbI | 2023-11-07
112 | 112 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte |  |  |  | 54687.828 | [5 empty] | 2023-11-07
113 | 113 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.07 | 6.07 | 347 | 63527.227 | 0 |  |  |  | bbI | 2023-11-07
114 | 114 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.07 | 6.733 | 385 | 81925.484 | 0 |  |  |  | bbI | 2023-11-07
115 | 115 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.05 | 12.484 | 452 | 77283.281 | 0.001 |  |  |  | bbI | 2023-11-07
116 | 116 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.06 | 2.375 | 131 | 77620.695 | 0 |  |  |  | bbI | 2023-11-07
117 | 117 | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.02 | 3.587 | 216 | 1.785 | 8.038 |  | 133.6 |  | bb | 2023-11-07
118 | 118 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.06 | 4.094 | 252 | 86903.203 | 0 |  |  |  | bbI | 2023-11-07
119 | 119 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte |  |  |  | 64245.207 | [5 empty] | 2023-11-07
120 | 120 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.06 | 20.189 | 1028 | 84458.141 | 0.001 |  |  |  | bbI | 2023-11-07
121 | 121 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.06 | 68.218 | 3260 | 87830.766 | 0.003 |  |  |  | bbI | 2023-11-07
122 | 122 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.07 | 17.151 | 687 | 82120.75 | 0.001 |  |  |  | bbI | 2023-11-07
123 | 123 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.07 | 17.723 | 611 | 76499.594 | 0.001 |  |  |  | bbI | 2023-11-07
124 | 124 | 20231106_HepCl_ReRun_124 | Blank | Blank | 4.06 | 3.813 | 215 | 8.867 | 1.72 |  | 28.09 |  | bb | 2023-11-07
125 | 125 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte |  |  |  | 73346.039 | [5 empty] | 2023-11-07
126 | 126 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.06 | 14.537 | 793 | 68280.641 | 0.001 |  |  |  | bbI | 2023-11-07
127 | 127 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.06 | 11.696 | 573 | 73499.289 | 0.001 |  |  |  | bbI | 2023-11-07
128 | 128 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte |  |  |  | 55749.684 | [5 empty] | 2023-11-07
129 | 129 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte |  |  |  | 30788.695 | [5 empty] | 2023-11-07
130 | 130 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.07 | 7.596 | 468 | 54407.445 | 0.001 |  |  |  | bbI | 2023-11-07
131 | 131 | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.06 | 0.554 | 114 | 0.21 | 10.552 |  | 175.59 |  | bb | 2023-11-07
132 | 132 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 4.06 | 17263.254 | 649328 | 21559.432 | 3.203 | 30.5 | 52.86 | 73.3 | bb | 2023-11-07
133 | 133 | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.09 | 2.798 | 175 | 1.004 | 11.147 |  | 185.53 |  | bb | 2023-11-07
134 | 134 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte |  |  |  | 76163.773 | [5 empty] | 2023-11-07
135 | 135 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.07 | 1.967 | 217 | 92251.773 | 0 |  |  |  | bbI | 2023-11-07
136 | 136 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.05 | 4.014 | 281 | 83986.945 | 0 |  |  |  | bbI | 2023-11-07
137 | 137 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.06 | 10.213 | 486 | 86446.32 | 0 |  |  |  | bbI | 2023-11-07
138 | 138 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte |  |  |  | 89926.844 | [5 empty] | 2023-11-07
139 | 139 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.09 | 16.609 | 464 | 91128.219 | 0.001 |  |  |  | bbI | 2023-11-07
140 | 140 | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.06 | 0.123 | 16 | 3.553 | 0.138 |  | 1.68 |  | bb | 2023-11-07
141 | 141 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 4.06 | 18161.553 | 695896 | 19427.527 | 3.739 | 30.5 | 61.82 | 102.7 | bb | 2023-11-07
142 | 142 | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.1 | 3.614 | 178 | 4.462 | 3.24 |  | 53.47 |  | bb | 2023-11-07
143 | 143 | 20231106_HepCl_ReRun_143 | Blank | Blank | [9 empty] | 2023-11-07
144 | 144 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.06 | 81.691 | 3339 | 107375.094 | 0.003 |  |  |  | bbI | 2023-11-07
145 | 145 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.06 | 71.185 | 2151 | 104266.414 | 0.003 |  |  |  | bbI | 2023-11-07
146 | 146 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.06 | 105.631 | 3989 | 105538.766 | 0.004 |  |  |  | bbI | 2023-11-07
147 | 147 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.07 | 57.227 | 2282 | 84474.836 | 0.003 |  |  |  | bbI | 2023-11-07
148 | 148 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.07 | 6.926 | 306 | 87116.578 | 0 |  |  |  | bbI | 2023-11-07
149 | 149 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.06 | 3.727 | 259 | 83468.359 | 0 |  |  |  | bbI | 2023-11-07
150 | 150 | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.05 | 4.853 | 276 | 6.476 | 2.998 |  | 49.43 |  | bb | 2023-11-07
151 | 151 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.06 | 12.387 | 330 | 89927.414 | 0.001 |  |  |  | bbI | 2023-11-07
152 | 152 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.06 | 10.963 | 199 | 86421.602 | 0.001 |  |  |  | bbI | 2023-11-07
153 | 153 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.07 | 13.604 | 430 | 91154.789 | 0.001 |  |  |  | bbI | 2023-11-07
154 | 154 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.03 | 16.884 | 380 | 89026.438 | 0.001 |  |  |  | bbI | 2023-11-07
155 | 155 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.06 | 13.921 | 612 | 89150.375 | 0.001 |  |  |  | bbI | 2023-11-07
156 | 156 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.07 | 29.097 | 1028 | 89808.383 | 0.001 |  |  |  | bbI | 2023-11-07
157 | 157 | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.02 | 4.052 | 238 | 0.75 | 21.611 |  | 360.26 |  | bb | 2023-11-07
158 | 158 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.06 | 5.17 | 309 | 78967.586 | 0 |  |  |  | bbI | 2023-11-07
159 | 159 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.03 | 0.403 | 62 | 83951.203 | 0 |  |  |  | bbI | 2023-11-07
160 | 160 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.06 | 68.997 | 2377 | 82722.727 | 0.003 |  |  |  | bbI | 2023-11-07
161 | 161 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.06 | 0.798 | 93 | 80214.852 | 0 |  |  |  | bbI | 2023-11-07
162 | 162 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte |  |  |  | 77124.359 | [5 empty] | 2023-11-07
163 | 163 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.07 | 2.654 | 165 | 74714.813 | 0 |  |  |  | bbI | 2023-11-07
164 | 164 | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.04 | 8.785 | 356 | 3.511 | 10.009 |  | 166.51 |  | bb | 2023-11-07
165 | 165 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 4.06 | 15258.011 | 543415 | 19198.016 | 3.179 | 30.5 | 52.46 | 72 | bb | 2023-11-07
166 | 166 | 20231106_HepCl_ReRun_168 | Blank | Blank | [9 empty] | 2023-11-07
167 | 167 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.06 | 13.18 | 449 | 99319.828 | 0.001 |  |  |  | bbI | 2023-11-07
168 | 168 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte |  |  |  | 94711.563 | [5 empty] | 2023-11-07
169 | 169 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.06 | 7.088 | 331 | 90822.211 | 0 |  |  |  | bbI | 2023-11-07
170 | 170 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.07 | 5.175 | 302 | 84107.078 | 0 |  |  |  | bbI | 2023-11-07
171 | 171 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte |  |  |  | 82892.469 | [5 empty] | 2023-11-07
172 | 172 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte |  |  |  | 89697.148 | [5 empty] | 2023-11-07
173 | 173 | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.02 | 0.679 | 63 | 6.395 | 0.425 |  | 6.46 |  | bb | 2023-11-07
174 | 174 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.07 | 8.217 | 261 | 84959.609 | 0 |  |  |  | bbI | 2023-11-07
175 | 175 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte |  |  |  | 86283.852 | [5 empty] | 2023-11-07
176 | 176 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.07 | 13.263 | 417 | 100046.359 | 0.001 |  |  |  | bbI | 2023-11-07
177 | 177 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.07 | 7.463 | 338 | 95674.383 | 0 |  |  |  | bbI | 2023-11-07
178 | 178 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte |  |  |  | 93804.273 | [5 empty] | 2023-11-07
179 | 179 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.11 | 1.75 | 98 | 75229.672 | 0 |  |  |  | bbI | 2023-11-07
180 | 180 | 20231106_HepCl_ReRun_182 | Blank | Blank |  |  |  | 0.43 | [5 empty] | 2023-11-07
181 | 181 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.06 | 38.406 | 1610 | 87124.805 | 0.002 |  |  |  | bbI | 2023-11-07
182 | 182 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.02 | 2.931 | 191 | 77005.422 | 0 |  |  |  | bbI | 2023-11-07
183 | 183 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.08 | 9.604 | 240 | 88369.641 | 0 |  |  |  | bbI | 2023-11-07
184 | 184 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.06 | 0.168 | 31 | 76210.016 | 0 |  |  |  | bbI | 2023-11-07
185 | 185 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.09 | 0.105 | 18 | 52721.07 | 0 |  |  |  | bbI | 2023-11-07
186 | 186 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.07 | 9.736 | 352 | 55817.023 | 0.001 |  |  |  | bbI | 2023-11-07
187 | 187 | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.1 | 4.923 | 175 | 2.904 | 6.781 |  | 112.61 |  | bb | 2023-11-07
188 | 188 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 4.07 | 16958.822 | 643563 | 21716.965 | 3.124 | 30.5 | 51.53 | 69 | bb | 2023-11-07
189 | 189 | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.04 | 7.196 | 210 | 2.254 | 12.77 |  | 212.63 |  | bb | 2023-11-07
190 | 190 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte |  |  |  | 99776.836 | [5 empty] | 2023-11-07
191 | 191 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.06 | 55.627 | 2127 | 101979.219 | 0.002 |  |  |  | bbI | 2023-11-07
192 | 192 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.07 | 3.523 | 218 | 114534.242 | 0 |  |  |  | bbI | 2023-11-07
193 | 193 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.06 | 8.106 | 373 | 79844.398 | 0 |  |  |  | bbI | 2023-11-07
194 | 194 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.06 | 9.177 | 373 | 86026.953 | 0 |  |  |  | bbI | 2023-11-07
195 | 195 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.06 | 3.263 | 202 | 90566.922 | 0 |  |  |  | bbI | 2023-11-07
196 | 196 | 20231106_HepCl_ReRun_198 | Blank | Blank |  |  |  | 0.96 | [5 empty] | 2023-11-08
197 | 197 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 4.07 | 17834.35 | 672807 | 19258.826 | 3.704 | 30.5 | 61.23 | 100.7 | bb | 2023-11-08
198 | 198 | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.04 | 2.128 | 95 | 1.102 | 7.724 |  | 128.36 |  | bb | 2023-11-08
199 | 199 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.07 | 84765.813 | 3149459 | 97325.922 | 3.484 |  | 57.55 |  | bb | 2023-11-08
200 | 200 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.06 | 89158.578 | 3440956 | 102662.727 | 3.474 |  | 57.38 |  | bb | 2023-11-08
201 | 201 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.06 | 89658.922 | 3423202 | 99524.109 | 3.604 |  | 59.55 |  | bb | 2023-11-08
202 | 202 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.06 | 83136.398 | 3056431 | 104638.984 | 3.178 |  | 52.44 |  | bb | 2023-11-08
203 | 203 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.06 | 73273.953 | 2781368 | 95722.523 | 3.062 |  | 50.5 |  | bb | 2023-11-08
204 | 204 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.07 | 91674.398 | 3524289 | 104475.281 | 3.51 |  | 57.98 |  | bb | 2023-11-08
205 | 205 | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.1 | 3.972 | 181 | 1.83 | 8.682 |  | 144.36 |  | bb | 2023-11-08
206 | 206 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.06 | 69076.906 | 2644574 | 96895.539 | 2.852 |  | 46.99 |  | bb | 2023-11-08
207 | 207 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.06 | 80887.133 | 3075320 | 90586.703 | 3.572 |  | 59.02 |  | bb | 2023-11-08
208 | 208 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.06 | 88501.672 | 3452261 | 85220.258 | 4.154 |  | 68.74 |  | bb | 2023-11-08
209 | 209 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.06 | 83890.813 | 3278919 | 87958.188 | 3.815 |  | 63.08 |  | bb | 2023-11-08
210 | 210 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.06 | 74294.75 | 2865533 | 73946.117 | 4.019 |  | 66.48 |  | bb | 2023-11-08
211 | 211 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.06 | 88411.766 | 3434985 | 88629.414 | 3.99 |  | 66 |  | bb | 2023-11-08
212 | 212 | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.07 | 12.603 | 469 | 1.23 | 40.985 |  | 683.81 |  | bb | 2023-11-08
213 | 213 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.06 | 82879.844 | 2973760 | 93904.141 | 3.53 |  | 58.33 |  | bb | 2023-11-08
214 | 214 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.06 | 83693.422 | 3138311 | 99092.422 | 3.378 |  | 55.79 |  | bb | 2023-11-08
215 | 215 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.07 | 85406.219 | 3482632 | 102824.828 | 3.322 |  | 54.85 |  | bb | 2023-11-08
216 | 216 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.07 | 60181.426 | 2252242 | 74625.852 | 3.226 |  | 53.24 |  | bb | 2023-11-08
217 | 217 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.07 | 66792.984 | 2494126 | 78271.125 | 3.413 |  | 56.37 |  | bb | 2023-11-08
218 | 218 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.07 | 63163.129 | 2425594 | 82232.203 | 3.072 |  | 50.68 |  | bb | 2023-11-08
219 | 219 | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.02 | 10.347 | 233 | 0.419 | 98.778 |  | 1648.94 |  | bb | 2023-11-08
220 | 220 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 4.07 | 17293.111 | 658177 | 19805.637 | 3.493 | 30.5 | 57.69 | 89.2 | bb | 2023-11-08
221 | 221 | 20231106_HepCl_ReRun_223 | Blank | Blank | 4.05 | 3.697 | 187 | 5.807 | 2.547 |  | 41.9 |  | bb | 2023-11-08
222 | 222 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.06 | 83874.047 | 3258557 | 68976.289 | 4.864 |  | 80.6 |  | bb | 2023-11-08
223 | 223 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.07 | 72234.156 | 2659945 | 80912.125 | 3.571 |  | 59 |  | bb | 2023-11-08
224 | 224 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.06 | 76960.344 | 2921204 | 64799.211 | 4.751 |  | 78.7 |  | bb | 2023-11-08
225 | 225 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.07 | 93146.758 | 3627092 | 95344.883 | 3.908 |  | 64.63 |  | bb | 2023-11-08
226 | 226 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.07 | 77323.891 | 2759422 | 98936.914 | 3.126 |  | 51.58 |  | bb | 2023-11-08
227 | 227 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.06 | 88055.625 | 3283423 | 96992.188 | 3.631 |  | 60.01 |  | bb | 2023-11-08
228 | 228 | 20231106_HepCl_ReRun_230 | Blank | Blank | 4.08 | 13.86 | 401 | 2.226 | 24.906 |  | 415.29 |  | bb | 2023-11-08
229 | 229 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.07 | 96175.375 | 3901113 | 108030.414 | 3.561 |  | 58.84 |  | bb | 2023-11-08
230 | 230 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.06 | 85629.828 | 3304683 | 100591.141 | 3.405 |  | 56.23 |  | bb | 2023-11-08
231 | 231 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.07 | 81606.148 | 3017306 | 69586.211 | 4.691 |  | 77.71 |  | bb | 2023-11-08
232 | 232 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.06 | 87272.586 | 3324550 | 104467.281 | 3.342 |  | 55.17 |  | bb | 2023-11-08
233 | 233 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.06 | 83227.961 | 3148725 | 48669.789 | 6.84 |  | 113.6 |  | bb | 2023-11-08
234 | 234 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.06 | 73634.305 | 2678980 | 39606.652 | 7.437 |  | 123.56 |  | bb | 2023-11-08
235 | 235 | 20231106_HepCl_ReRun_237 | Blank | Blank | 4.06 | 2.31 | 145 | 8.198 | 1.127 |  | 18.19 |  | bb | 2023-11-08
236 | 236 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.06 | 82758.57 | 3205700 | 100054.281 | 3.309 |  | 54.62 |  | bb | 2023-11-08
237 | 237 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.07 | 84048.07 | 3297586 | 91328.695 | 3.681 |  | 60.84 |  | bb | 2023-11-08
238 | 238 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.07 | 90018.523 | 3535425 | 102183.391 | 3.524 |  | 58.22 |  | bb | 2023-11-08
239 | 239 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.07 | 68525.758 | 2553414 | 95108.883 | 2.882 |  | 47.5 |  | bb | 2023-11-08
240 | 240 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.07 | 58218.891 | 2193878 | 53921.77 | 4.319 |  | 71.49 |  | bb | 2023-11-08
241 | 241 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.07 | 65500.477 | 2518845 | 61000.227 | 4.295 |  | 71.1 |  | bb | 2023-11-08
242 | 242 | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 7.8 | [5 empty] | 2023-11-08
243 | 243 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 4.07 | 18489.309 | 735656 | 20601.318 | 3.59 | 30.5 | 59.32 | 94.5 | bb | 2023-11-08
244 | 244 | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.08 | 0.727 | 65 | 4.036 | 0.721 |  | 11.4 |  | bb | 2023-11-08
245 | 245 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.06 | 98320.117 | 3693365 | 180164.406 | 2.183 |  | 35.82 |  | bb | 2023-11-08
246 | 246 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.07 | 84453.852 | 3030661 | 128731.063 | 2.624 |  | 43.19 |  | bb | 2023-11-08
247 | 247 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.06 | 98009.695 | 3569567 | 119917.117 | 3.269 |  | 53.96 |  | bb | 2023-11-08
248 | 248 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.06 | 66279.32 | 2477570 | 88866.891 | 2.983 |  | 49.19 |  | bb | 2023-11-08
249 | 249 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.06 | 66022.867 | 2635104 | 88157.313 | 2.996 |  | 49.4 |  | bb | 2023-11-08
250 | 250 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.06 | 59553.82 | 2256071 | 98938.063 | 2.408 |  | 39.58 |  | bb | 2023-11-08
251 | 251 | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.06 | 0.597 | 60 | 1.238 | 1.929 |  | 31.58 |  | bb | 2023-11-08
252 | 252 | 20231106_HepCl_ReRun_254 | Blank | Blank | 4.08 | 4.24 | 210 | 7.609 | 2.229 |  | 36.59 |  | bb | 2023-11-08
253 | 253 | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.01 | 0.367 | 74 | 3 | 0.489 |  | 7.54 |  | bb | 2023-11-08
Compound 9: PFUnDA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20231106_HepCl_ReRun_001 | Blank | Blank | 4.42 | 18.895 | 1390 | [5 empty] | bb | 2023-11-06
2 | 2 | 20231106_HepCl_ReRun_002 | Blank | Blank | 4.41 | 62.547 | 1907 | [5 empty] | bb | 2023-11-06
3 | 3 | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.39 | 1.965 | 356 | 6.982 | 1.126 |  | 3.36 |  | bb | 2023-11-06
4 | 4 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.43 | 13.412 | 1130 | 5.073 | 10.575 |  | 36.17 |  | bb | 2023-11-06
5 | 5 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.37 | 140.794 | 2229 | 3.127 | 180.101 |  | 624.63 |  | bb | 2023-11-06
6 | 6 | 20231106_HepCl_ReRun_006 | Blank | Blank | 4.39 | 12.852 | 1235 | [5 empty] | bb | 2023-11-06
7 | 7 | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.35 | 36.801 | 1623 | 7.306 | 20.148 |  | 69.4 |  | bb | 2023-11-06
8 | 8 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 4.4 | 2611.916 | 113399 | 49528.125 | 0.211 | 0.175 | 0.19 | 7.5 | bb | 2023-11-06
9 | 9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 4.4 | 2187.755 | 92047 | 47045.578 | 0.186 | 0.281 | 0.1 | -63.9 | bb | 2023-11-06
10 | 10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 4.4 | 2251.235 | 90945 | 40876.117 | 0.22 | 0.444 | 0.22 | -50.3 | bb | 2023-11-06
11 | 11 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 4.4 | 5824.572 | 243158 | 45863.699 | 0.508 | 0.713 | 1.22 | 71.1 | bb | 2023-11-06
12 | 12 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 4.4 | 4827.899 | 196931 | 43603.629 | 0.443 | 1.138 | 0.99 | -12.7 | bb | 2023-11-06
13 | 13 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.4 | 7043 | 277350 | 40008.617 | 0.704 | 1.825 | 1.9 | 4.1 | bb | 2023-11-06
14 | 14 | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.42 | 1.451 | 372 | 0.639 | 9.083 |  | 30.99 |  | bb | 2023-11-06
15 | 15 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.4 | 10383.962 | 423591 | 42977.938 | 0.966 | 2.919 | 2.81 | -3.7 | bb | 2023-11-06
16 | 16 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.4 | 16841.48 | 684362 | 44416.688 | 1.517 | 4.669 | 4.72 | 1.1 | bb | 2023-11-06
17 | 17 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.4 | 23083.746 | 919942 | 41387.938 | 2.231 | 7.45 | 7.2 | -3.4 | bb | 2023-11-06
18 | 18 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.4 | 36668.535 | 1499343 | 43287.93 | 3.388 | 11.875 | 11.22 | -5.5 | bb | 2023-11-06
19 | 19 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.4 | 63299.324 | 2670024 | 43650.934 | 5.801 | 19.125 | 19.59 | 2.4 | bb | 2023-11-06
20 | 20 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.39 | 110621.594 | 4408315 | 43166.441 | 10.251 | 30.5 | 35.04 | 14.9 | bb | 2023-11-06
21 | 21 | 20231106_HepCl_ReRun_021 | Blank | Blank | 4.43 | 18.559 | 1477 | [5 empty] | bb | 2023-11-06
22 | 22 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.4 | 142737.063 | 5848686 | 44292.91 | 12.89 | 48.75 | 44.2 | -9.3 | bb | 2023-11-06
23 | 23 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.4 | 284172 | 11155492 | 50112.5 | 22.683 | 78.125 | 78.19 | 0.1 | bb | 2023-11-06
24 | 24 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.4 | 376289.313 | 15172176 | 38008.148 | 39.601 | 125 | 136.92 | 9.5 | bb | 2023-11-06
25 | 25 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 4.4 | 619418.688 | 22182238 | 52133.781 | 47.525 | 200 | 164.43 | -17.8 | bb | 2023-11-06
26 | 26 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 4.4 | 1052163.5 | 39965172 | 43163.68 | 97.505 | 325.625 | 337.92 | 3.8 | bb | 2023-11-06
27 | 27 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.38 | 59.634 | 2149 | 16.466 | 14.487 |  | 49.74 |  | bb | 2023-11-06
28 | 28 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.36 | 23.582 | 2277 | 4.24 | 22.247 |  | 76.68 |  | bb | 2023-11-06
29 | 29 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 4.4 | 5001.823 | 214961 | 38003.336 | 0.526 | 0.713 | 1.28 | 80.1 | bb | 2023-11-06
30 | 30 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 4.4 | 4156.563 | 170807 | 35276.152 | 0.471 | 1.138 | 1.09 | -4 | bb | 2023-11-06
31 | 31 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.4 | 5588.62 | 236229 | 33664.371 | 0.664 | 1.825 | 1.76 | -3.5 | bb | 2023-11-06
32 | 32 | 20231106_HepCl_ReRun_032 | Blank | Blank |  |  |  | 4.114 | [5 empty] | 2023-11-06
33 | 33 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 4.4 | 4821.958 | 198195 | 37383.543 | 0.516 | 0.713 | 1.25 | 75 | bb | 2023-11-06
34 | 34 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 4.4 | 4101.203 | 162869 | 33432.625 | 0.491 | 1.138 | 1.16 | 1.9 | bb | 2023-11-06
35 | 35 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.4 | 5433.467 | 207631 | 31861.572 | 0.682 | 1.825 | 1.82 | -0.1 | bb | 2023-11-06
36 | 36 | 20231106_HepCl_ReRun_036 | Blank | Blank |  |  |  | 7.725 | [5 empty] | 2023-11-06
37 | 37 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 4.4 | 4405.595 | 180554 | 34334.141 | 0.513 | 0.713 | 1.24 | 73.7 | bb | 2023-11-06
38 | 38 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 4.39 | 3882.14 | 161436 | 32710.328 | 0.475 | 1.138 | 1.1 | -3 | bb | 2023-11-06
39 | 39 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.4 | 5458.212 | 220757 | 30194.342 | 0.723 | 1.825 | 1.97 | 7.7 | bb | 2023-11-06
40 | 40 | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.44 | 65.969 | 2411 | 5.681 | 46.449 |  | 160.69 |  | bb | 2023-11-06
41 | 41 | 20231106_HepCl_ReRun_041 | QC A | QC | 4.4 | 3405.577 | 139548 | 49355.238 | 0.276 | 0.625 | 0.41 | -33.8 | bb | 2023-11-06
42 | 42 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.4 | 10822.062 | 445443 | 51642.734 | 0.838 | 2.506 | 2.37 | -5.6 | bb | 2023-11-06
43 | 43 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.4 | 26934.248 | 1136498 | 53094.781 | 2.029 | 6.269 | 6.5 | 3.7 | bb | 2023-11-06
44 | 44 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.4 | 85835.969 | 3549459 | 46975.297 | 7.309 | 25 | 24.83 | -0.7 | bb | 2023-11-06
45 | 45 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank |  |  |  | 9.694 | [5 empty] | 2023-11-06
46 | 46 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank |  |  |  | 1.228 | [5 empty] | 2023-11-06
47 | 47 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 4.4 | 1618.11 | 66118 | 31262.33 | 0.207 | 0.175 | 0.17 | -0.3 | bb | 2023-11-06
48 | 48 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 4.4 | 1399.847 | 59255 | 31310.842 | 0.179 | 0.281 | 0.08 | -72.8 | bb | 2023-11-06
49 | 49 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 4.4 | 1610.389 | 65901 | 28202.996 | 0.228 | 0.444 | 0.25 | -44 | bb | 2023-11-06
50 | 50 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 4.4 | 3810.877 | 157128 | 29024.73 | 0.525 | 0.713 | 1.28 | 79.5 | bb | 2023-11-06
51 | 51 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 4.4 | 3269.672 | 137616 | 28126.109 | 0.465 | 1.138 | 1.07 | -5.9 | bb | 2023-11-06
52 | 52 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.4 | 4747.757 | 201851 | 27641.32 | 0.687 | 1.825 | 1.84 | 0.9 | bb | 2023-11-06
53 | 53 | 20231106_HepCl_ReRun_053 | Blank | Blank | 4.43 | 2.767 | 722 | 3.035 | 3.647 |  | 12.11 |  | bb | 2023-11-06
54 | 54 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.39 | 6970.027 | 276378 | 30703.615 | 0.908 | 2.919 | 2.61 | -10.7 | bb | 2023-11-06
55 | 55 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.4 | 11831.235 | 485726 | 30452.596 | 1.554 | 4.669 | 4.85 | 3.9 | bb | 2023-11-06
56 | 56 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.4 | 16286.906 | 675195 | 29372.057 | 2.218 | 7.45 | 7.16 | -4 | bb | 2023-11-06
57 | 57 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.4 | 23608.922 | 980457 | 28786.756 | 3.281 | 11.875 | 10.84 | -8.7 | bb | 2023-11-06
58 | 58 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.4 | 45632.473 | 1810515 | 31747.754 | 5.749 | 19.125 | 19.41 | 1.5 | bb | 2023-11-06
59 | 59 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.4 | 81080.563 | 3271567 | 31395.316 | 10.33 | 30.5 | 35.31 | 15.8 | bb | 2023-11-06
60 | 60 | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.43 | 61.475 | 1351 | 2.979 | 82.544 |  | 285.99 |  | bb | 2023-11-06
61 | 61 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.4 | 110454.477 | 4597438 | 34492.879 | 12.809 | 48.75 | 43.92 | -9.9 | bb | 2023-11-06
62 | 62 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.4 | 222048.391 | 9055133 | 41397.715 | 21.455 | 78.125 | 73.93 | -5.4 | bb | 2023-11-06
63 | 63 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.4 | 327021.594 | 13200246 | 38112.547 | 34.322 | 125 | 118.6 | -5.1 | bb | 2023-11-06
64 | 64 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 4.4 | 652959.25 | 26559178 | 57215.652 | 45.649 | 200 | 157.92 | -21 | bb | 2023-11-06
65 | 65 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 4.4 | 930327.5 | 35965604 | 38809.316 | 95.887 | 325.625 | 332.3 | 2.1 | bb | 2023-11-07
66 | 66 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4.43 | 0.86 | 403 | 6.959 | 0.494 |  | 1.17 |  | bb | 2023-11-07
67 | 67 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank |  |  |  | 1.469 | [5 empty] | 2023-11-07
68 | 68 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 4.4 | 3619.746 | 142326 | 26842.035 | 0.539 | 0.713 | 1.33 | 86.4 | bb | 2023-11-07
69 | 69 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 4.4 | 3183.645 | 129738 | 25787.34 | 0.494 | 1.138 | 1.17 | 2.9 | bb | 2023-11-07
70 | 70 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.4 | 4767.315 | 194365 | 26664.004 | 0.715 | 1.825 | 1.94 | 6.2 | bb | 2023-11-07
71 | 71 | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.45 | 47.228 | 2391 | 13.529 | 13.963 |  | 47.93 |  | bb | 2023-11-07
72 | 72 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 4.4 | 3276.574 | 133935 | 25236.123 | 0.519 | 0.713 | 1.26 | 76.7 | bb | 2023-11-07
73 | 73 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 4.4 | 3061.236 | 126429 | 25884.043 | 0.473 | 1.138 | 1.1 | -3.5 | bb | 2023-11-07
74 | 74 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.4 | 4664.719 | 193278 | 27006.176 | 0.691 | 1.825 | 1.85 | 1.6 | bb | 2023-11-07
75 | 75 | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.38 | 69.248 | 2207 | 8.646 | 32.037 |  | 110.66 |  | bb | 2023-11-07
76 | 76 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 4.4 | 3055.906 | 131424 | 24837.654 | 0.492 | 0.713 | 1.16 | 63.4 | bb | 2023-11-07
77 | 77 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 4.4 | 2840.368 | 113787 | 23961.09 | 0.474 | 1.138 | 1.1 | -3.1 | bb | 2023-11-07
78 | 78 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.39 | 4229.273 | 162426 | 24890.996 | 0.68 | 1.825 | 1.82 | -0.5 | bb | 2023-11-07
79 | 79 | 20231106_HepCl_ReRun_079 | Blank | Blank |  |  |  | 10.156 | [5 empty] | 2023-11-07
80 | 80 | 20231106_HepCl_ReRun_080 | QC A | QC | 4.4 | 2607.647 | 110211 | 36659.156 | 0.285 | 0.625 | 0.44 | -29 | bb | 2023-11-07
81 | 81 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.4 | 7522.85 | 298162 | 37154.508 | 0.81 | 2.506 | 2.27 | -9.5 | bb | 2023-11-07
82 | 82 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.4 | 21356.891 | 869820 | 41341.777 | 2.066 | 6.269 | 6.63 | 5.7 | bb | 2023-11-07
83 | 83 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.4 | 60284.25 | 2577048 | 32931.387 | 7.322 | 25 | 24.87 | -0.5 | bb | 2023-11-07
84 | 84 | 20231106_HepCl_ReRun_084 | Blank | Blank |  |  |  | 4.803 | [5 empty] | 2023-11-07
85 | 85 | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.36 | 4.824 | 525 | 0.242 | 79.736 |  | 276.24 |  | bb | 2023-11-07
86 | 86 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 4.4 | 64428.176 | 2543358 | 25853.348 | 9.968 | 30.5 | 34.06 | 11.7 | bb | 2023-11-07
87 | 87 | 20231106_HepCl_ReRun_087 | Blank | Blank | 4.36 | 22.948 | 1505 | [5 empty] | bb | 2023-11-07
88 | 88 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.4 | 288979.969 | 11903002 | 98002.758 | 11.795 |  | 40.4 |  | bb | 2023-11-07
89 | 89 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.4 | 252387.266 | 10245247 | 88635.797 | 11.39 |  | 38.99 |  | bb | 2023-11-07
90 | 90 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.4 | 232170.688 | 9274514 | 89190.406 | 10.412 |  | 35.6 |  | bb | 2023-11-07
91 | 91 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.4 | 267149.969 | 10917079 | 89103.305 | 11.993 |  | 41.09 |  | bb | 2023-11-07
92 | 92 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.4 | 231363.469 | 9350834 | 83877.867 | 11.033 |  | 37.76 |  | bb | 2023-11-07
93 | 93 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.39 | 257262.188 | 10705604 | 86799.516 | 11.855 |  | 40.61 |  | bb | 2023-11-07
94 | 94 | 20231106_HepCl_ReRun_094 | Blank | Blank |  |  |  | 0.119 | [5 empty] | 2023-11-07
95 | 95 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.4 | 279341.906 | 11378870 | 93797.602 | 11.913 |  | 40.81 |  | bb | 2023-11-07
96 | 96 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.4 | 282657.531 | 11606108 | 95630.289 | 11.823 |  | 40.5 |  | bb | 2023-11-07
97 | 97 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.4 | 274644.281 | 11201561 | 88288.648 | 12.443 |  | 42.65 |  | bb | 2023-11-07
98 | 98 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.4 | 279947.906 | 11400369 | 88574.773 | 12.642 |  | 43.34 |  | bb | 2023-11-07
99 | 99 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.4 | 300004.281 | 12091701 | 95312.484 | 12.59 |  | 43.16 |  | bb | 2023-11-07
100 | 100 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.4 | 299780.094 | 11869348 | 94841.609 | 12.643 |  | 43.34 |  | bb | 2023-11-07
101 | 101 | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.42 | 49.545 | 2242 | 1.22 | 162.443 |  | 563.34 |  | bb | 2023-11-07
102 | 102 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.4 | 205967.813 | 8300433 | 76076.914 | 10.829 |  | 37.05 |  | bb | 2023-11-07
103 | 103 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.4 | 206321.797 | 8091617 | 76258 | 10.822 |  | 37.02 |  | bb | 2023-11-07
104 | 104 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.4 | 225792.75 | 9374209 | 80176.891 | 11.265 |  | 38.56 |  | bb | 2023-11-07
105 | 105 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.4 | 189770.188 | 7714787 | 65362.707 | 11.613 |  | 39.77 |  | bb | 2023-11-07
106 | 106 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.4 | 188079.766 | 7687524 | 66185.563 | 11.367 |  | 38.91 |  | bb | 2023-11-07
107 | 107 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.4 | 188022.875 | 7440712 | 64907.961 | 11.587 |  | 39.68 |  | bb | 2023-11-07
108 | 108 | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.44 | 120.509 | 2214 | 7.197 | 66.977 |  | 231.95 |  | bb | 2023-11-07
109 | 109 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 4.4 | 57135.223 | 2375849 | 23824.189 | 9.593 | 30.5 | 32.75 | 7.4 | bb | 2023-11-07
110 | 110 | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.42 | 11.766 | 1260 | 1.133 | 41.539 |  | 143.65 |  | bb | 2023-11-07
111 | 111 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.4 | 228470.469 | 9286333 | 78984.391 | 11.57 |  | 39.62 |  | bb | 2023-11-07
112 | 112 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.4 | 145300.688 | 5942828 | 54687.828 | 10.628 |  | 36.35 |  | bb | 2023-11-07
113 | 113 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.4 | 171869.766 | 6608780 | 63527.227 | 10.822 |  | 37.02 |  | bb | 2023-11-07
114 | 114 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.4 | 232343.172 | 9388697 | 81925.484 | 11.344 |  | 38.83 |  | bb | 2023-11-07
115 | 115 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.4 | 238910.734 | 9909917 | 77283.281 | 12.365 |  | 42.38 |  | bb | 2023-11-07
116 | 116 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.4 | 228967.797 | 9634266 | 77620.695 | 11.799 |  | 40.41 |  | bb | 2023-11-07
117 | 117 | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.4 | 73.177 | 2401 | 1.785 | 163.982 |  | 568.68 |  | bb | 2023-11-07
118 | 118 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.4 | 257647.172 | 10769629 | 86903.203 | 11.859 |  | 40.62 |  | bb | 2023-11-07
119 | 119 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.4 | 179824.234 | 6961351 | 64245.207 | 11.196 |  | 38.32 |  | bb | 2023-11-07
120 | 120 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.4 | 246133.094 | 10120721 | 84458.141 | 11.657 |  | 39.92 |  | bb | 2023-11-07
121 | 121 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.4 | 278046.969 | 11063472 | 87830.766 | 12.663 |  | 43.41 |  | bb | 2023-11-07
122 | 122 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.4 | 275360.031 | 10737526 | 82120.75 | 13.412 |  | 46.01 |  | bb | 2023-11-07
123 | 123 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.4 | 245384.563 | 10366516 | 76499.594 | 12.831 |  | 43.99 |  | bb | 2023-11-07
124 | 124 | 20231106_HepCl_ReRun_124 | Blank | Blank |  |  |  | 8.867 | [5 empty] | 2023-11-07
125 | 125 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.4 | 199021.969 | 7933975 | 73346.039 | 10.854 |  | 37.13 |  | bb | 2023-11-07
126 | 126 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.39 | 195954.266 | 7834061 | 68280.641 | 11.479 |  | 39.3 |  | bb | 2023-11-07
127 | 127 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.4 | 209089.75 | 8472446 | 73499.289 | 11.379 |  | 38.96 |  | bb | 2023-11-07
128 | 128 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.4 | 163660.906 | 6955397 | 55749.684 | 11.743 |  | 40.22 |  | bb | 2023-11-07
129 | 129 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.4 | 87596.305 | 3600304 | 30788.695 | 11.38 |  | 38.96 |  | bb | 2023-11-07
130 | 130 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.4 | 144037.141 | 5669520 | 54407.445 | 10.59 |  | 36.21 |  | bb | 2023-11-07
131 | 131 | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.43 | 5.209 | 616 | 0.21 | 99.219 |  | 343.87 |  | bb | 2023-11-07
132 | 132 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 4.4 | 52353.785 | 2100787 | 21559.432 | 9.713 | 30.5 | 33.17 | 8.8 | bb | 2023-11-07
133 | 133 | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.39 | 20.099 | 1631 | 1.004 | 80.076 |  | 277.42 |  | bb | 2023-11-07
134 | 134 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.4 | 198084.219 | 8400724 | 76163.773 | 10.403 |  | 35.57 |  | bb | 2023-11-07
135 | 135 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.4 | 241833.688 | 9651093 | 92251.773 | 10.486 |  | 35.85 |  | bb | 2023-11-07
136 | 136 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.4 | 218327.75 | 8751073 | 83986.945 | 10.398 |  | 35.55 |  | bb | 2023-11-07
137 | 137 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.4 | 238265.75 | 9592995 | 86446.32 | 11.025 |  | 37.73 |  | bb | 2023-11-07
138 | 138 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.4 | 262954.25 | 10368118 | 89926.844 | 11.696 |  | 40.06 |  | bb | 2023-11-07
139 | 139 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.4 | 266459.938 | 11257610 | 91128.219 | 11.696 |  | 40.06 |  | bb | 2023-11-07
140 | 140 | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.43 | 18.612 | 2037 | 3.553 | 20.954 |  | 72.19 |  | bb | 2023-11-07
141 | 141 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 4.4 | 46493.293 | 1956441 | 19427.527 | 9.573 | 30.5 | 32.69 | 7.2 | bb | 2023-11-07
142 | 142 | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.41 | 27.864 | 1977 | 4.462 | 24.979 |  | 86.16 |  | bb | 2023-11-07
143 | 143 | 20231106_HepCl_ReRun_143 | Blank | Blank | [9 empty] | 2023-11-07
144 | 144 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.4 | 27795.627 | 1163048 | 107375.094 | 1.035 |  | 3.05 |  | bb | 2023-11-07
145 | 145 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.4 | 4972.751 | 197217 | 104266.414 | 0.191 |  | 0.12 |  | bb | 2023-11-07
146 | 146 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.4 | 5474.621 | 219701 | 105538.766 | 0.207 |  | 0.18 |  | bb | 2023-11-07
147 | 147 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.4 | 2002.274 | 89927 | 84474.836 | 0.095 |  |  |  | bbI | 2023-11-07
148 | 148 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.4 | 2397.686 | 94151 | 87116.578 | 0.11 |  |  |  | bbI | 2023-11-07
149 | 149 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.4 | 2073.612 | 89065 | 83468.359 | 0.099 |  |  |  | bbI | 2023-11-07
150 | 150 | 20231106_HepCl_ReRun_152 | Blank | Blank |  |  |  | 6.476 | [5 empty] | 2023-11-07
151 | 151 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.4 | 2722.367 | 114256 | 89927.414 | 0.121 |  |  |  | bbI | 2023-11-07
152 | 152 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.4 | 2252.211 | 97347 | 86421.602 | 0.104 |  |  |  | bbI | 2023-11-07
153 | 153 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.39 | 3439.174 | 128733 | 91154.789 | 0.151 |  |  |  | bbI | 2023-11-07
154 | 154 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.4 | 2381.84 | 88870 | 89026.438 | 0.107 |  |  |  | bbI | 2023-11-07
155 | 155 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.4 | 6998.725 | 272006 | 89150.375 | 0.314 |  | 0.55 |  | bb | 2023-11-07
156 | 156 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.4 | 14859.413 | 603578 | 89808.383 | 0.662 |  | 1.75 |  | bb | 2023-11-07
157 | 157 | 20231106_HepCl_ReRun_159 | Blank | Blank |  |  |  | 0.75 | [5 empty] | 2023-11-07
158 | 158 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.4 | 1691.597 | 72046 | 78967.586 | 0.086 |  |  |  | bbI | 2023-11-07
159 | 159 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.4 | 2561.379 | 93938 | 83951.203 | 0.122 |  |  |  | bbI | 2023-11-07
160 | 160 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.39 | 4321.598 | 173179 | 82722.727 | 0.209 |  | 0.18 |  | bb | 2023-11-07
161 | 161 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.4 | 2641.276 | 106059 | 80214.852 | 0.132 |  |  |  | bbI | 2023-11-07
162 | 162 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.4 | 2297.828 | 98469 | 77124.359 | 0.119 |  |  |  | bbI | 2023-11-07
163 | 163 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.4 | 1735.724 | 69695 | 74714.813 | 0.093 |  |  |  | bbI | 2023-11-07
164 | 164 | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.41 | 1.558 | 302 | 3.511 | 1.775 |  | 5.62 |  | bb | 2023-11-07
165 | 165 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 4.4 | 46230.961 | 1938281 | 19198.016 | 9.632 | 30.5 | 32.89 | 7.8 | bb | 2023-11-07
166 | 166 | 20231106_HepCl_ReRun_168 | Blank | Blank | [9 empty] | 2023-11-07
167 | 167 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.4 | 4481.813 | 193533 | 99319.828 | 0.181 |  | 0.08 |  | bb | 2023-11-07
168 | 168 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.4 | 7830.961 | 317929 | 94711.563 | 0.331 |  | 0.6 |  | bb | 2023-11-07
169 | 169 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.4 | 9123.27 | 397031 | 90822.211 | 0.402 |  | 0.85 |  | bb | 2023-11-07
170 | 170 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.4 | 2338.676 | 92326 | 84107.078 | 0.111 |  |  |  | bbI | 2023-11-07
171 | 171 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.4 | 1933.72 | 74930 | 82892.469 | 0.093 |  |  |  | bbI | 2023-11-07
172 | 172 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.4 | 2124.865 | 95781 | 89697.148 | 0.095 |  |  |  | bbI | 2023-11-07
173 | 173 | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.43 | 88.735 | 1658 | 6.395 | 55.503 |  | 192.12 |  | bb | 2023-11-07
174 | 174 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.4 | 2743.245 | 102340 | 84959.609 | 0.129 |  |  |  | bbI | 2023-11-07
175 | 175 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.4 | 3430.022 | 134885 | 86283.852 | 0.159 |  | 0.01 |  | bb | 2023-11-07
176 | 176 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.4 | 2008.583 | 88863 | 100046.359 | 0.08 |  |  |  | bbI | 2023-11-07
177 | 177 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.4 | 14150.011 | 605117 | 95674.383 | 0.592 |  | 1.51 |  | bb | 2023-11-07
178 | 178 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.4 | 6172.626 | 255818 | 93804.273 | 0.263 |  | 0.37 |  | bb | 2023-11-07
179 | 179 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.4 | 5040.559 | 197915 | 75229.672 | 0.268 |  | 0.39 |  | bb | 2023-11-07
180 | 180 | 20231106_HepCl_ReRun_182 | Blank | Blank |  |  |  | 0.43 | [5 empty] | 2023-11-07
181 | 181 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.4 | 2383.649 | 96799 | 87124.805 | 0.109 |  |  |  | bbI | 2023-11-07
182 | 182 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.4 | 1627.43 | 65057 | 77005.422 | 0.085 |  |  |  | bbI | 2023-11-07
183 | 183 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.4 | 3482.942 | 128910 | 88369.641 | 0.158 |  | 0 |  | bb | 2023-11-07
184 | 184 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.4 | 1741.286 | 67916 | 76210.016 | 0.091 |  |  |  | bbI | 2023-11-07
185 | 185 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.4 | 1187.428 | 48088 | 52721.07 | 0.09 |  |  |  | bbI | 2023-11-07
186 | 186 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.4 | 1416.868 | 47558 | 55817.023 | 0.102 |  |  |  | bbI | 2023-11-07
187 | 187 | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.43 | 14.778 | 1378 | 2.904 | 20.355 |  | 70.11 |  | bb | 2023-11-07
188 | 188 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 4.4 | 50770.645 | 1929792 | 21716.965 | 9.351 | 30.5 | 31.92 | 4.6 | bb | 2023-11-07
189 | 189 | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.41 | 3.016 | 787 | 2.254 | 5.352 |  | 18.03 |  | bb | 2023-11-07
190 | 190 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.4 | 2355.618 | 92222 | 99776.836 | 0.094 |  |  |  | bbI | 2023-11-07
191 | 191 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.4 | 3240.659 | 122106 | 101979.219 | 0.127 |  |  |  | bbI | 2023-11-07
192 | 192 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.4 | 16277.225 | 691768 | 114534.242 | 0.568 |  | 1.43 |  | bb | 2023-11-07
193 | 193 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.4 | 3266.747 | 141507 | 79844.398 | 0.164 |  | 0.02 |  | bb | 2023-11-07
194 | 194 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.4 | 3671.382 | 155139 | 86026.953 | 0.171 |  | 0.05 |  | bb | 2023-11-07
195 | 195 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.4 | 3701.034 | 150082 | 90566.922 | 0.163 |  | 0.02 |  | bb | 2023-11-07
196 | 196 | 20231106_HepCl_ReRun_198 | Blank | Blank | 4.35 | 76.1 | 1422 | 0.96 | 317.083 |  | 1100.13 |  | bb | 2023-11-08
197 | 197 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 4.4 | 46421.027 | 1889998 | 19258.826 | 9.642 | 30.5 | 32.92 | 7.9 | bb | 2023-11-08
198 | 198 | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.43 | 185.614 | 2410 | 1.102 | 673.735 |  | 2338.17 |  | bb | 2023-11-08
199 | 199 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.4 | 4395.498 | 171055 | 97325.922 | 0.181 |  | 0.08 |  | bb | 2023-11-08
200 | 200 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.4 | 4938.487 | 182785 | 102662.727 | 0.192 |  | 0.12 |  | bb | 2023-11-08
201 | 201 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.4 | 3017.959 | 122122 | 99524.109 | 0.121 |  |  |  | bbI | 2023-11-08
202 | 202 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.4 | 2407.487 | 92725 | 104638.984 | 0.092 |  |  |  | bbI | 2023-11-08
203 | 203 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.4 | 2357.775 | 96783 | 95722.523 | 0.099 |  |  |  | bbI | 2023-11-08
204 | 204 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.4 | 2659.664 | 107623 | 104475.281 | 0.102 |  |  |  | bbI | 2023-11-08
205 | 205 | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.4 | 24.258 | 1634 | 1.83 | 53.023 |  | 183.51 |  | bb | 2023-11-08
206 | 206 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.4 | 6274.098 | 254249 | 96895.539 | 0.259 |  | 0.35 |  | bb | 2023-11-08
207 | 207 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.4 | 2641.203 | 106814 | 90586.703 | 0.117 |  |  |  | bbI | 2023-11-08
208 | 208 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.4 | 2226.196 | 79075 | 85220.258 | 0.104 |  |  |  | bbI | 2023-11-08
209 | 209 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.4 | 2858.172 | 106053 | 87958.188 | 0.13 |  |  |  | bbI | 2023-11-08
210 | 210 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.4 | 3113.741 | 112647 | 73946.117 | 0.168 |  | 0.04 |  | bb | 2023-11-08
211 | 211 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.4 | 5364.324 | 218230 | 88629.414 | 0.242 |  | 0.3 |  | bb | 2023-11-08
212 | 212 | 20231106_HepCl_ReRun_214 | Blank | Blank |  |  |  | 1.23 | [5 empty] | 2023-11-08
213 | 213 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.4 | 2398.448 | 106629 | 93904.141 | 0.102 |  |  |  | bbI | 2023-11-08
214 | 214 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.4 | 2452.679 | 97232 | 99092.422 | 0.099 |  |  |  | bbI | 2023-11-08
215 | 215 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.4 | 3060.941 | 112753 | 102824.828 | 0.119 |  |  |  | bbI | 2023-11-08
216 | 216 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.4 | 2650.645 | 109312 | 74625.852 | 0.142 |  |  |  | bbI | 2023-11-08
217 | 217 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.4 | 2487.767 | 106960 | 78271.125 | 0.127 |  |  |  | bbI | 2023-11-08
218 | 218 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.4 | 1859.349 | 76758 | 82232.203 | 0.09 |  |  |  | bbI | 2023-11-08
219 | 219 | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.36 | 269.191 | 2916 | 0.419 | 2569.842 |  | 8920.05 |  | bb | 2023-11-08
220 | 220 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 4.4 | 48423.484 | 2028444 | 19805.637 | 9.78 | 30.5 | 33.4 | 9.5 | bb | 2023-11-08
221 | 221 | 20231106_HepCl_ReRun_223 | Blank | Blank |  |  |  | 5.807 | [5 empty] | 2023-11-08
222 | 222 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.4 | 2407.742 | 103954 | 68976.289 | 0.14 |  |  |  | bbI | 2023-11-08
223 | 223 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.4 | 1676.365 | 65562 | 80912.125 | 0.083 |  |  |  | bbI | 2023-11-08
224 | 224 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.4 | 1451.845 | 55029 | 64799.211 | 0.09 |  |  |  | bbI | 2023-11-08
225 | 225 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.4 | 2270.83 | 89051 | 95344.883 | 0.095 |  |  |  | bbI | 2023-11-08
226 | 226 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.4 | 2288.021 | 85144 | 98936.914 | 0.093 |  |  |  | bbI | 2023-11-08
227 | 227 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.4 | 2269.482 | 96527 | 96992.188 | 0.094 |  |  |  | bbI | 2023-11-08
228 | 228 | 20231106_HepCl_ReRun_230 | Blank | Blank |  |  |  | 2.226 | [5 empty] | 2023-11-08
229 | 229 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.4 | 2530.248 | 100776 | 108030.414 | 0.094 |  |  |  | bbI | 2023-11-08
230 | 230 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.4 | 3093.332 | 115462 | 100591.141 | 0.123 |  |  |  | bbI | 2023-11-08
231 | 231 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.4 | 1622.044 | 59654 | 69586.211 | 0.093 |  |  |  | bbI | 2023-11-08
232 | 232 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.4 | 121838.508 | 4179704 | 104467.281 | 4.665 |  | 15.65 |  | bb | 2023-11-08
233 | 233 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.4 | 4144.071 | 183465 | 48669.789 | 0.341 |  | 0.64 |  | bb | 2023-11-08
234 | 234 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.4 | 34674.074 | 1378169 | 39606.652 | 3.502 |  | 11.61 |  | bb | 2023-11-08
235 | 235 | 20231106_HepCl_ReRun_237 | Blank | Blank |  |  |  | 8.198 | [5 empty] | 2023-11-08
236 | 236 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.4 | 2658.295 | 114635 | 100054.281 | 0.106 |  |  |  | bbI | 2023-11-08
237 | 237 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.4 | 2295.008 | 85333 | 91328.695 | 0.101 |  |  |  | bbI | 2023-11-08
238 | 238 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.4 | 2382.871 | 91499 | 102183.391 | 0.093 |  |  |  | bbI | 2023-11-08
239 | 239 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.4 | 2309.284 | 84392 | 95108.883 | 0.097 |  |  |  | bbI | 2023-11-08
240 | 240 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.4 | 1318.323 | 47264 | 53921.77 | 0.098 |  |  |  | bbI | 2023-11-08
241 | 241 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.4 | 1408.505 | 61221 | 61000.227 | 0.092 |  |  |  | bbI | 2023-11-08
242 | 242 | 20231106_HepCl_ReRun_244 | Blank | Blank |  |  |  | 7.8 | [5 empty] | 2023-11-08
243 | 243 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 4.4 | 50404.621 | 1991044 | 20601.318 | 9.787 | 30.5 | 33.43 | 9.6 | bb | 2023-11-08
244 | 244 | 20231106_HepCl_ReRun_246 | Blank | Blank |  |  |  | 4.036 | [5 empty] | 2023-11-08
245 | 245 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.4 | 6702.993 | 251062 | 180164.406 | 0.149 |  |  |  | bbI | 2023-11-08
246 | 246 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.4 | 3526.294 | 137227 | 128731.063 | 0.11 |  |  |  | bbI | 2023-11-08
247 | 247 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.4 | 3763.688 | 148501 | 119917.117 | 0.126 |  |  |  | bbI | 2023-11-08
248 | 248 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.4 | 9193.139 | 358467 | 88866.891 | 0.414 |  | 0.89 |  | bb | 2023-11-08
249 | 249 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.4 | 9416.714 | 361946 | 88157.313 | 0.427 |  | 0.94 |  | bb | 2023-11-08
250 | 250 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.4 | 10473.873 | 423184 | 98938.063 | 0.423 |  | 0.93 |  | bb | 2023-11-08
251 | 251 | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.43 | 12.221 | 1633 | 1.238 | 39.486 |  | 136.52 |  | bb | 2023-11-08
252 | 252 | 20231106_HepCl_ReRun_254 | Blank | Blank |  |  |  | 7.609 | [5 empty] | 2023-11-08
253 | 253 | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.43 | 36.339 | 1486 | 3 | 48.452 |  | 167.65 |  | bb | 2023-11-08
Compound 10: M7PFUdA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20231106_HepCl_ReRun_001 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
2 | 2 | 20231106_HepCl_ReRun_002 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
3 | 3 | 20231106_HepCl_ReRun_003 | Blank | Blank | 4.39 | 6.982 | 235 |  | 6.982 | 4 | 0 | -100 | bb | 2023-11-06
4 | 4 | 20231106_HepCl_ReRun_004 | Matrix Blank | Blank | 4.41 | 5.073 | 225 |  | 5.073 | 4 | 0 | -100 | bb | 2023-11-06
5 | 5 | 20231106_HepCl_ReRun_005 | Crash Blank | Blank | 4.38 | 3.127 | 303 |  | 3.127 | 4 | 0 | -100 | bb | 2023-11-06
6 | 6 | 20231106_HepCl_ReRun_006 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
7 | 7 | 20231106_HepCl_ReRun_007 | Blank | Blank | 4.4 | 7.306 | 320 |  | 7.306 | 4 | 0 | -100 | bb | 2023-11-06
8 | 8 | 20231106_HepCl_ReRun_008 | CC01 | Analyte | 4.4 | 49528.125 | 2117331 |  | 49528.125 | 4 | 5.54 | 38.5 | bb | 2023-11-06
9 | 9 | 20231106_HepCl_ReRun_009 | CC02 | Analyte | 4.4 | 47045.578 | 1943068 |  | 47045.578 | 4 | 5.26 | 31.6 | bb | 2023-11-06
10 | 10 | 20231106_HepCl_ReRun_010 | CC03 | Analyte | 4.4 | 40876.117 | 1656387 |  | 40876.117 | 4 | 4.57 | 14.3 | bb | 2023-11-06
11 | 11 | 20231106_HepCl_ReRun_011 | CC04 | Analyte | 4.4 | 45863.699 | 1923637 |  | 45863.699 | 4 | 5.13 | 28.3 | bb | 2023-11-06
12 | 12 | 20231106_HepCl_ReRun_012 | CC05 | Analyte | 4.4 | 43603.629 | 1755856 |  | 43603.629 | 4 | 4.88 | 21.9 | bb | 2023-11-06
13 | 13 | 20231106_HepCl_ReRun_013 | CC06 | Standard | 4.4 | 40008.617 | 1617090 |  | 40008.617 | 4 | 4.48 | 11.9 | bb | 2023-11-06
14 | 14 | 20231106_HepCl_ReRun_014 | Blank | Blank | 4.37 | 0.639 | 71 |  | 0.639 | 4 | 0 | -100 | bb | 2023-11-06
15 | 15 | 20231106_HepCl_ReRun_015 | CC07 | Standard | 4.4 | 42977.938 | 1740176 |  | 42977.938 | 4 | 4.81 | 20.2 | bb | 2023-11-06
16 | 16 | 20231106_HepCl_ReRun_016 | CC08 | Standard | 4.4 | 44416.688 | 1823227 |  | 44416.688 | 4 | 4.97 | 24.2 | bb | 2023-11-06
17 | 17 | 20231106_HepCl_ReRun_017 | CC09 | Standard | 4.4 | 41387.938 | 1671639 |  | 41387.938 | 4 | 4.63 | 15.7 | bb | 2023-11-06
18 | 18 | 20231106_HepCl_ReRun_018 | CC10 | Standard | 4.4 | 43287.93 | 1755786 |  | 43287.93 | 4 | 4.84 | 21.1 | bb | 2023-11-06
19 | 19 | 20231106_HepCl_ReRun_019 | CC11 | Standard | 4.4 | 43650.934 | 1729358 |  | 43650.934 | 4 | 4.88 | 22.1 | bb | 2023-11-06
20 | 20 | 20231106_HepCl_ReRun_020 | CC12 | Standard | 4.4 | 43166.441 | 1736996 |  | 43166.441 | 4 | 4.83 | 20.7 | bb | 2023-11-06
21 | 21 | 20231106_HepCl_ReRun_021 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-06
22 | 22 | 20231106_HepCl_ReRun_022 | CC13 | Standard | 4.4 | 44292.91 | 1771998 |  | 44292.91 | 4 | 4.95 | 23.9 | bb | 2023-11-06
23 | 23 | 20231106_HepCl_ReRun_023 | CC14 | Standard | 4.4 | 50112.5 | 1895970 |  | 50112.5 | 4 | 5.61 | 40.1 | bb | 2023-11-06
24 | 24 | 20231106_HepCl_ReRun_024 | CC15 | Standard | 4.4 | 38008.148 | 1745132 |  | 38008.148 | 4 | 4.25 | 6.3 | MM | 2023-11-06
25 | 25 | 20231106_HepCl_ReRun_025 | CC16 | Analyte | 4.4 | 52133.781 | 1791967 |  | 52133.781 | 4 | 5.83 | 45.8 | bb | 2023-11-06
26 | 26 | 20231106_HepCl_ReRun_026 | CC17 | Analyte | 4.4 | 43163.68 | 1662867 |  | 43163.68 | 4 | 4.83 | 20.7 | bb | 2023-11-06
27 | 27 | 20231106_HepCl_ReRun_027 | Matrix Blank | Blank | 4.41 | 16.466 | 203 |  | 16.466 | 4 | 0 | -100 | bb | 2023-11-06
28 | 28 | 20231106_HepCl_ReRun_028 | Crash Blank | Blank | 4.39 | 4.24 | 142 |  | 4.24 | 4 | 0 | -100 | bb | 2023-11-06
29 | 29 | 20231106_HepCl_ReRun_029 | CC04 | Analyte | 4.4 | 38003.336 | 1571169 |  | 38003.336 | 4 | 4.25 | 6.3 | bb | 2023-11-06
30 | 30 | 20231106_HepCl_ReRun_030 | CC05 | Analyte | 4.4 | 35276.152 | 1423350 |  | 35276.152 | 4 | 3.95 | -1.3 | bb | 2023-11-06
31 | 31 | 20231106_HepCl_ReRun_031 | CC06 | Standard | 4.4 | 33664.371 | 1373684 |  | 33664.371 | 4 | 3.77 | -5.9 | bb | 2023-11-06
32 | 32 | 20231106_HepCl_ReRun_032 | Blank | Blank | 4.35 | 4.114 | 195 |  | 4.114 | 4 | 0 | -100 | bb | 2023-11-06
33 | 33 | 20231106_HepCl_ReRun_033 | CC04 | Analyte | 4.4 | 37383.543 | 1509787 |  | 37383.543 | 4 | 4.18 | 4.5 | bb | 2023-11-06
34 | 34 | 20231106_HepCl_ReRun_034 | CC05 | Analyte | 4.4 | 33432.625 | 1300150 |  | 33432.625 | 4 | 3.74 | -6.5 | bb | 2023-11-06
35 | 35 | 20231106_HepCl_ReRun_035 | CC06 | Standard | 4.4 | 31861.572 | 1318096 |  | 31861.572 | 4 | 3.56 | -10.9 | bb | 2023-11-06
36 | 36 | 20231106_HepCl_ReRun_036 | Blank | Blank | 4.41 | 7.725 | 173 |  | 7.725 | 4 | 0 | -100 | bb | 2023-11-06
37 | 37 | 20231106_HepCl_ReRun_037 | CC04 | Analyte | 4.39 | 34334.141 | 1351660 |  | 34334.141 | 4 | 3.84 | -4 | bb | 2023-11-06
38 | 38 | 20231106_HepCl_ReRun_038 | CC05 | Analyte | 4.39 | 32710.328 | 1335617 |  | 32710.328 | 4 | 3.66 | -8.5 | bb | 2023-11-06
39 | 39 | 20231106_HepCl_ReRun_039 | CC06 | Standard | 4.4 | 30194.342 | 1255627 |  | 30194.342 | 4 | 3.38 | -15.6 | bb | 2023-11-06
40 | 40 | 20231106_HepCl_ReRun_040 | Blank | Blank | 4.4 | 5.681 | 201 |  | 5.681 | 4 | 0 | -100 | bb | 2023-11-06
41 | 41 | 20231106_HepCl_ReRun_041 | QC A | QC | 4.4 | 49355.238 | 1984290 |  | 49355.238 | 4 | 5.52 | 38 | bb | 2023-11-06
42 | 42 | 20231106_HepCl_ReRun_042 | QC B | QC | 4.4 | 51642.734 | 2142571 |  | 51642.734 | 4 | 5.78 | 44.4 | bb | 2023-11-06
43 | 43 | 20231106_HepCl_ReRun_043 | QC C | QC | 4.4 | 53094.781 | 2272471 |  | 53094.781 | 4 | 5.94 | 48.5 | bb | 2023-11-06
44 | 44 | 20231106_HepCl_ReRun_044 | QC D | QC | 4.4 | 46975.297 | 1959518 |  | 46975.297 | 4 | 5.25 | 31.4 | bb | 2023-11-06
45 | 45 | 20231106_HepCl_ReRun_045 | Matrix Blank | Blank | 4.43 | 9.694 | 259 |  | 9.694 | 4 | 0 | -100 | bb | 2023-11-06
46 | 46 | 20231106_HepCl_ReRun_046 | Crash Blank | Blank | 4.42 | 1.228 | 138 |  | 1.228 | 4 | 0 | -100 | bb | 2023-11-06
47 | 47 | 20231106_HepCl_ReRun_047 | CC01 | Analyte | 4.4 | 31262.33 | 1253093 |  | 31262.33 | 4 | 3.5 | -12.6 | bb | 2023-11-06
48 | 48 | 20231106_HepCl_ReRun_048 | CC02 | Analyte | 4.4 | 31310.842 | 1311558 |  | 31310.842 | 4 | 3.5 | -12.4 | bb | 2023-11-06
49 | 49 | 20231106_HepCl_ReRun_049 | CC03 | Analyte | 4.4 | 28202.996 | 1139326 |  | 28202.996 | 4 | 3.15 | -21.1 | bb | 2023-11-06
50 | 50 | 20231106_HepCl_ReRun_050 | CC04 | Analyte | 4.4 | 29024.73 | 1111171 |  | 29024.73 | 4 | 3.25 | -18.8 | bb | 2023-11-06
51 | 51 | 20231106_HepCl_ReRun_051 | CC05 | Analyte | 4.4 | 28126.109 | 1122780 |  | 28126.109 | 4 | 3.15 | -21.3 | bb | 2023-11-06
52 | 52 | 20231106_HepCl_ReRun_052 | CC06 | Standard | 4.4 | 27641.32 | 1122443 |  | 27641.32 | 4 | 3.09 | -22.7 | bb | 2023-11-06
53 | 53 | 20231106_HepCl_ReRun_053 | Blank | Blank | 4.39 | 3.035 | 208 |  | 3.035 | 4 | 0 | -100 | bb | 2023-11-06
54 | 54 | 20231106_HepCl_ReRun_054 | CC07 | Standard | 4.39 | 30703.615 | 1272896 |  | 30703.615 | 4 | 3.43 | -14.1 | bb | 2023-11-06
55 | 55 | 20231106_HepCl_ReRun_055 | CC08 | Standard | 4.4 | 30452.596 | 1229852 |  | 30452.596 | 4 | 3.41 | -14.8 | bb | 2023-11-06
56 | 56 | 20231106_HepCl_ReRun_056 | CC09 | Standard | 4.4 | 29372.057 | 1219939 |  | 29372.057 | 4 | 3.29 | -17.9 | bb | 2023-11-06
57 | 57 | 20231106_HepCl_ReRun_057 | CC10 | Standard | 4.4 | 28786.756 | 1209493 |  | 28786.756 | 4 | 3.22 | -19.5 | bb | 2023-11-06
58 | 58 | 20231106_HepCl_ReRun_058 | CC11 | Standard | 4.4 | 31747.754 | 1296720 |  | 31747.754 | 4 | 3.55 | -11.2 | bb | 2023-11-06
59 | 59 | 20231106_HepCl_ReRun_059 | CC12 | Standard | 4.4 | 31395.316 | 1232338 |  | 31395.316 | 4 | 3.51 | -12.2 | bb | 2023-11-06
60 | 60 | 20231106_HepCl_ReRun_060 | Blank | Blank | 4.41 | 2.979 | 184 |  | 2.979 | 4 | 0 | -100 | bb | 2023-11-06
61 | 61 | 20231106_HepCl_ReRun_061 | CC13 | Standard | 4.4 | 34492.879 | 1450823 |  | 34492.879 | 4 | 3.86 | -3.5 | bb | 2023-11-06
62 | 62 | 20231106_HepCl_ReRun_062 | CC14 | Standard | 4.4 | 41397.715 | 1662759 |  | 41397.715 | 4 | 4.63 | 15.8 | bb | 2023-11-06
63 | 63 | 20231106_HepCl_ReRun_063 | CC15 | Standard | 4.4 | 38112.547 | 1543227 |  | 38112.547 | 4 | 4.26 | 6.6 | bb | 2023-11-06
64 | 64 | 20231106_HepCl_ReRun_064 | CC16 | Analyte | 4.4 | 57215.652 | 2313839 |  | 57215.652 | 4 | 6.4 | 60 | bb | 2023-11-06
65 | 65 | 20231106_HepCl_ReRun_065 | CC17 | Analyte | 4.4 | 38809.316 | 1505012 |  | 38809.316 | 4 | 4.34 | 8.5 | bb | 2023-11-07
66 | 66 | 20231106_HepCl_ReRun_066 | Matrix Blank | Blank | 4.37 | 6.959 | 192 |  | 6.959 | 4 | 0 | -100 | bb | 2023-11-07
67 | 67 | 20231106_HepCl_ReRun_067 | Crash Blank | Blank | 4.39 | 1.469 | 120 |  | 1.469 | 4 | 0 | -100 | bb | 2023-11-07
68 | 68 | 20231106_HepCl_ReRun_068 | CC04 | Analyte | 4.4 | 26842.035 | 1085838 |  | 26842.035 | 4 | 3 | -24.9 | bb | 2023-11-07
69 | 69 | 20231106_HepCl_ReRun_069 | CC05 | Analyte | 4.4 | 25787.34 | 1064513 |  | 25787.34 | 4 | 2.88 | -27.9 | bb | 2023-11-07
70 | 70 | 20231106_HepCl_ReRun_070 | CC06 | Standard | 4.4 | 26664.004 | 1080929 |  | 26664.004 | 4 | 2.98 | -25.4 | bb | 2023-11-07
71 | 71 | 20231106_HepCl_ReRun_071 | Blank | Blank | 4.37 | 13.529 | 246 |  | 13.529 | 4 | 0 | -100 | bb | 2023-11-07
72 | 72 | 20231106_HepCl_ReRun_072 | CC04 | Analyte | 4.4 | 25236.123 | 989841 |  | 25236.123 | 4 | 2.82 | -29.4 | bb | 2023-11-07
73 | 73 | 20231106_HepCl_ReRun_073 | CC05 | Analyte | 4.4 | 25884.043 | 1063068 |  | 25884.043 | 4 | 2.9 | -27.6 | bb | 2023-11-07
74 | 74 | 20231106_HepCl_ReRun_074 | CC06 | Standard | 4.4 | 27006.176 | 1109378 |  | 27006.176 | 4 | 3.02 | -24.5 | bb | 2023-11-07
75 | 75 | 20231106_HepCl_ReRun_075 | Blank | Blank | 4.4 | 8.646 | 340 |  | 8.646 | 4 | 0 | -100 | bb | 2023-11-07
76 | 76 | 20231106_HepCl_ReRun_076 | CC04 | Analyte | 4.4 | 24837.654 | 1011861 |  | 24837.654 | 4 | 2.78 | -30.5 | bb | 2023-11-07
77 | 77 | 20231106_HepCl_ReRun_077 | CC05 | Analyte | 4.4 | 23961.09 | 987606 |  | 23961.09 | 4 | 2.68 | -33 | bb | 2023-11-07
78 | 78 | 20231106_HepCl_ReRun_078 | CC06 | Standard | 4.4 | 24890.996 | 992358 |  | 24890.996 | 4 | 2.78 | -30.4 | bb | 2023-11-07
79 | 79 | 20231106_HepCl_ReRun_079 | Blank | Blank | 4.36 | 10.156 | 202 |  | 10.156 | 4 | 0 | -100 | bb | 2023-11-07
80 | 80 | 20231106_HepCl_ReRun_080 | QC A | QC | 4.4 | 36659.156 | 1533681 |  | 36659.156 | 4 | 4.1 | 2.5 | bb | 2023-11-07
81 | 81 | 20231106_HepCl_ReRun_081 | QC B | QC | 4.4 | 37154.508 | 1560233 |  | 37154.508 | 4 | 4.16 | 3.9 | bb | 2023-11-07
82 | 82 | 20231106_HepCl_ReRun_082 | QC C | QC | 4.4 | 41341.777 | 1699979 |  | 41341.777 | 4 | 4.62 | 15.6 | bb | 2023-11-07
83 | 83 | 20231106_HepCl_ReRun_083 | QC D | QC | 4.4 | 32931.387 | 1374550 |  | 32931.387 | 4 | 3.68 | -7.9 | bb | 2023-11-07
84 | 84 | 20231106_HepCl_ReRun_084 | Blank | Blank | 4.4 | 4.803 | 181 |  | 4.803 | 4 | 0 | -100 | bb | 2023-11-07
85 | 85 | 20231106_HepCl_ReRun_085 | Blank | Blank | 4.42 | 0.242 | 58 |  | 0.242 | 4 | 0 | -100 | bb | 2023-11-07
86 | 86 | 20231106_HepCl_ReRun_086 | CC12 | Analyte | 4.4 | 25853.348 | 1085943 |  | 25853.348 | 4 | 2.89 | -27.7 | bb | 2023-11-07
87 | 87 | 20231106_HepCl_ReRun_087 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
88 | 88 | 20231106_HepCl_ReRun_088 | Mix 1 Active T0 A | Analyte | 4.4 | 98002.758 | 4043580 |  | 98002.758 | 4 | 10.96 | 174.1 | bb | 2023-11-07
89 | 89 | 20231106_HepCl_ReRun_089 | Mix 1 Active T0 B | Analyte | 4.4 | 88635.797 | 3597215 |  | 88635.797 | 4 | 9.92 | 147.9 | bb | 2023-11-07
90 | 90 | 20231106_HepCl_ReRun_090 | Mix 1 Active T0 C | Analyte | 4.4 | 89190.406 | 3620419 |  | 89190.406 | 4 | 9.98 | 149.4 | bb | 2023-11-07
91 | 91 | 20231106_HepCl_ReRun_091 | Mix 1 Active T15 A | Analyte | 4.4 | 89103.305 | 3621246 |  | 89103.305 | 4 | 9.97 | 149.2 | bb | 2023-11-07
92 | 92 | 20231106_HepCl_ReRun_092 | Mix 1 Active T15 B | Analyte | 4.4 | 83877.867 | 3228451 |  | 83877.867 | 4 | 9.38 | 134.6 | bb | 2023-11-07
93 | 93 | 20231106_HepCl_ReRun_093 | Mix 1 Active T15 C | Analyte | 4.4 | 86799.516 | 3602688 |  | 86799.516 | 4 | 9.71 | 142.7 | bb | 2023-11-07
94 | 94 | 20231106_HepCl_ReRun_094 | Blank | Blank | 4.4 | 0.119 | 22 |  | 0.119 | 4 | 0 | -100 | bb | 2023-11-07
95 | 95 | 20231106_HepCl_ReRun_095 | Mix 1 Active T30 A | Analyte | 4.4 | 93797.602 | 3755169 |  | 93797.602 | 4 | 10.49 | 162.3 | bb | 2023-11-07
96 | 96 | 20231106_HepCl_ReRun_096 | Mix 1 Active T30 B | Analyte | 4.4 | 95630.289 | 3956639 |  | 95630.289 | 4 | 10.7 | 167.4 | bb | 2023-11-07
97 | 97 | 20231106_HepCl_ReRun_097 | Mix 1 Active T30 C | Analyte | 4.4 | 88288.648 | 3610306 |  | 88288.648 | 4 | 9.88 | 146.9 | bb | 2023-11-07
98 | 98 | 20231106_HepCl_ReRun_098 | Mix 1 Active T60 A | Analyte | 4.4 | 88574.773 | 3453310 |  | 88574.773 | 4 | 9.91 | 147.7 | bb | 2023-11-07
99 | 99 | 20231106_HepCl_ReRun_099 | Mix 1 Active T60 B | Analyte | 4.4 | 95312.484 | 3746577 |  | 95312.484 | 4 | 10.66 | 166.6 | bb | 2023-11-07
100 | 100 | 20231106_HepCl_ReRun_100 | Mix 1 Active T60 C | Analyte | 4.4 | 94841.609 | 3775499 |  | 94841.609 | 4 | 10.61 | 165.2 | bb | 2023-11-07
101 | 101 | 20231106_HepCl_ReRun_101 | Blank | Blank | 4.39 | 1.22 | 126 |  | 1.22 | 4 | 0 | -100 | bb | 2023-11-07
102 | 102 | 20231106_HepCl_ReRun_102 | Mix 1 Active T120 A | Analyte | 4.4 | 76076.914 | 2958757 |  | 76076.914 | 4 | 8.51 | 112.8 | bb | 2023-11-07
103 | 103 | 20231106_HepCl_ReRun_103 | Mix 1 Active T120 B | Analyte | 4.4 | 76258 | 2982523 |  | 76258 | 4 | 8.53 | 113.3 | bb | 2023-11-07
104 | 104 | 20231106_HepCl_ReRun_104 | Mix 1 Active T120 C | Analyte | 4.4 | 80176.891 | 3248475 |  | 80176.891 | 4 | 8.97 | 124.2 | bb | 2023-11-07
105 | 105 | 20231106_HepCl_ReRun_105 | Mix 1 Active T240 A | Analyte | 4.4 | 65362.707 | 2663057 |  | 65362.707 | 4 | 7.31 | 82.8 | bb | 2023-11-07
106 | 106 | 20231106_HepCl_ReRun_106 | Mix 1 Active T240 B | Analyte | 4.4 | 66185.563 | 2694615 |  | 66185.563 | 4 | 7.4 | 85.1 | bb | 2023-11-07
107 | 107 | 20231106_HepCl_ReRun_107 | Mix 1 Active T240 C | Analyte | 4.4 | 64907.961 | 2536898 |  | 64907.961 | 4 | 7.26 | 81.5 | bb | 2023-11-07
108 | 108 | 20231106_HepCl_ReRun_108 | Blank | Blank | 4.41 | 7.197 | 270 |  | 7.197 | 4 | 0 | -100 | bb | 2023-11-07
109 | 109 | 20231106_HepCl_ReRun_109 | CC12 | Analyte | 4.4 | 23824.189 | 1007547 |  | 23824.189 | 4 | 2.67 | -33.4 | bb | 2023-11-07
110 | 110 | 20231106_HepCl_ReRun_110 | Blank | Blank | 4.38 | 1.133 | 113 |  | 1.133 | 4 | 0 | -100 | bb | 2023-11-07
111 | 111 | 20231106_HepCl_ReRun_111 | Mix 1 Media T0 A | Analyte | 4.4 | 78984.391 | 3150051 |  | 78984.391 | 4 | 8.84 | 120.9 | bb | 2023-11-07
112 | 112 | 20231106_HepCl_ReRun_112 | Mix 1 Media T0 B | Analyte | 4.4 | 54687.828 | 2268070 |  | 54687.828 | 4 | 6.12 | 52.9 | bb | 2023-11-07
113 | 113 | 20231106_HepCl_ReRun_113 | Mix 1 Media T0 C | Analyte | 4.4 | 63527.227 | 2347999 |  | 63527.227 | 4 | 7.11 | 77.7 | bb | 2023-11-07
114 | 114 | 20231106_HepCl_ReRun_114 | Mix 1 Media T15 A | Analyte | 4.4 | 81925.484 | 3419540 |  | 81925.484 | 4 | 9.16 | 129.1 | bb | 2023-11-07
115 | 115 | 20231106_HepCl_ReRun_115 | Mix 1 Media T15 B | Analyte | 4.4 | 77283.281 | 3151650 |  | 77283.281 | 4 | 8.65 | 116.1 | bb | 2023-11-07
116 | 116 | 20231106_HepCl_ReRun_116 | Mix 1 Media T15 C | Analyte | 4.4 | 77620.695 | 3186330 |  | 77620.695 | 4 | 8.68 | 117.1 | bb | 2023-11-07
117 | 117 | 20231106_HepCl_ReRun_117 | Blank | Blank | 4.35 | 1.785 | 169 |  | 1.785 | 4 | 0 | -100 | bb | 2023-11-07
118 | 118 | 20231106_HepCl_ReRun_118 | Mix 1 Media T30 A | Analyte | 4.4 | 86903.203 | 3634190 |  | 86903.203 | 4 | 9.72 | 143 | bb | 2023-11-07
119 | 119 | 20231106_HepCl_ReRun_119 | Mix 1 Media T30 B | Analyte | 4.4 | 64245.207 | 2644733 |  | 64245.207 | 4 | 7.19 | 79.7 | bb | 2023-11-07
120 | 120 | 20231106_HepCl_ReRun_120 | Mix 1 Media T30 C | Analyte | 4.4 | 84458.141 | 3363683 |  | 84458.141 | 4 | 9.45 | 136.2 | bb | 2023-11-07
121 | 121 | 20231106_HepCl_ReRun_121 | Mix 1 Media T60 A | Analyte | 4.4 | 87830.766 | 3422934 |  | 87830.766 | 4 | 9.83 | 145.6 | bb | 2023-11-07
122 | 122 | 20231106_HepCl_ReRun_122 | Mix 1 Media T60 B | Analyte | 4.4 | 82120.75 | 3182716 |  | 82120.75 | 4 | 9.19 | 129.7 | bb | 2023-11-07
123 | 123 | 20231106_HepCl_ReRun_123 | Mix 1 Media T60 C | Analyte | 4.4 | 76499.594 | 3151665 |  | 76499.594 | 4 | 8.56 | 113.9 | bb | 2023-11-07
124 | 124 | 20231106_HepCl_ReRun_124 | Blank | Blank | 4.39 | 8.867 | 284 |  | 8.867 | 4 | 0 | -100 | bb | 2023-11-07
125 | 125 | 20231106_HepCl_ReRun_125 | Mix 1 Media T120 A | Analyte | 4.4 | 73346.039 | 2836304 |  | 73346.039 | 4 | 8.2 | 105.1 | bb | 2023-11-07
126 | 126 | 20231106_HepCl_ReRun_126 | Mix 1 Media T120 B | Analyte | 4.4 | 68280.641 | 2797197 |  | 68280.641 | 4 | 7.64 | 91 | bb | 2023-11-07
127 | 127 | 20231106_HepCl_ReRun_127 | Mix 1 Media T120 C | Analyte | 4.4 | 73499.289 | 2960621 |  | 73499.289 | 4 | 8.22 | 105.5 | bb | 2023-11-07
128 | 128 | 20231106_HepCl_ReRun_128 | Mix 1 Media T240 A | Analyte | 4.4 | 55749.684 | 2264068 |  | 55749.684 | 4 | 6.24 | 55.9 | bb | 2023-11-07
129 | 129 | 20231106_HepCl_ReRun_129 | Mix 1 Media T240 B | Analyte | 4.4 | 30788.695 | 1220278 |  | 30788.695 | 4 | 3.44 | -13.9 | bb | 2023-11-07
130 | 130 | 20231106_HepCl_ReRun_130 | Mix 1 Media T240 C | Analyte | 4.4 | 54407.445 | 2019079 |  | 54407.445 | 4 | 6.09 | 52.2 | bb | 2023-11-07
131 | 131 | 20231106_HepCl_ReRun_131 | Blank | Blank | 4.44 | 0.21 | 68 |  | 0.21 | 4 | 0 | -100 | bb | 2023-11-07
132 | 132 | 20231106_HepCl_ReRun_132 | CC12 | Analyte | 4.4 | 21559.432 | 889577 |  | 21559.432 | 4 | 2.41 | -39.7 | bb | 2023-11-07
133 | 133 | 20231106_HepCl_ReRun_133 | Blank | Blank | 4.38 | 1.004 | 93 |  | 1.004 | 4 | 0 | -100 | bb | 2023-11-07
134 | 134 | 20231106_HepCl_ReRun_134 | Mix 1 Inactive T0 A | Analyte | 4.4 | 76163.773 | 3220525 |  | 76163.773 | 4 | 8.52 | 113 | bb | 2023-11-07
135 | 135 | 20231106_HepCl_ReRun_135 | Mix 1 Inactive T0 B | Analyte | 4.4 | 92251.773 | 3670430 |  | 92251.773 | 4 | 10.32 | 158 | bb | 2023-11-07
136 | 136 | 20231106_HepCl_ReRun_136 | Mix 1 Inactive T0 C | Analyte | 4.4 | 83986.945 | 3315792 |  | 83986.945 | 4 | 9.4 | 134.9 | bb | 2023-11-07
137 | 137 | 20231106_HepCl_ReRun_137 | Mix 1 Inactive T240 A | Analyte | 4.4 | 86446.32 | 3432628 |  | 86446.32 | 4 | 9.67 | 141.8 | bb | 2023-11-07
138 | 138 | 20231106_HepCl_ReRun_138 | Mix 1 Inactive T240 B | Analyte | 4.4 | 89926.844 | 3532382 |  | 89926.844 | 4 | 10.06 | 151.5 | bb | 2023-11-07
139 | 139 | 20231106_HepCl_ReRun_139 | Mix 1 Inactive T240 C | Analyte | 4.4 | 91128.219 | 3617198 |  | 91128.219 | 4 | 10.19 | 154.9 | bb | 2023-11-07
140 | 140 | 20231106_HepCl_ReRun_140 | Blank | Blank | 4.4 | 3.553 | 153 |  | 3.553 | 4 | 0 | -100 | bb | 2023-11-07
141 | 141 | 20231106_HepCl_ReRun_141 | CC12 | Analyte | 4.4 | 19427.527 | 791137 |  | 19427.527 | 4 | 2.17 | -45.7 | bb | 2023-11-07
142 | 142 | 20231106_HepCl_ReRun_142 | Blank | Blank | 4.39 | 4.462 | 185 |  | 4.462 | 4 | 0 | -100 | bb | 2023-11-07
143 | 143 | 20231106_HepCl_ReRun_143 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
144 | 144 | 20231106_HepCl_ReRun_146 | Mix 2 Active T0 A | Analyte | 4.4 | 107375.094 | 4487870 |  | 107375.094 | 4 | 12.01 | 200.3 | bb | 2023-11-07
145 | 145 | 20231106_HepCl_ReRun_147 | Mix 2 Active T0 B | Analyte | 4.4 | 104266.414 | 4273133 |  | 104266.414 | 4 | 11.66 | 191.6 | bb | 2023-11-07
146 | 146 | 20231106_HepCl_ReRun_148 | Mix 2 Active T0 C | Analyte | 4.4 | 105538.766 | 4219927 |  | 105538.766 | 4 | 11.81 | 195.2 | bb | 2023-11-07
147 | 147 | 20231106_HepCl_ReRun_149 | Mix 2 Active T15 A | Analyte | 4.4 | 84474.836 | 3404028 |  | 84474.836 | 4 | 9.45 | 136.2 | bb | 2023-11-07
148 | 148 | 20231106_HepCl_ReRun_150 | Mix 2 Active T15 B | Analyte | 4.4 | 87116.578 | 3474043 |  | 87116.578 | 4 | 9.75 | 143.6 | bb | 2023-11-07
149 | 149 | 20231106_HepCl_ReRun_151 | Mix 2 Active T15 C | Analyte | 4.4 | 83468.359 | 3439578 |  | 83468.359 | 4 | 9.34 | 133.4 | bb | 2023-11-07
150 | 150 | 20231106_HepCl_ReRun_152 | Blank | Blank | 4.4 | 6.476 | 159 |  | 6.476 | 4 | 0 | -100 | bb | 2023-11-07
151 | 151 | 20231106_HepCl_ReRun_153 | Mix 2 Active T30 A | Analyte | 4.4 | 89927.414 | 3486817 |  | 89927.414 | 4 | 10.06 | 151.5 | bb | 2023-11-07
152 | 152 | 20231106_HepCl_ReRun_154 | Mix 2 Active T30 B | Analyte | 4.4 | 86421.602 | 3416719 |  | 86421.602 | 4 | 9.67 | 141.7 | bb | 2023-11-07
153 | 153 | 20231106_HepCl_ReRun_155 | Mix 2 Active T30 C | Analyte | 4.4 | 91154.789 | 3585221 |  | 91154.789 | 4 | 10.2 | 154.9 | bb | 2023-11-07
154 | 154 | 20231106_HepCl_ReRun_156 | Mix 2 Active T60 A | Analyte | 4.4 | 89026.438 | 3511176 |  | 89026.438 | 4 | 9.96 | 149 | bb | 2023-11-07
155 | 155 | 20231106_HepCl_ReRun_157 | Mix 2 Active T60 B | Analyte | 4.4 | 89150.375 | 3327173 |  | 89150.375 | 4 | 9.97 | 149.3 | bb | 2023-11-07
156 | 156 | 20231106_HepCl_ReRun_158 | Mix 2 Active T60 C | Analyte | 4.4 | 89808.383 | 3663026 |  | 89808.383 | 4 | 10.05 | 151.2 | bb | 2023-11-07
157 | 157 | 20231106_HepCl_ReRun_159 | Blank | Blank | 4.4 | 0.75 | 88 |  | 0.75 | 4 | 0 | -100 | bb | 2023-11-07
158 | 158 | 20231106_HepCl_ReRun_160 | Mix 2 Active T120 A | Analyte | 4.4 | 78967.586 | 3313229 |  | 78967.586 | 4 | 8.83 | 120.8 | bb | 2023-11-07
159 | 159 | 20231106_HepCl_ReRun_161 | Mix 2 Active T120 B | Analyte | 4.4 | 83951.203 | 3239414 |  | 83951.203 | 4 | 9.39 | 134.8 | bb | 2023-11-07
160 | 160 | 20231106_HepCl_ReRun_162 | Mix 2 Active T120 C | Analyte | 4.4 | 82722.727 | 3247101 |  | 82722.727 | 4 | 9.25 | 131.3 | bb | 2023-11-07
161 | 161 | 20231106_HepCl_ReRun_163 | Mix 2 Active T240 A | Analyte | 4.4 | 80214.852 | 3212792 |  | 80214.852 | 4 | 8.97 | 124.3 | bb | 2023-11-07
162 | 162 | 20231106_HepCl_ReRun_164 | Mix 2 Active T240 B | Analyte | 4.4 | 77124.359 | 3314510 |  | 77124.359 | 4 | 8.63 | 115.7 | bb | 2023-11-07
163 | 163 | 20231106_HepCl_ReRun_165 | Mix 2 Active T240 C | Analyte | 4.4 | 74714.813 | 3064155 |  | 74714.813 | 4 | 8.36 | 108.9 | bb | 2023-11-07
164 | 164 | 20231106_HepCl_ReRun_166 | Blank | Blank | 4.41 | 3.511 | 188 |  | 3.511 | 4 | 0 | -100 | bb | 2023-11-07
165 | 165 | 20231106_HepCl_ReRun_167 | CC12 | Analyte | 4.4 | 19198.016 | 739716 |  | 19198.016 | 4 | 2.15 | -46.3 | bb | 2023-11-07
166 | 166 | 20231106_HepCl_ReRun_168 | Blank | Blank | [5 empty] | 4 |  |  |  | 2023-11-07
167 | 167 | 20231106_HepCl_ReRun_169 | Mix 2 Media T0 A | Analyte | 4.4 | 99319.828 | 4043231 |  | 99319.828 | 4 | 11.11 | 177.8 | bb | 2023-11-07
168 | 168 | 20231106_HepCl_ReRun_170 | Mix 2 Media T0 B | Analyte | 4.4 | 94711.563 | 3750901 |  | 94711.563 | 4 | 10.59 | 164.9 | bb | 2023-11-07
169 | 169 | 20231106_HepCl_ReRun_171 | Mix 2 Media T0 C | Analyte | 4.4 | 90822.211 | 3627143 |  | 90822.211 | 4 | 10.16 | 154 | bb | 2023-11-07
170 | 170 | 20231106_HepCl_ReRun_172 | Mix 2 Media T15 A | Analyte | 4.4 | 84107.078 | 3258718 |  | 84107.078 | 4 | 9.41 | 135.2 | bb | 2023-11-07
171 | 171 | 20231106_HepCl_ReRun_173 | Mix 2 Media T15 B | Analyte | 4.4 | 82892.469 | 3306039 |  | 82892.469 | 4 | 9.27 | 131.8 | bb | 2023-11-07
172 | 172 | 20231106_HepCl_ReRun_174 | Mix 2 Media T15 C | Analyte | 4.4 | 89697.148 | 3831743 |  | 89697.148 | 4 | 10.03 | 150.8 | bb | 2023-11-07
173 | 173 | 20231106_HepCl_ReRun_175 | Blank | Blank | 4.4 | 6.395 | 288 |  | 6.395 | 4 | 0 | -100 | bb | 2023-11-07
174 | 174 | 20231106_HepCl_ReRun_176 | Mix 2 Media T30 A | Analyte | 4.4 | 84959.609 | 3222198 |  | 84959.609 | 4 | 9.5 | 137.6 | bb | 2023-11-07
175 | 175 | 20231106_HepCl_ReRun_177 | Mix 2 Media T30 B | Analyte | 4.4 | 86283.852 | 3504027 |  | 86283.852 | 4 | 9.65 | 141.3 | bb | 2023-11-07
176 | 176 | 20231106_HepCl_ReRun_178 | Mix 2 Media T30 C | Analyte | 4.4 | 100046.359 | 4210013 |  | 100046.359 | 4 | 11.19 | 179.8 | bb | 2023-11-07
177 | 177 | 20231106_HepCl_ReRun_179 | Mix 2 Media T60 A | Analyte | 4.4 | 95674.383 | 4084776 |  | 95674.383 | 4 | 10.7 | 167.6 | bb | 2023-11-07
178 | 178 | 20231106_HepCl_ReRun_180 | Mix 2 Media T60 B | Analyte | 4.4 | 93804.273 | 3910266 |  | 93804.273 | 4 | 10.49 | 162.3 | bb | 2023-11-07
179 | 179 | 20231106_HepCl_ReRun_181 | Mix 2 Media T60 C | Analyte | 4.4 | 75229.672 | 2819570 |  | 75229.672 | 4 | 8.42 | 110.4 | bb | 2023-11-07
180 | 180 | 20231106_HepCl_ReRun_182 | Blank | Blank | 4.4 | 0.43 | 81 |  | 0.43 | 4 | 0 | -100 | bb | 2023-11-07
181 | 181 | 20231106_HepCl_ReRun_183 | Mix 2 Media T120 A | Analyte | 4.4 | 87124.805 | 3374102 |  | 87124.805 | 4 | 9.75 | 143.7 | bb | 2023-11-07
182 | 182 | 20231106_HepCl_ReRun_184 | Mix 2 Media T120 B | Analyte | 4.4 | 77005.422 | 3313557 |  | 77005.422 | 4 | 8.61 | 115.4 | bb | 2023-11-07
183 | 183 | 20231106_HepCl_ReRun_185 | Mix 2 Media T120 C | Analyte | 4.4 | 88369.641 | 3280962 |  | 88369.641 | 4 | 9.89 | 147.1 | bb | 2023-11-07
184 | 184 | 20231106_HepCl_ReRun_186 | Mix 2 Media T240 A | Analyte | 4.4 | 76210.016 | 3155056 |  | 76210.016 | 4 | 8.53 | 113.1 | bb | 2023-11-07
185 | 185 | 20231106_HepCl_ReRun_187 | Mix 2 Media T240 B | Analyte | 4.4 | 52721.07 | 2267876 |  | 52721.07 | 4 | 5.9 | 47.4 | bb | 2023-11-07
186 | 186 | 20231106_HepCl_ReRun_188 | Mix 2 Media T240 C | Analyte | 4.4 | 55817.023 | 2277350 |  | 55817.023 | 4 | 6.24 | 56.1 | bb | 2023-11-07
187 | 187 | 20231106_HepCl_ReRun_189 | Blank | Blank | 4.42 | 2.904 | 87 |  | 2.904 | 4 | 0 | -100 | bb | 2023-11-07
188 | 188 | 20231106_HepCl_ReRun_190 | CC12 | Analyte | 4.4 | 21716.965 | 892153 |  | 21716.965 | 4 | 2.43 | -39.3 | bb | 2023-11-07
189 | 189 | 20231106_HepCl_ReRun_191 | Blank | Blank | 4.39 | 2.254 | 110 |  | 2.254 | 4 | 0 | -100 | bb | 2023-11-07
190 | 190 | 20231106_HepCl_ReRun_192 | Mix 2 Inactive T0 A | Analyte | 4.4 | 99776.836 | 4067719 |  | 99776.836 | 4 | 11.16 | 179 | bb | 2023-11-07
191 | 191 | 20231106_HepCl_ReRun_193 | Mix 2 Inactive T0 B | Analyte | 4.4 | 101979.219 | 4124754 |  | 101979.219 | 4 | 11.41 | 185.2 | bb | 2023-11-07
192 | 192 | 20231106_HepCl_ReRun_194 | Mix 2 Inactive T0 C | Analyte | 4.4 | 114534.242 | 4796443 |  | 114534.242 | 4 | 12.81 | 220.3 | bb | 2023-11-07
193 | 193 | 20231106_HepCl_ReRun_195 | Mix 2 Inactive T240 A | Analyte | 4.4 | 79844.398 | 3250808 |  | 79844.398 | 4 | 8.93 | 123.3 | bb | 2023-11-07
194 | 194 | 20231106_HepCl_ReRun_196 | Mix 2 Inactive T240 B | Analyte | 4.4 | 86026.953 | 3398443 |  | 86026.953 | 4 | 9.62 | 140.6 | bb | 2023-11-07
195 | 195 | 20231106_HepCl_ReRun_197 | Mix 2 Inactive T240 C | Analyte | 4.4 | 90566.922 | 3678126 |  | 90566.922 | 4 | 10.13 | 153.3 | bb | 2023-11-07
196 | 196 | 20231106_HepCl_ReRun_198 | Blank | Blank | 4.37 | 0.96 | 105 |  | 0.96 | 4 | 0 | -100 | bb | 2023-11-08
197 | 197 | 20231106_HepCl_ReRun_199 | CC12 | Analyte | 4.4 | 19258.826 | 790180 |  | 19258.826 | 4 | 2.15 | -46.1 | bb | 2023-11-08
198 | 198 | 20231106_HepCl_ReRun_200 | Blank | Blank | 4.4 | 1.102 | 122 |  | 1.102 | 4 | 0 | -100 | bb | 2023-11-08
199 | 199 | 20231106_HepCl_ReRun_201 | Mix 3 Active T0 A | Analyte | 4.4 | 97325.922 | 4022685 |  | 97325.922 | 4 | 10.89 | 172.2 | bb | 2023-11-08
200 | 200 | 20231106_HepCl_ReRun_202 | Mix 3 Active T0 B | Analyte | 4.4 | 102662.727 | 3738008 |  | 102662.727 | 4 | 11.48 | 187.1 | bb | 2023-11-08
201 | 201 | 20231106_HepCl_ReRun_203 | Mix 3 Active T0 C | Analyte | 4.4 | 99524.109 | 3978316 |  | 99524.109 | 4 | 11.13 | 178.3 | bb | 2023-11-08
202 | 202 | 20231106_HepCl_ReRun_204 | Mix 3 Active T15 A | Analyte | 4.4 | 104638.984 | 4058627 |  | 104638.984 | 4 | 11.71 | 192.6 | bb | 2023-11-08
203 | 203 | 20231106_HepCl_ReRun_205 | Mix 3 Active T15 B | Analyte | 4.4 | 95722.523 | 3923375 |  | 95722.523 | 4 | 10.71 | 167.7 | bb | 2023-11-08
204 | 204 | 20231106_HepCl_ReRun_206 | Mix 3 Active T15 C | Analyte | 4.4 | 104475.281 | 4375782 |  | 104475.281 | 4 | 11.69 | 192.2 | bb | 2023-11-08
205 | 205 | 20231106_HepCl_ReRun_207 | Blank | Blank | 4.4 | 1.83 | 193 |  | 1.83 | 4 | 0 | -100 | bb | 2023-11-08
206 | 206 | 20231106_HepCl_ReRun_208 | Mix 3 Active T30 A | Analyte | 4.4 | 96895.539 | 3608709 |  | 96895.539 | 4 | 10.84 | 171 | bb | 2023-11-08
207 | 207 | 20231106_HepCl_ReRun_209 | Mix 3 Active T30 B | Analyte | 4.4 | 90586.703 | 3675424 |  | 90586.703 | 4 | 10.13 | 153.3 | bb | 2023-11-08
208 | 208 | 20231106_HepCl_ReRun_210 | Mix 3 Active T30 C | Analyte | 4.4 | 85220.258 | 3006224 |  | 85220.258 | 4 | 9.53 | 138.3 | bb | 2023-11-08
209 | 209 | 20231106_HepCl_ReRun_211 | Mix 3 Active T60 A | Analyte | 4.4 | 87958.188 | 3322711 |  | 87958.188 | 4 | 9.84 | 146 | bb | 2023-11-08
210 | 210 | 20231106_HepCl_ReRun_212 | Mix 3 Active T60 B | Analyte | 4.4 | 73946.117 | 2831281 |  | 73946.117 | 4 | 8.27 | 106.8 | bb | 2023-11-08
211 | 211 | 20231106_HepCl_ReRun_213 | Mix 3 Active T60 C | Analyte | 4.4 | 88629.414 | 3533247 |  | 88629.414 | 4 | 9.91 | 147.9 | bb | 2023-11-08
212 | 212 | 20231106_HepCl_ReRun_214 | Blank | Blank | 4.43 | 1.23 | 122 |  | 1.23 | 4 | 0 | -100 | bb | 2023-11-08
213 | 213 | 20231106_HepCl_ReRun_215 | Mix 3 Active T120 A | Analyte | 4.4 | 93904.141 | 4091944 |  | 93904.141 | 4 | 10.5 | 162.6 | bb | 2023-11-08
214 | 214 | 20231106_HepCl_ReRun_216 | Mix 3 Active T120 B | Analyte | 4.4 | 99092.422 | 3960975 |  | 99092.422 | 4 | 11.08 | 177.1 | bb | 2023-11-08
215 | 215 | 20231106_HepCl_ReRun_217 | Mix 3 Active T120 C | Analyte | 4.4 | 102824.828 | 3514491 |  | 102824.828 | 4 | 11.5 | 187.6 | bb | 2023-11-08
216 | 216 | 20231106_HepCl_ReRun_218 | Mix 3 Active T240 A | Analyte | 4.4 | 74625.852 | 2961373 |  | 74625.852 | 4 | 8.35 | 108.7 | bb | 2023-11-08
217 | 217 | 20231106_HepCl_ReRun_219 | Mix 3 Active T240 B | Analyte | 4.4 | 78271.125 | 3317273 |  | 78271.125 | 4 | 8.76 | 118.9 | bb | 2023-11-08
218 | 218 | 20231106_HepCl_ReRun_220 | Mix 3 Active T240 C | Analyte | 4.4 | 82232.203 | 3383647 |  | 82232.203 | 4 | 9.2 | 130 | bb | 2023-11-08
219 | 219 | 20231106_HepCl_ReRun_221 | Blank | Blank | 4.38 | 0.419 | 76 |  | 0.419 | 4 | 0 | -100 | bb | 2023-11-08
220 | 220 | 20231106_HepCl_ReRun_222 | CC12 | Analyte | 4.4 | 19805.637 | 827016 |  | 19805.637 | 4 | 2.22 | -44.6 | bb | 2023-11-08
221 | 221 | 20231106_HepCl_ReRun_223 | Blank | Blank | 4.39 | 5.807 | 229 |  | 5.807 | 4 | 0 | -100 | bb | 2023-11-08
222 | 222 | 20231106_HepCl_ReRun_224 | Mix 3 Media T0 A | Analyte | 4.4 | 68976.289 | 2929824 |  | 68976.289 | 4 | 7.72 | 92.9 | bb | 2023-11-08
223 | 223 | 20231106_HepCl_ReRun_225 | Mix 3 Media T0 B | Analyte | 4.4 | 80912.125 | 3194553 |  | 80912.125 | 4 | 9.05 | 126.3 | bb | 2023-11-08
224 | 224 | 20231106_HepCl_ReRun_226 | Mix 3 Media T0 C | Analyte | 4.4 | 64799.211 | 2501513 |  | 64799.211 | 4 | 7.25 | 81.2 | bb | 2023-11-08
225 | 225 | 20231106_HepCl_ReRun_227 | Mix 3 Media T15 A | Analyte | 4.4 | 95344.883 | 3645451 |  | 95344.883 | 4 | 10.67 | 166.6 | bb | 2023-11-08
226 | 226 | 20231106_HepCl_ReRun_228 | Mix 3 Media T15 B | Analyte | 4.4 | 98936.914 | 3873731 |  | 98936.914 | 4 | 11.07 | 176.7 | bb | 2023-11-08
227 | 227 | 20231106_HepCl_ReRun_229 | Mix 3 Media T15 C | Analyte | 4.4 | 96992.188 | 3737935 |  | 96992.188 | 4 | 10.85 | 171.3 | bb | 2023-11-08
228 | 228 | 20231106_HepCl_ReRun_230 | Blank | Blank | 4.39 | 2.226 | 148 |  | 2.226 | 4 | 0 | -100 | bb | 2023-11-08
229 | 229 | 20231106_HepCl_ReRun_231 | Mix 3 Media T30 A | Analyte | 4.4 | 108030.414 | 4106710 |  | 108030.414 | 4 | 12.08 | 202.1 | bb | 2023-11-08
230 | 230 | 20231106_HepCl_ReRun_232 | Mix 3 Media T30 B | Analyte | 4.4 | 100591.141 | 3646289 |  | 100591.141 | 4 | 11.25 | 181.3 | bb | 2023-11-08
231 | 231 | 20231106_HepCl_ReRun_233 | Mix 3 Media T30 C | Analyte | 4.4 | 69586.211 | 2536447 |  | 69586.211 | 4 | 7.78 | 94.6 | bb | 2023-11-08
232 | 232 | 20231106_HepCl_ReRun_234 | Mix 3 Media T60 A | Analyte | 4.4 | 104467.281 | 3438160 |  | 104467.281 | 4 | 11.69 | 192.2 | bb | 2023-11-08
233 | 233 | 20231106_HepCl_ReRun_235 | Mix 3 Media T60 B | Analyte | 4.4 | 48669.789 | 2086193 |  | 48669.789 | 4 | 5.44 | 36.1 | bb | 2023-11-08
234 | 234 | 20231106_HepCl_ReRun_236 | Mix 3 Media T60 C | Analyte | 4.4 | 39606.652 | 1617206 |  | 39606.652 | 4 | 4.43 | 10.8 | bb | 2023-11-08
235 | 235 | 20231106_HepCl_ReRun_237 | Blank | Blank | 4.37 | 8.198 | 144 |  | 8.198 | 4 | 0 | -100 | bb | 2023-11-08
236 | 236 | 20231106_HepCl_ReRun_238 | Mix 3 Media T120 A | Analyte | 4.4 | 100054.281 | 4122043 |  | 100054.281 | 4 | 11.19 | 179.8 | bb | 2023-11-08
237 | 237 | 20231106_HepCl_ReRun_239 | Mix 3 Media T120 B | Analyte | 4.4 | 91328.695 | 3601070 |  | 91328.695 | 4 | 10.22 | 155.4 | bb | 2023-11-08
238 | 238 | 20231106_HepCl_ReRun_240 | Mix 3 Media T120 C | Analyte | 4.4 | 102183.391 | 3971211 |  | 102183.391 | 4 | 11.43 | 185.8 | bb | 2023-11-08
239 | 239 | 20231106_HepCl_ReRun_241 | Mix 3 Media T240 A | Analyte | 4.4 | 95108.883 | 3551584 |  | 95108.883 | 4 | 10.64 | 166 | bb | 2023-11-08
240 | 240 | 20231106_HepCl_ReRun_242 | Mix 3 Media T240 B | Analyte | 4.4 | 53921.77 | 2013740 |  | 53921.77 | 4 | 6.03 | 50.8 | bb | 2023-11-08
241 | 241 | 20231106_HepCl_ReRun_243 | Mix 3 Media T240 C | Analyte | 4.4 | 61000.227 | 2395311 |  | 61000.227 | 4 | 6.82 | 70.6 | bb | 2023-11-08
242 | 242 | 20231106_HepCl_ReRun_244 | Blank | Blank | 4.44 | 7.8 | 151 |  | 7.8 | 4 | 0 | -100 | bb | 2023-11-08
243 | 243 | 20231106_HepCl_ReRun_245 | CC12 | Analyte | 4.4 | 20601.318 | 819968 |  | 20601.318 | 4 | 2.3 | -42.4 | bb | 2023-11-08
244 | 244 | 20231106_HepCl_ReRun_246 | Blank | Blank | 4.4 | 4.036 | 240 |  | 4.036 | 4 | 0 | -100 | bb | 2023-11-08
245 | 245 | 20231106_HepCl_ReRun_247 | Mix 3 Inactive T0 A | Analyte | 4.4 | 180164.406 | 6972059 |  | 180164.406 | 4 | 20.15 | 403.9 | bb | 2023-11-08
246 | 246 | 20231106_HepCl_ReRun_248 | Mix 3 Inactive T0 B | Analyte | 4.4 | 128731.063 | 5090614 |  | 128731.063 | 4 | 14.4 | 260 | bb | 2023-11-08
247 | 247 | 20231106_HepCl_ReRun_249 | Mix 3 Inactive T0 C | Analyte | 4.4 | 119917.117 | 4419736 |  | 119917.117 | 4 | 13.41 | 235.4 | bb | 2023-11-08
248 | 248 | 20231106_HepCl_ReRun_250 | Mix 3 Inactive T240 A | Analyte | 4.4 | 88866.891 | 3528454 |  | 88866.891 | 4 | 9.94 | 148.5 | bb | 2023-11-08
249 | 249 | 20231106_HepCl_ReRun_251 | Mix 3 Inactive T240 B | Analyte | 4.4 | 88157.313 | 3428768 |  | 88157.313 | 4 | 9.86 | 146.5 | bb | 2023-11-08
250 | 250 | 20231106_HepCl_ReRun_252 | Mix 3 Inactive T240 C | Analyte | 4.4 | 98938.063 | 4006047 |  | 98938.063 | 4 | 11.07 | 176.7 | bb | 2023-11-08
251 | 251 | 20231106_HepCl_ReRun_253 | Blank | Blank | 4.4 | 1.238 | 124 |  | 1.238 | 4 | 0 | -100 | bb | 2023-11-08
252 | 252 | 20231106_HepCl_ReRun_254 | Blank | Blank | 4.39 | 7.609 | 217 |  | 7.609 | 4 | 0 | -100 | bb | 2023-11-08
253 | 253 | 20231106_HepCl_ReRun_255 | Blank | Blank | 4.41 | 3 | 179 |  | 3 | 4 | 0 | -100 | bb | 2023-11-08

## Sheet9
6.4714 | 6.4714 | 6.4359 | 6.3449
6.4359 | 6.4884 | 6.4038 | 6.4766
6.3449 | 6.4815 | 6.4739 | 6.5256
6.4884 | 6.5417 | 6.5375 | 6.5417
6.4038 | 6.3849 | 6.384 | 6.4248
6.4766 | 6.4557 | 6.4338 | 6.4534
6.4815
6.4739
6.5256
6.5417
6.5375
6.5417
6.3849
6.384
6.4248
6.4557
6.4338
6.4534
6.4519 | 6.4519 | 6.3658 | 6.384
6.3658 | 6.4318 | 6.5193 | 6.4717
6.384 | 6.4768 | 6.4186 | 6.4595
6.4318 | 6.5433 | 6.6014 | 6.5566
6.5193 | 6.387 | 6.4438 | 6.4351
6.4717 | 6.467 | 6.4351 | 6.3619
6.4768
6.4186
6.4595
6.5433
6.6014
6.5566
6.387
6.4438
6.4351
6.467
6.4351
6.3619
6.3441 | 6.3441 | 6.3519 | 6.3435
6.3519
6.3435
6.403
6.463
6.463 | 6.403 | 6.463 | 6.463