## CoverSheet
Analytical Analysis Cover Sheet for Hepatic Clearance Assessment of PFAS Compounds
QAPP: | In Vitro Toxicokinetic (TK) Assessments of PFAS (D-CCED-0031343-1-1)
PI: | Barbara A. Wetmore
Analysis Set ID: | HepCl_20240206
Analytes: | K-PFHxS, Na PFOA, PFOSA
Calibration Curve Prepared Date: | 2024-01-22
Preparer: | KR
Sample Type: | Calibration Curve (CC1 - CC17), Curve Precision Check (QC-A - QC-D)
Hepatic Clearance Assay Conducted Date: | 2024-01-22
Preparer: | KR, MP
Sample Type: | Hepatic Clearance Samples
Analytical Analysis Date: | 2/6-9/2024
Analytical Technique: | UPLC-MS/MS
Instrument: | Waters Xevo TQ-S micro (QEB0036)
Analyst: | KR
Relevant SOPs: | I-CCED-ETTB-4568; I-CCED-ETTB-SOP-3632
Notebook Entries: | Hepatocytic Clearance (Human) of Select Compounds
Total Samples Number (including blanks): | 151
Actual Sample Number: | 112
Method Detection Limit (MDL) Definition: | Minimum measured concentration of a substance that can be reported with 99% confidence that the measured concentration is distinguishable from method blank results
MDLXi = method detection limit determined for target analyte X from solution i = t(n-1,1-α=0.99) x Si
t(n-1,1-α=0.99) = 3.143, the student's t-value for a 99% confidence level and a standard deviation estimate with six degrees of freedom (n=7) |  |  |  | 40 CFR Part 136 | EPA Definition and Procedure for the Determination of the Method Detection Limit, Revision 2 (December 2016)
Si = standard deviation of the seven replicate analyses for solution i |  |  |  | https://www.ecfr.gov/cgi-bin/text-idx?SID=a68a02b6d783f9356758b5ff0ed106&mc=true&node=pt40.25.136&rgn=div5
estimated Limit of Quantitation (eLOQ) Definition: | Lowest Calibration Curve point at which precision checks can be quantitated against curve ± 30%, provided other curve points have accuracy to ± 25%
QC Samples and Acceptance Criteria
QC Type | Frequency | Acceptance Criteria | Type | Corrective Action if QC Fails
Calibration Curve Linearity (Analytical) | 1 each run | 7-point curve min; r2≥0.98 | Linearity | Re-run samples and curve
Curve Precision Check (Analytical) | 2-3 samples per run | 75-125% of expected | Accuracy | Flag in comments
Blanks (Analytical) | 1 every 6-10 injections | <1/2 MDL of analytes | Specificity | Flag in comments, assess and resolve
Technical Replicate Assessment (Assay) | Every assay sample | Individual Replicates within ± 25% of Sample Mean | Precision | Flag in comments
Positive Control Reference Chemical(s) for Protein Binding (Assay) | 1 per assay batch | ± 20% of historical or published values | Binding Assay Performance | Flag in comments
Note: one of the three blanks (solvent, crash blank, crash mixed matrix) should be run every 6 injections
Analytes, Internal Standards Assessed
Analyte (Common Name) | DTXSID | Sample ID | Average Molecular Weight (g/mol) | Internal Standard | Peak Assignment from CoA | Mix | Concentration at Instrument (pg/uL ¦ µM)
Potassium perfluorohexanesulfonate | DTXSID3037709 | 1210314796 | 438.2 | M3PFHxS | G | 1 | 0.01
Sodium perfluorooctanoate | DTXSID40880025 | 1208990462 | 436.052 | M8PFOA | I | 1 | 0.01
Perfluorooctanesulfonamide | DTXSID3038939 | 1210314819 | 499.14 | M8FOSA | 0 | 1 | 0.01
Phenacetin | DTXSID1021116 |  | 179.219 | M8PFOA | I | 1 | 0.01
Propranolol | DTXSID6023525 |  | 259.349 | M8PFOA | I | 1 | 0.01
Internal Standard Information
Standard | Vendor | Catalog Number | Lot Number | Other Descriptors
MPFAC-24ES | Wellington | MPFAC-23ES | MPFAC24ES0123
Hepatocytes | Vendor | Catalog Number | Lot Number | Species
LiverPool 20 Donor Hepatocytes Mixed Donor | BioIVT | X008000 | WGR | Human
Sample ID Key
CC## | Calibration Curve
QC-X | Quality Check
Human Mix # Active T# X | Species Mix # SampleType Time# Replicate | [6 empty] | An Example Certificate of Analysis for Wellington MPFAC-24ES Mix (Mass-labelled PFAS Solution)
*MP = Matt Phillips
*KR = Krishna Ravindra
*BAW = Barbara Wetmore
Additional notes
Concentrations listed are at the instrument, unless otherwise specified

## Dilution Factors
PFHxS, PFOA |  | PFOSA |  | Phenacetin
Step | Dilution Factor | Step | Dilution Factor | Step | Dilution Factor
Acetonitrile Crash | 3 | Acetonitrile Crash | 3 | Acetonitrile Crash | 3
Mix Prep | 3 | Mix Prep | 1.5 | Mix Prep | 2
Mobile Phase | 4 | Mobile Phase | 4 | Mobile Phase | 4
Final Factor | 36 | Final Factor | 18 | Final Factor | 24

## Clint Data Summary
Hepatic Clearance Data Summary
Human | [21 empty] | p-value adjusted Cl
DTXSID | Common Name | Sample ID | Mix/SPE | Analyte | Best-Fit Values | [8 empty] | No Cells BkgdCl |  |  |  | HT BkgdCl |  |  |  | Clearance (µL/min/106 hepatocytes) | Bkgd-Adj Clearance
[5 empty] | Slope | r2 | Sy.x | F | DFn | DFd | P value | Half Life | Clearance | pvalue | Slope | Half Life | Clearance | pvalue | Slope | Half Life | Clearance
DTXSID3037709 | Potassium perfluorohexanesulfonate | 1210314796 | 1 | K-PFHxS | -0.0004644 | 0.02235 | 0.2683 | 0.3657 | 1 | 16 | 0.5538 | 1492.2481 | 0.9288 | 0.9364 | 6.116e-05 | -11330.9353 | -0.12232 | 0.1554 | 0.001338 | -517.9372 | -2.676 | 0 | 0
DTXSID40880025 | Sodium perfluorooctanoate | 1208990462 | 1 | Na PFOA | -0.001666 | 0.4792 | 0.1517 | 14.72 | 1 | 16 | 0.0015 | 415.9664 | 3.332 | 0.0271 | -0.00166 | 417.4699 | 3.32 | 0.7849 | 0.0001319 | -5253.9803 | -0.2638 | 3.332 | 0.012
DTXSID3038939 | Perfluorooctanesulfonamide | 1210314819 | 1 | PFOSA | -0.004457 | 0.966 | 0.07305 | 454.4 | 1 | 16 | <0.0001 | 155.4858 | 8.914 | 0.7147 | -0.0002024 | 3423.913 | 0.4048 | 0.0906 | -0.0006092 | 1137.5575 | 1.2184 | 8.914 | 8.91
DTXSID1021116 | Phenacetin | N/A | Ref | Phenacetin | -0.004215 | 0.7828 | 0.1939 | 57.66 | 1 | 16 | <0.0001 | 164.4128 | 8.43 | 0.1579 | -0.001191 | 581.864 | 2.382 | 0.7553 | 0.0001879 | -3688.132 | -0.3758 | 8.43 | 8.43
DTXSID6023525 | Propranolol | N/A | Ref | Propranolol | -0.006102 | 0.9188 | 0.1585 | 181 | 1 | 16 | <0.0001 | 113.5693 | 12.204 | 0.1902 | -0.001073 | 645.8527 | 2.146 | 0.7465 | 0.0001567 | -4422.4633 | -0.3134 | 12.204 | 12.2

## K-PFHxS
Human Active Cells | [8 empty] | Media Only | [7 empty] | Compound 1: PFHxS
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20240206_HepCl_007 | CC01 | Standard | 3.82 | 665.396 | 15303 | 13228.613 | 5.03 | 0.175 | 0.33 | 88.2
0 | Active T0 A | 10.86 | 5.9686 | 390.96 | 377.52 | 28.4134 | 7.526% |  | 0 | Media T0 A | 11.77 | 6.0491 | 423.72 | 411.6 | 10.7539 | 2.613% | 20240206_HepCl_050 | CC01 | Standard | 3.82 | 644.485 | 14194 | 13672.755 | 4.714 | 0.175 | 0.31 | 75.2
| Active T0 B | 11.02 | 5.9832 | 396.72 | [5 empty] | Media T0 B | 11.2 | 5.9994 | 403.2 |  |  |  | 20240206_HepCl_008 | CC02 | Standard | 3.82 | 592.364 | 12612 | 13060.495 | 4.536 | 0.281 | 0.29 | 4.5
| Active T0 C | 9.58 | 5.8432 | 344.88 | [5 empty] | Media T0 C | 11.33 | 6.011 | 407.88 |  |  |  | 20240206_HepCl_051 | CC02 | Standard | 3.82 | 651.616 | 14548 | 14006.428 | 4.652 | 0.281 | 0.3 | 7.4
15 | Active T15 A | 18.77 | 6.5158 | 675.72 | 633.84 | 170.0528 | 26.83% |  | 15 | Media T15 A | 15.28 | 6.3101 | 550.08 | 426 | 109.6059 | 25.73% | 20240206_HepCl_009 | CC03 | Standard | 3.82 | 907.246 | 21508 | 12383.586 | 7.326 | 0.444 | 0.49 | 11.5
| Active T15 B | 12.41 | 6.102 | 446.76 | [5 empty] | Media T15 B | 10.71 | 5.9547 | 385.56 |  |  |  | 20240206_HepCl_052 | CC03 | Standard | 3.82 | 811.079 | 17729 | 12324.795 | 6.581 | 0.444 | 0.44 | -0.6
| Active T15 C | 21.64 | 6.6581 | 779.04 | [5 empty] | Media T15 C | 9.51 | 5.8359 | 342.36 |  |  |  | 20240206_HepCl_010 | CC04 | Standard | 3.82 | 1198.078 | 30750 | 14313.511 | 8.37 | 0.713 | 0.57 | -20
30 | Active T30 A | 13.21 | 6.1645 | 475.56 | 364.08 | 98.1106 | 26.95% |  | 30 | Media T30 A | 8.43 | 5.7153 | 303.48 | 264.6 | 51.7061 | 19.54% | 20240206_HepCl_034 | CC04 | Standard | 3.81 | 1166.829 | 28096 | 11906.045 | 9.8 | 0.713 | 0.67 | -5.6
| Active T30 B | 8.08 | 5.6729 | 290.88 | [5 empty] | Media T30 B | 5.72 | 5.3275 | 205.92 |  |  |  | 20240206_HepCl_039 | CC04 | Standard | 3.82 | 1174.397 | 28083 | 12918.972 | 9.09 | 0.713 | 0.62 | -12.8
| Active T30 C | 9.05 | 5.7863 | 325.8 | [5 empty] | Media T30 C | 7.9 | 5.6504 | 284.4 |  |  |  | 20240206_HepCl_044 | CC04 | Standard | 3.82 | 1285.76 | 31434 | 12008.481 | 10.707 | 0.713 | 0.74 | 3.6
60 | Active T60 A | 12.21 | 6.0858 | 439.56 | 398.52 | 54.8371 | 13.76% |  | 60 | Media T60 A | 12.02 | 6.0701 | 432.72 | 385.32 | 105.3468 | 27.34% | 20240206_HepCl_053 | CC04 | Standard | 3.82 | 1271.823 | 32126 | 12797.363 | 9.938 | 0.713 | 0.68 | -4.2
| Active T60 B | 9.34 | 5.8178 | 336.24 | [5 empty] | Media T60 B | 12.74 | 6.1283 | 458.64 |  |  |  | 20240206_HepCl_077 | CC04 | Standard | 3.82 | 1257.506 | 31257 | 10298.73 | 12.21 | 0.713 | 0.85 | 18.7
| Active T60 C | 11.66 | 6.0397 | 419.76 | [5 empty] | Media T60 C | 7.35 | 5.5782 | 264.6 |  |  |  | 20240206_HepCl_082 | CC04 | Standard | 3.82 | 1128.267 | 27893 | 12015.393 | 9.39 | 0.713 | 0.64 | -9.7
120 | Active T120 A | 9.48 | 5.8327 | 341.28 | 334.32 | 29.0717 | 8.696% |  | 120 | Media T120 A | 9.57 | 5.8422 | 344.52 | 287.28 | 52.7999 | 18.38% | 20240206_HepCl_087 | CC04 | Standard | 3.82 | 1347.186 | 35163 | 15304.867 | 8.802 | 0.713 | 0.6 | -15.7
| Active T120 B | 8.4 | 5.7118 | 302.4 | [5 empty] | Media T120 B | 6.68 | 5.4826 | 240.48 |  |  |  | 20240206_HepCl_011 | CC05 | Standard | 3.82 | 2166.731 | 54951 | 10278.519 | 21.08 | 1.138 | 1.48 | 30.5
| Active T120 C | 9.98 | 5.8841 | 359.28 | [5 empty] | Media T120 C | 7.69 | 5.6234 | 276.84 |  |  |  | 20240206_HepCl_035 | CC05 | Standard | 3.82 | 2021.19 | 50930 | 12084.42 | 16.726 | 1.138 | 1.17 | 2.9
240 | Active T240 A | 11.11 | 5.9914 | 399.96 | 420.72 | 93.7206 | 22.28% |  | 240 | Media T240 A | 10.74 | 5.9575 | 386.64 | 402.84 | 47.6521 | 11.83% | 20240206_HepCl_040 | CC05 | Standard | 3.82 | 1709.357 | 41074 | 11582.531 | 14.758 | 1.138 | 1.03 | -9.5
| Active T240 B | 14.53 | 6.2597 | 523.08 | [5 empty] | Media T240 B | 12.68 | 6.1235 | 456.48 |  |  |  | 20240206_HepCl_045 | CC05 | Standard | 3.82 | 1746.354 | 40763 | 11747.538 | 14.866 | 1.138 | 1.04 | -8.8
| Active T240 C | 9.42 | 5.8264 | 339.12 | [5 empty] | Media T240 C | 10.15 | 5.901 | 365.4 |  |  |  | 20240206_HepCl_054 | CC05 | Standard | 3.82 | 1961.893 | 48988 | 14087.408 | 13.927 | 1.138 | 0.97 | -14.8
[17 empty] | 20240206_HepCl_078 | CC05 | Standard | 3.82 | 2021.016 | 50230 | 13693.741 | 14.759 | 1.138 | 1.03 | -9.5
Human Inactive Cells | [16 empty] | 20240206_HepCl_083 | CC05 | Standard | 3.82 | 2348.791 | 61487 | 15032.213 | 15.625 | 1.138 | 1.09 | -4
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20240206_HepCl_088 | CC05 | Standard | 3.82 | 2182.837 | 54618 | 13440.987 | 16.24 | 1.138 | 1.14 | -0.1
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20240206_HepCl_012 | CC06 | Standard | 3.82 | 3070.819 | 77651 | 12826.483 | 23.941 | 1.825 | 1.69 | -7.4
0 | Inactive T0 A | 9.78 | 5.8639 | 352.08 | 266.16 | 77.381 | 29.07% | [9 empty] | 20240206_HepCl_036 | CC06 | Standard | 3.81 | 3175.081 | 83981 | 11320.455 | 28.047 | 1.825 | 1.99 | 8.8
| Inactive T0 B | 5.61 | 5.3081 | 201.96 | [12 empty] | 20240206_HepCl_041 | CC06 | Standard | 3.82 | 3226.61 | 87559 | 13133.145 | 24.568 | 1.825 | 1.74 | -4.9
| Inactive T0 C | 6.79 | 5.499 | 244.44 | [12 empty] | 20240206_HepCl_046 | CC06 | Standard | 3.82 | 2751.479 | 67568 | 11973.123 | 22.98 | 1.825 | 1.62 | -11.2
240 | Inactive T240 A | 11.41 | 6.018 | 410.76 | 358.74 | 73.5674 | 20.51% | [9 empty] | 20240206_HepCl_055 | CC06 | Standard | 3.82 | 3377.518 | 86556 | 12849.26 | 26.286 | 1.825 | 1.86 | 1.8
| Inactive T240 B | 8.52 | 5.7259 | 306.72 | [12 empty] | 20240206_HepCl_079 | CC06 | Standard | 3.82 | 3237.091 | 88521 | 13200.004 | 24.523 | 1.825 | 1.73 | -5.1
| Inactive T240 C | 10.04 | 5.8901 | 361.44 | [12 empty] | 20240206_HepCl_084 | CC06 | Standard | 3.82 | 3132.317 | 79741 | 12810.994 | 24.45 | 1.825 | 1.73 | -5.4
[17 empty] | 20240206_HepCl_089 | CC06 | Standard | 3.82 | 3372.854 | 88752 | 13731.132 | 24.564 | 1.825 | 1.73 | -4.9
| Calibration Curve | [15 empty] | 20240206_HepCl_014 | CC07 | Standard | 3.82 | 4642.819 | 118697 | 13677.68 | 33.944 | 2.919 | 2.41 | -17.4
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20240206_HepCl_057 | CC07 | Standard | 3.82 | 5172.834 | 138072 | 11294.601 | 45.799 | 2.919 | 3.26 | 11.8
| 20240206_HepCl_011 | CC05 | 1.138 | 1.48 | 1.1187 | 0.15968 | 0.14273 | 0.50188 | -0.016916 | [7 empty] | 20240206_HepCl_015 | CC08 | Standard | 3.82 | 7219.885 | 193161 | 12671.938 | 56.975 | 4.669 | 4.07 | -12.9
| 20240206_HepCl_035 | CC05 | 1.138 | 1.17 | [12 empty] | 20240206_HepCl_058 | CC08 | Standard | 3.82 | 8000.987 | 213848 | 13495.462 | 59.286 | 4.669 | 4.23 | -9.3
| 20240206_HepCl_040 | CC05 | 1.138 | 1.03 | [12 empty] | 20240206_HepCl_016 | CC09 | Standard | 3.82 | 9527.564 | 245663 | 9804.405 | 97.176 | 7.45 | 6.96 | -6.6
| 20240206_HepCl_045 | CC05 | 1.138 | 1.04 | [12 empty] | 20240206_HepCl_059 | CC09 | Standard | 3.82 | 10462.401 | 267364 | 12311.349 | 84.982 | 7.45 | 6.08 | -18.4
| 20240206_HepCl_054 | CC05 | 1.138 | 0.97 | [12 empty] | 20240206_HepCl_017 | CC10 | Standard | 3.82 | 17529.984 | 443192 | 14090.989 | 124.406 | 11.875 | 8.92 | -24.9
| 20240206_HepCl_078 | CC05 | 1.138 | 1.03 | [12 empty] | 20240206_HepCl_060 | CC10 | Standard | 3.82 | 20564.145 | 543171 | 12715.516 | 161.725 | 11.875 | 11.6 | -2.3
| 20240206_HepCl_083 | CC05 | 1.138 | 1.09 | [12 empty] | 20240206_HepCl_018 | CC11 | Standard | 3.82 | 30891.338 | 783737 | 12546.156 | 246.222 | 19.125 | 17.68 | -7.5
| 20240206_HepCl_088 | CC05 | 1.138 | 1.14 | [12 empty] | 20240206_HepCl_061 | CC11 | Standard | 3.82 | 33561.285 | 903724 | 13536.77 | 247.927 | 19.125 | 17.81 | -6.9
[17 empty] | 20240206_HepCl_019 | CC12 | Standard | 3.82 | 48357.563 | 1291990 | 12496.379 | 386.973 | 30.5 | 27.81 | -8.8
| Blanks should be less than… | 0.25094 |  | nM | [12 empty] | 20240206_HepCl_062 | CC12 | Standard | 3.82 | 59905.824 | 1629002 | 13279.041 | 451.131 | 30.5 | 32.43 | 6.3
| eLOQ | 1.138 |  | nM | [12 empty] | 20240206_HepCl_107 | CC12 | Standard | 3.82 | 50414.809 | 1320994 | 12644.87 | 398.698 | 30.5 | 28.65 | -6.1
[17 empty] | 20240206_HepCl_116 | CC12 | Standard | 3.82 | 56819.621 | 1537002 | 13688.918 | 415.078 | 30.5 | 29.83 | -2.2
| QC FLAGS | [15 empty] | 20240206_HepCl_132 | CC12 | Standard | 3.82 | 62562.719 | 1743878 | 13353.789 | 468.502 | 30.5 | 33.68 | 10.4
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20240206_HepCl_141 | CC12 | Standard | 3.82 | 63911.715 | 1723573 | 14119.17 | 452.659 | 30.5 | 32.54 | 6.7
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20240206_HepCl_021 | CC13 | Standard | 3.82 | 80208.766 | 2132800 | 12645.211 | 634.302 | 48.75 | 45.61 | -6.4
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20240206_HepCl_037 | CC13 | Standard | 3.82 | 73982.078 | 1876210 | 12121.979 | 610.314 | 48.75 | 43.88 | -10
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20240206_HepCl_042 | CC13 | Standard | 3.82 | 85995.836 | 2358451 | 12904.606 | 666.396 | 48.75 | 47.92 | -1.7
| Beloq eLoQ | [15 empty] | 20240206_HepCl_047 | CC13 | Standard | 3.82 | 85420.461 | 2232592 | 12636.182 | 675.999 | 48.75 | 48.61 | -0.3
[17 empty] | 20240206_HepCl_064 | CC13 | Standard | 3.82 | 88756.984 | 2438055 | 12511.845 | 709.384 | 48.75 | 51.01 | 4.6
[17 empty] | 20240206_HepCl_080 | CC13 | Standard | 3.82 | 93043.977 | 2559129 | 12701.495 | 732.544 | 48.75 | 52.68 | 8.1
[17 empty] | 20240206_HepCl_085 | CC13 | Standard | 3.82 | 80929.953 | 2105538 | 12612.149 | 641.682 | 48.75 | 46.14 | -5.4
[17 empty] | 20240206_HepCl_090 | CC13 | Standard | 3.82 | 92521.656 | 2488279 | 11894.017 | 777.884 | 48.75 | 55.94 | 14.7
[17 empty] | 20240206_HepCl_022 | CC14 | Standard | 3.82 | 137819.438 | 3808929 | 13375.271 | 1030.405 | 78.125 | 74.11 | -5.1
[17 empty] | 20240206_HepCl_065 | CC14 | Standard | 3.82 | 135522 | 3683825 | 14586.896 | 929.067 | 78.125 | 66.82 | -14.5
[17 empty] | 20240206_HepCl_023 | CC15 | Standard | 3.82 | 221276.516 | 6151594 | 11024.747 | 2007.089 | 125 | 144.38 | 15.5
[17 empty] | 20240206_HepCl_066 | CC15 | Standard | 3.82 | 234000.469 | 6380038 | 13164.907 | 1777.456 | 125 | 127.86 | 2.3
[17 empty] | 20240206_HepCl_024 | CC16 | Standard | 3.82 | 355811.5 | 9871077 | 15599.353 | 2280.938 | 200 | 164.09 | -18
[17 empty] | 20240206_HepCl_067 | CC16 | Standard | 3.82 | 372184.063 | 10188026 | 15479.135 | 2404.424 | 200 | 172.97 | -13.5
[17 empty] | 20240206_HepCl_025 | CC17 | Standard | 3.82 | 572068.563 | 15426399 | 11725.054 | 4879.027 | 325.625 | 351.02 | 7.8
[17 empty] | 20240206_HepCl_068 | CC17 | Standard | 3.82 | 638518.063 | 17553762 | 12682.382 | 5034.686 | 325.625 | 362.22 | 11.2
[17 empty] | 20240206_HepCl_028 | CCQC A | QC | 3.82 | 1670.837 | 42937 | 15079.548 | 11.08 | 0.625 | 0.76 | 22.3
[17 empty] | 20240206_HepCl_071 | CCQC A | QC | 3.82 | 1721.165 | 40731 | 17512.383 | 9.828 | 0.625 | 0.67 | 7.9
[17 empty] | 20240206_HepCl_029 | CCQC B | QC | 3.82 | 6524.17 | 174808 | 17196.908 | 37.938 | 2.506 | 2.7 | 7.6
[17 empty] | 20240206_HepCl_072 | CCQC B | QC | 3.82 | 6749.785 | 185797 | 17804.043 | 37.912 | 2.506 | 2.7 | 7.5
[17 empty] | 20240206_HepCl_030 | CCQC C | QC | 3.82 | 12457.592 | 324712 | 14326.982 | 86.952 | 6.269 | 6.22 | -0.7
[17 empty] | 20240206_HepCl_073 | CCQC C | QC | 3.82 | 14612.47 | 390227 | 15293.799 | 95.545 | 6.269 | 6.84 | 9.1
[17 empty] | 20240206_HepCl_031 | CCQC D | QC | 3.82 | 52733.43 | 1380563 | 13883.457 | 379.829 | 25 | 27.3 | 9.2
[17 empty] | 20240206_HepCl_074 | CCQC D | QC | 3.82 | 62387.379 | 1673291 | 14440.207 | 432.039 | 25 | 31.05 | 24.2
[17 empty] | 20240206_HepCl_093 | Active T0 A | Analyte | 3.82 | 86594.836 | 2325177 | 57196.434 | 151.399 |  | 10.86
[17 empty] | 20240206_HepCl_094 | Active T0 B | Analyte | 3.82 | 80860.594 | 1947570 | 52622.742 | 153.661 |  | 11.02
[17 empty] | 20240206_HepCl_095 | Active T0 C | Analyte | 3.82 | 84127 | 2268312 | 62977.305 | 133.583 |  | 9.58
[17 empty] | 20240206_HepCl_096 | Active T15 A | Analyte | 3.82 | 142816.328 | 3862124 | 54653.637 | 261.312 |  | 18.77
[17 empty] | 20240206_HepCl_097 | Active T15 B | Analyte | 3.82 | 97267.898 | 2538648 | 56251.129 | 172.917 |  | 12.41
[17 empty] | 20240206_HepCl_098 | Active T15 C | Analyte | 3.82 | 140795.516 | 3873275 | 46738.797 | 301.239 |  | 21.64
[17 empty] | 20240206_HepCl_100 | Active T30 A | Analyte | 3.82 | 100394.141 | 2828198 | 54554.977 | 184.024 |  | 13.21
[17 empty] | 20240206_HepCl_101 | Active T30 B | Analyte | 3.82 | 65320.938 | 1655156 | 57943.566 | 112.732 |  | 8.08
[17 empty] | 20240206_HepCl_102 | Active T30 C | Analyte | 3.82 | 69604.961 | 1912436 | 55111.348 | 126.299 |  | 9.05
[17 empty] | 20240206_HepCl_103 | Active T60 A | Analyte | 3.82 | 101476.094 | 2799021 | 59652.234 | 170.113 |  | 12.21
[17 empty] | 20240206_HepCl_104 | Active T60 B | Analyte | 3.82 | 77268.211 | 1955454 | 59296.051 | 130.309 |  | 9.34
[17 empty] | 20240206_HepCl_105 | Active T60 C | Analyte | 3.82 | 106937.984 | 2930513 | 65826.789 | 162.454 |  | 11.66
[17 empty] | 20240206_HepCl_109 | Active T120 A | Analyte | 3.82 | 93797.734 | 2549652 | 70924.125 | 132.251 |  | 9.48
[17 empty] | 20240206_HepCl_110 | Active T120 B | Analyte | 3.82 | 84825.398 | 2274040 | 72407.617 | 117.15 |  | 8.4
[17 empty] | 20240206_HepCl_111 | Active T120 C | Analyte | 3.82 | 83003.328 | 2307404 | 59619.336 | 139.222 |  | 9.98
[17 empty] | 20240206_HepCl_112 | Active T240 A | Analyte | 3.82 | 100694.75 | 2752351 | 65011.285 | 154.888 |  | 11.11
[17 empty] | 20240206_HepCl_113 | Active T240 B | Analyte | 3.82 | 113709.984 | 3235892 | 56173.52 | 202.426 |  | 14.53
[17 empty] | 20240206_HepCl_114 | Active T240 C | Analyte | 3.82 | 78901.656 | 2163128 | 60057.676 | 131.376 |  | 9.42
[17 empty] | 20240206_HepCl_118 | Media T0 A | Analyte | 3.82 | 100251.555 | 2701375 | 61092.672 | 164.098 |  | 11.77
[17 empty] | 20240206_HepCl_119 | Media T0 B | Analyte | 3.82 | 90407.914 | 2479259 | 57920.395 | 156.09 |  | 11.2
[17 empty] | 20240206_HepCl_120 | Media T0 C | Analyte | 3.82 | 78154.32 | 2074097 | 49501.477 | 157.883 |  | 11.33
[17 empty] | 20240206_HepCl_121 | Media T15 A | Analyte | 3.82 | 107893.641 | 2836428 | 50702.512 | 212.797 |  | 15.28
[17 empty] | 20240206_HepCl_122 | Media T15 B | Analyte | 3.82 | 73800.031 | 1945417 | 49429.43 | 149.304 |  | 10.71
[17 empty] | 20240206_HepCl_123 | Media T15 C | Analyte | 3.82 | 77809.773 | 2146174 | 58678.934 | 132.603 |  | 9.51
[17 empty] | 20240206_HepCl_125 | Media T30 A | Analyte | 3.82 | 72522.531 | 1973974 | 61647.859 | 117.64 |  | 8.43
[17 empty] | 20240206_HepCl_126 | Media T30 B | Analyte | 3.82 | 48810.473 | 1231894 | 61024.07 | 79.986 |  | 5.72
[17 empty] | 20240206_HepCl_127 | Media T30 C | Analyte | 3.82 | 67735.219 | 1841966 | 61411.234 | 110.298 |  | 7.9
[17 empty] | 20240206_HepCl_128 | Media T60 A | Analyte | 3.82 | 103834.203 | 2896273 | 61999.914 | 167.475 |  | 12.02
[17 empty] | 20240206_HepCl_129 | Media T60 B | Analyte | 3.82 | 88792.891 | 2441441 | 50020.82 | 177.512 |  | 12.74
[17 empty] | 20240206_HepCl_130 | Media T60 C | Analyte | 3.82 | 63867.574 | 1681176 | 62273.973 | 102.559 |  | 7.35
[17 empty] | 20240206_HepCl_134 | Media T120 A | Analyte | 3.82 | 89899.516 | 2354959 | 67364.336 | 133.453 |  | 9.57
[17 empty] | 20240206_HepCl_135 | Media T120 B | Analyte | 3.82 | 69891.273 | 1889217 | 74916.945 | 93.292 |  | 6.68
[17 empty] | 20240206_HepCl_136 | Media T120 C | Analyte | 3.82 | 70295.148 | 1937328 | 65483.613 | 107.348 |  | 7.69
[17 empty] | 20240206_HepCl_137 | Media T240 A | Analyte | 3.82 | 86548.836 | 2110591 | 57828.598 | 149.664 |  | 10.74
[17 empty] | 20240206_HepCl_138 | Media T240 B | Analyte | 3.82 | 107041.172 | 2904692 | 60564.098 | 176.74 |  | 12.68
[17 empty] | 20240206_HepCl_139 | Media T240 C | Analyte | 3.82 | 86867.195 | 2313675 | 61393.453 | 141.493 |  | 10.15
[17 empty] | 20240206_HepCl_143 | Inactive T0 A | Analyte | 3.82 | 69812.531 | 1812805 | 51183.547 | 136.396 |  | 9.78
[17 empty] | 20240206_HepCl_144 | Inactive T0 B | Analyte | 3.82 | 47744.313 | 1223507 | 60897.484 | 78.401 |  | 5.61
[17 empty] | 20240206_HepCl_145 | Inactive T0 C | Analyte | 3.82 | 54972.5 | 1454558 | 57956.52 | 94.851 |  | 6.79
[17 empty] | 20240206_HepCl_146 | Inactive T240 A | Analyte | 3.82 | 86538.484 | 2347388 | 54418.652 | 159.024 |  | 11.41
[17 empty] | 20240206_HepCl_147 | Inactive T240 B | Analyte | 3.82 | 64077.941 | 1700892 | 53906.863 | 118.868 |  | 8.52
[17 empty] | 20240206_HepCl_148 | Inactive T240 C | Analyte | 3.82 | 83679.945 | 2268970 | 59785.922 | 139.966 |  | 10.04
[17 empty] | 20240206_HepCl_004 | Crash Blank | Blank | 3.9 | 85.059 | 1917 | 129.917 | 65.472 |  | 4.68
[17 empty] | 20240206_HepCl_003 | Matrix Blank | Blank | 3.9 | 90.429 | 1328 | 160.252 | 56.429 |  | 4.03
[17 empty] | 20240206_HepCl_001 | Blank | Blank | 3.9 | 83.006 | 1481 | 81.312 | 102.083 |  | 7.31
[17 empty] | 20240206_HepCl_002 | Blank | Blank | 3.9 | 105.249 | 1243 | 121.729 | 86.462 |  | 6.19
[17 empty] | 20240206_HepCl_005 | Blank | Blank | 3.89 | 81.687 | 1547 | 157.713 | 51.795 |  | 3.69
[17 empty] | 20240206_HepCl_006 | Blank | Blank | 3.89 | 81.964 | 1491 | 23.43 | 349.825 |  | 25.14
[17 empty] | 20240206_HepCl_013 | Blank | Blank | 3.89 | 74.91 | 1535 | 119.286 | 62.799 |  | 4.49
[17 empty] | 20240206_HepCl_020 | Blank | Blank | 3.89 | 90.015 | 1678 | 185.395 | 48.553 |  | 3.46
[17 empty] | 20240206_HepCl_026 | Blank | Blank | 3.89 | 67.843 | 1608 | 79.602 | 85.228 |  | 6.1
[17 empty] | 20240206_HepCl_027 | Blank | Blank | 3.89 | 88.088 | 1727 | 189.619 | 46.455 |  | 3.31
[17 empty] | 20240206_HepCl_032 | Blank | Blank | 3.89 | 65.47 | 1652 | 0.622 | 10525.723 |  | 757.31
[17 empty] | 20240206_HepCl_033 | Blank | Blank | 3.89 | 69.892 | 1803 | 148.775 | 46.978 |  | 3.35
[17 empty] | 20240206_HepCl_038 | Blank | Blank | 3.89 | 95.655 | 1716 | 193.399 | 49.46 |  | 3.53
[17 empty] | 20240206_HepCl_043 | Blank | Blank | 3.88 | 95.703 | 1741 | 91.56 | 104.525 |  | 7.49
[17 empty] | 20240206_HepCl_048 | Blank | Blank | 3.89 | 72.437 | 1606 | 45.974 | 157.561 |  | 11.3
[17 empty] | 20240206_HepCl_049 | Blank | Blank | 3.89 | 102.721 | 1862 | 30.281 | 339.226 |  | 24.38
[17 empty] | 20240206_HepCl_056 | Blank | Blank | 3.88 | 89.901 | 1862 | 41.153 | 218.456 |  | 15.69
[17 empty] | 20240206_HepCl_063 | Blank | Blank | 3.89 | 89.378 | 1697 | 128.157 | 69.741 |  | 4.99
[17 empty] | 20240206_HepCl_069 | Blank | Blank | 3.9 | 63.475 | 1854 | 147.617 | 43 |  | 3.06
[17 empty] | 20240206_HepCl_070 | Blank | Blank | 3.89 | 78.396 | 1431 | 171.114 | 45.815 |  | 3.26
[17 empty] | 20240206_HepCl_075 | Blank | Blank | 3.88 | 85.324 | 1906 | 212.657 | 40.123 |  | 2.85
[17 empty] | 20240206_HepCl_076 | Blank | Blank | 3.89 | 113.44 | 1819 | 110.334 | 102.815 |  | 7.37
[17 empty] | 20240206_HepCl_081 | Blank | Blank | 3.89 | 114.742 | 1644 | 184.305 | 62.257 |  | 4.45
[17 empty] | 20240206_HepCl_086 | Blank | Blank | 3.9 | 59.277 | 1597 | 48.337 | 122.633 |  | 8.79
[17 empty] | 20240206_HepCl_091 | Blank | Blank | 3.89 | 66.831 | 1575 | 141.691 | 47.167 |  | 3.36
[17 empty] | 20240206_HepCl_092 | Blank | Blank | 3.89 | 140.583 | 1667 | 193.432 | 72.678 |  | 5.2
[17 empty] | 20240206_HepCl_099 | Blank | Blank | 3.89 | 66.924 | 1651 | 74.173 | 90.227 |  | 6.46
[17 empty] | 20240206_HepCl_106 | Blank | Blank | 3.9 | 80.859 | 1522 | 178.652 | 45.261 |  | 3.22
[17 empty] | 20240206_HepCl_108 | Blank | Blank | 3.89 | 99.444 | 1228 | 145.36 | 68.412 |  | 4.89
[17 empty] | 20240206_HepCl_115 | Blank | Blank | 3.89 | 138.643 | 3232 | 192.55 | 72.004 |  | 5.15
[17 empty] | 20240206_HepCl_117 | Blank | Blank | 3.89 | 81.174 | 1365 | 124.869 | 65.007 |  | 4.64
[17 empty] | 20240206_HepCl_124 | Blank | Blank | 3.9 | 75.368 | 1168 | 173.451 | 43.452 |  | 3.09
[17 empty] | 20240206_HepCl_131 | Blank | Blank | 3.9 | 91.398 | 1266 | 91.41 | 99.987 |  | 7.16
[17 empty] | 20240206_HepCl_133 | Blank | Blank | 3.74 | 10.297 | 251 | 139.391 | 7.387 |  | 0.5
[17 empty] | 20240206_HepCl_140 | Blank | Blank | 3.89 | 49.966 | 1137 | 54.732 | 91.292 |  | 6.54
[17 empty] | 20240206_HepCl_142 | Blank | Blank | 3.73 | 4.838 | 165 | 122.332 | 3.955 |  | 0.25
[17 empty] | 20240206_HepCl_149 | Blank | Blank | 3.9 | 68.14 | 1366 | 228.384 | 29.836 |  | 2.11
[17 empty] | 20240206_HepCl_150 | Blank | Blank | 3.89 | 61.63 | 1495 | 26.952 | 228.666 |  | 16.42
[17 empty] | 20240206_HepCl_151 | Blank | Blank | 3.89 | 61.384 | 1342 | 176.011 | 34.875 |  | 2.48
Compound 2: M3PFHxS
[17 empty] | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
[17 empty] | 20240206_HepCl_007 | CC01 | Standard | 3.82 | 13228.613 | 413852 |  | 13228.613 | 100 | 102.85 | 2.9
[17 empty] | 20240206_HepCl_050 | CC01 | Standard | 3.82 | 13672.755 | 422542 |  | 13672.755 | 100 | 106.31 | 6.3
[17 empty] | 20240206_HepCl_008 | CC02 | Standard | 3.82 | 13060.495 | 418054 |  | 13060.495 | 100 | 101.55 | 1.5
[17 empty] | 20240206_HepCl_051 | CC02 | Standard | 3.82 | 14006.428 | 444985 |  | 14006.428 | 100 | 108.9 | 8.9
[17 empty] | 20240206_HepCl_009 | CC03 | Standard | 3.82 | 12383.586 | 377565 |  | 12383.586 | 100 | 96.28 | -3.7
[17 empty] | 20240206_HepCl_052 | CC03 | Standard | 3.82 | 12324.795 | 377716 |  | 12324.795 | 100 | 95.83 | -4.2
[17 empty] | 20240206_HepCl_010 | CC04 | Standard | 3.82 | 14313.511 | 451057 |  | 14313.511 | 100 | 111.29 | 11.3
[17 empty] | 20240206_HepCl_034 | CC04 | Standard | 3.81 | 11906.045 | 371992 |  | 11906.045 | 100 | 92.57 | -7.4
[17 empty] | 20240206_HepCl_039 | CC04 | Standard | 3.82 | 12918.972 | 393904 |  | 12918.972 | 100 | 100.45 | 0.4
[17 empty] | 20240206_HepCl_044 | CC04 | Standard | 3.81 | 12008.481 | 363469 |  | 12008.481 | 100 | 93.37 | -6.6
[17 empty] | 20240206_HepCl_053 | CC04 | Standard | 3.82 | 12797.363 | 393599 |  | 12797.363 | 100 | 99.5 | -0.5
[17 empty] | 20240206_HepCl_077 | CC04 | Standard | 3.82 | 10298.73 | 301611 |  | 10298.73 | 100 | 80.07 | -19.9
[17 empty] | 20240206_HepCl_082 | CC04 | Standard | 3.82 | 12015.393 | 359388 |  | 12015.393 | 100 | 93.42 | -6.6
[17 empty] | 20240206_HepCl_087 | CC04 | Standard | 3.82 | 15304.867 | 492807 |  | 15304.867 | 100 | 119 | 19
[17 empty] | 20240206_HepCl_011 | CC05 | Standard | 3.82 | 10278.519 | 300601 |  | 10278.519 | 100 | 79.92 | -20.1
[17 empty] | 20240206_HepCl_035 | CC05 | Standard | 3.81 | 12084.42 | 372086 |  | 12084.42 | 100 | 93.96 | -6
[17 empty] | 20240206_HepCl_040 | CC05 | Standard | 3.82 | 11582.531 | 353675 |  | 11582.531 | 100 | 90.05 | -9.9
[17 empty] | 20240206_HepCl_045 | CC05 | Standard | 3.82 | 11747.538 | 362078 |  | 11747.538 | 100 | 91.34 | -8.7
[17 empty] | 20240206_HepCl_054 | CC05 | Standard | 3.82 | 14087.408 | 443564 |  | 14087.408 | 100 | 109.53 | 9.5
[17 empty] | 20240206_HepCl_078 | CC05 | Standard | 3.82 | 13693.741 | 424112 |  | 13693.741 | 100 | 106.47 | 6.5
[17 empty] | 20240206_HepCl_083 | CC05 | Standard | 3.82 | 15032.213 | 466578 |  | 15032.213 | 100 | 116.88 | 16.9
[17 empty] | 20240206_HepCl_088 | CC05 | Standard | 3.82 | 13440.987 | 419437 |  | 13440.987 | 100 | 104.5 | 4.5
[17 empty] | 20240206_HepCl_012 | CC06 | Standard | 3.82 | 12826.483 | 409590 |  | 12826.483 | 100 | 99.73 | -0.3
[17 empty] | 20240206_HepCl_036 | CC06 | Standard | 3.81 | 11320.455 | 343551 |  | 11320.455 | 100 | 88.02 | -12
[17 empty] | 20240206_HepCl_041 | CC06 | Standard | 3.82 | 13133.145 | 404806 |  | 13133.145 | 100 | 102.11 | 2.1
[17 empty] | 20240206_HepCl_046 | CC06 | Standard | 3.81 | 11973.123 | 361518 |  | 11973.123 | 100 | 93.09 | -6.9
[17 empty] | 20240206_HepCl_055 | CC06 | Standard | 3.82 | 12849.26 | 390870 |  | 12849.26 | 100 | 99.9 | -0.1
[17 empty] | 20240206_HepCl_079 | CC06 | Standard | 3.82 | 13200.004 | 412523 |  | 13200.004 | 100 | 102.63 | 2.6
[17 empty] | 20240206_HepCl_084 | CC06 | Standard | 3.82 | 12810.994 | 387593 |  | 12810.994 | 100 | 99.61 | -0.4
[17 empty] | 20240206_HepCl_089 | CC06 | Standard | 3.82 | 13731.132 | 432971 |  | 13731.132 | 100 | 106.76 | 6.8
[17 empty] | 20240206_HepCl_014 | CC07 | Standard | 3.82 | 13677.68 | 437952 |  | 13677.68 | 100 | 106.34 | 6.3
[17 empty] | 20240206_HepCl_057 | CC07 | Standard | 3.82 | 11294.601 | 335217 |  | 11294.601 | 100 | 87.82 | -12.2
[17 empty] | 20240206_HepCl_015 | CC08 | Standard | 3.82 | 12671.938 | 390296 |  | 12671.938 | 100 | 98.52 | -1.5
[17 empty] | 20240206_HepCl_058 | CC08 | Standard | 3.82 | 13495.462 | 416955 |  | 13495.462 | 100 | 104.93 | 4.9
[17 empty] | 20240206_HepCl_016 | CC09 | Standard | 3.82 | 9804.405 | 296585 |  | 9804.405 | 100 | 76.23 | -23.8
[17 empty] | 20240206_HepCl_059 | CC09 | Standard | 3.82 | 12311.349 | 379701 |  | 12311.349 | 100 | 95.72 | -4.3
[17 empty] | 20240206_HepCl_017 | CC10 | Standard | 3.82 | 14090.989 | 442205 |  | 14090.989 | 100 | 109.56 | 9.6
[17 empty] | 20240206_HepCl_060 | CC10 | Standard | 3.82 | 12715.516 | 385560 |  | 12715.516 | 100 | 98.86 | -1.1
[17 empty] | 20240206_HepCl_018 | CC11 | Standard | 3.82 | 12546.156 | 389787 |  | 12546.156 | 100 | 97.55 | -2.5
[17 empty] | 20240206_HepCl_061 | CC11 | Standard | 3.82 | 13536.77 | 411755 |  | 13536.77 | 100 | 105.25 | 5.2
[17 empty] | 20240206_HepCl_019 | CC12 | Standard | 3.82 | 12496.379 | 384143 |  | 12496.379 | 100 | 97.16 | -2.8
[17 empty] | 20240206_HepCl_062 | CC12 | Standard | 3.82 | 13279.041 | 410264 |  | 13279.041 | 100 | 103.24 | 3.2
[17 empty] | 20240206_HepCl_107 | CC12 | Standard | 3.82 | 12644.87 | 393543 |  | 12644.87 | 100 | 98.31 | -1.7
[17 empty] | 20240206_HepCl_116 | CC12 | Standard | 3.82 | 13688.918 | 427866 |  | 13688.918 | 100 | 106.43 | 6.4
[17 empty] | 20240206_HepCl_132 | CC12 | Standard | 3.82 | 13353.789 | 412544 |  | 13353.789 | 100 | 103.83 | 3.8
[17 empty] | 20240206_HepCl_141 | CC12 | Standard | 3.82 | 14119.17 | 438989 |  | 14119.17 | 100 | 109.78 | 9.8
[17 empty] | 20240206_HepCl_021 | CC13 | Standard | 3.82 | 12645.211 | 393316 |  | 12645.211 | 100 | 98.32 | -1.7
[17 empty] | 20240206_HepCl_037 | CC13 | Standard | 3.82 | 12121.979 | 372326 |  | 12121.979 | 100 | 94.25 | -5.8
[17 empty] | 20240206_HepCl_042 | CC13 | Standard | 3.82 | 12904.606 | 398315 |  | 12904.606 | 100 | 100.33 | 0.3
[17 empty] | 20240206_HepCl_047 | CC13 | Standard | 3.82 | 12636.182 | 381013 |  | 12636.182 | 100 | 98.25 | -1.8
[17 empty] | 20240206_HepCl_064 | CC13 | Standard | 3.82 | 12511.845 | 372955 |  | 12511.845 | 100 | 97.28 | -2.7
[17 empty] | 20240206_HepCl_080 | CC13 | Standard | 3.82 | 12701.495 | 388987 |  | 12701.495 | 100 | 98.75 | -1.2
[17 empty] | 20240206_HepCl_085 | CC13 | Standard | 3.82 | 12612.149 | 388733 |  | 12612.149 | 100 | 98.06 | -1.9
[17 empty] | 20240206_HepCl_090 | CC13 | Standard | 3.82 | 11894.017 | 349842 |  | 11894.017 | 100 | 92.48 | -7.5
[17 empty] | 20240206_HepCl_022 | CC14 | Standard | 3.82 | 13375.271 | 417579 |  | 13375.271 | 100 | 103.99 | 4
[17 empty] | 20240206_HepCl_065 | CC14 | Standard | 3.82 | 14586.896 | 459507 |  | 14586.896 | 100 | 113.41 | 13.4
[17 empty] | 20240206_HepCl_023 | CC15 | Standard | 3.82 | 11024.747 | 338200 |  | 11024.747 | 100 | 85.72 | -14.3
[17 empty] | 20240206_HepCl_066 | CC15 | Standard | 3.82 | 13164.907 | 405608 |  | 13164.907 | 100 | 102.36 | 2.4
[17 empty] | 20240206_HepCl_024 | CC16 | Standard | 3.82 | 15599.353 | 485161 |  | 15599.353 | 100 | 121.29 | 21.3
[17 empty] | 20240206_HepCl_067 | CC16 | Standard | 3.82 | 15479.135 | 480047 |  | 15479.135 | 100 | 120.35 | 20.4
[17 empty] | 20240206_HepCl_025 | CC17 | Standard | 3.82 | 11725.054 | 355157 |  | 11725.054 | 100 | 91.16 | -8.8
[17 empty] | 20240206_HepCl_068 | CC17 | Standard | 3.82 | 12682.382 | 373498 |  | 12682.382 | 100 | 98.61 | -1.4
[17 empty] | 20240206_HepCl_028 | CCQC A | QC | 3.82 | 15079.548 | 473390 |  | 15079.548 | 100 | 117.24 | 17.2
[17 empty] | 20240206_HepCl_071 | CCQC A | QC | 3.82 | 17512.383 | 543966 |  | 17512.383 | 100 | 136.16 | 36.2
[17 empty] | 20240206_HepCl_029 | CCQC B | QC | 3.82 | 17196.908 | 543029 |  | 17196.908 | 100 | 133.71 | 33.7
[17 empty] | 20240206_HepCl_072 | CCQC B | QC | 3.82 | 17804.043 | 566058 |  | 17804.043 | 100 | 138.43 | 38.4
[17 empty] | 20240206_HepCl_030 | CCQC C | QC | 3.82 | 14326.982 | 457026 |  | 14326.982 | 100 | 111.39 | 11.4
[17 empty] | 20240206_HepCl_073 | CCQC C | QC | 3.82 | 15293.799 | 475924 |  | 15293.799 | 100 | 118.91 | 18.9
[17 empty] | 20240206_HepCl_031 | CCQC D | QC | 3.82 | 13883.457 | 426754 |  | 13883.457 | 100 | 107.94 | 7.9
[17 empty] | 20240206_HepCl_074 | CCQC D | QC | 3.82 | 14440.207 | 442584 |  | 14440.207 | 100 | 112.27 | 12.3
[17 empty] | 20240206_HepCl_093 | Active T0 A | Analyte | 3.82 | 57196.434 | 1753631 |  | 57196.434 | 100 | 444.7 | 344.7
[17 empty] | 20240206_HepCl_094 | Active T0 B | Analyte | 3.82 | 52622.742 | 1567755 |  | 52622.742 | 100 | 409.14 | 309.1
[17 empty] | 20240206_HepCl_095 | Active T0 C | Analyte | 3.82 | 62977.305 | 2001805 |  | 62977.305 | 100 | 489.65 | 389.7
[17 empty] | 20240206_HepCl_096 | Active T15 A | Analyte | 3.82 | 54653.637 | 1667995 |  | 54653.637 | 100 | 424.93 | 324.9
[17 empty] | 20240206_HepCl_097 | Active T15 B | Analyte | 3.82 | 56251.129 | 1713435 |  | 56251.129 | 100 | 437.35 | 337.4
[17 empty] | 20240206_HepCl_098 | Active T15 C | Analyte | 3.82 | 46738.797 | 1393550 |  | 46738.797 | 100 | 363.4 | 263.4
[17 empty] | 20240206_HepCl_100 | Active T30 A | Analyte | 3.82 | 54554.977 | 1695409 |  | 54554.977 | 100 | 424.17 | 324.2
[17 empty] | 20240206_HepCl_101 | Active T30 B | Analyte | 3.82 | 57943.566 | 1843299 |  | 57943.566 | 100 | 450.51 | 350.5
[17 empty] | 20240206_HepCl_102 | Active T30 C | Analyte | 3.82 | 55111.348 | 1760062 |  | 55111.348 | 100 | 428.49 | 328.5
[17 empty] | 20240206_HepCl_103 | Active T60 A | Analyte | 3.82 | 59652.234 | 1889510 |  | 59652.234 | 100 | 463.8 | 363.8
[17 empty] | 20240206_HepCl_104 | Active T60 B | Analyte | 3.82 | 59296.051 | 1854851 |  | 59296.051 | 100 | 461.03 | 361
[17 empty] | 20240206_HepCl_105 | Active T60 C | Analyte | 3.82 | 65826.789 | 2056766 |  | 65826.789 | 100 | 511.81 | 411.8
[17 empty] | 20240206_HepCl_109 | Active T120 A | Analyte | 3.82 | 70924.125 | 2216592 |  | 70924.125 | 100 | 551.44 | 451.4
[17 empty] | 20240206_HepCl_110 | Active T120 B | Analyte | 3.82 | 72407.617 | 2285308 |  | 72407.617 | 100 | 562.97 | 463
[17 empty] | 20240206_HepCl_111 | Active T120 C | Analyte | 3.82 | 59619.336 | 1828233 |  | 59619.336 | 100 | 463.54 | 363.5
[17 empty] | 20240206_HepCl_112 | Active T240 A | Analyte | 3.82 | 65011.285 | 2034666 |  | 65011.285 | 100 | 505.46 | 405.5
[17 empty] | 20240206_HepCl_113 | Active T240 B | Analyte | 3.82 | 56173.52 | 1712859 |  | 56173.52 | 100 | 436.75 | 336.8
[17 empty] | 20240206_HepCl_114 | Active T240 C | Analyte | 3.82 | 60057.676 | 1861754 |  | 60057.676 | 100 | 466.95 | 367
[17 empty] | 20240206_HepCl_118 | Media T0 A | Analyte | 3.82 | 61092.672 | 1890146 |  | 61092.672 | 100 | 475 | 375
[17 empty] | 20240206_HepCl_119 | Media T0 B | Analyte | 3.82 | 57920.395 | 1824850 |  | 57920.395 | 100 | 450.33 | 350.3
[17 empty] | 20240206_HepCl_120 | Media T0 C | Analyte | 3.82 | 49501.477 | 1486998 |  | 49501.477 | 100 | 384.88 | 284.9
[17 empty] | 20240206_HepCl_121 | Media T15 A | Analyte | 3.82 | 50702.512 | 1553175 |  | 50702.512 | 100 | 394.21 | 294.2
[17 empty] | 20240206_HepCl_122 | Media T15 B | Analyte | 3.82 | 49429.43 | 1506484 |  | 49429.43 | 100 | 384.32 | 284.3
[17 empty] | 20240206_HepCl_123 | Media T15 C | Analyte | 3.82 | 58678.934 | 1838720 |  | 58678.934 | 100 | 456.23 | 356.2
[17 empty] | 20240206_HepCl_125 | Media T30 A | Analyte | 3.82 | 61647.859 | 1945396 |  | 61647.859 | 100 | 479.31 | 379.3
[17 empty] | 20240206_HepCl_126 | Media T30 B | Analyte | 3.82 | 61024.07 | 1875551 |  | 61024.07 | 100 | 474.46 | 374.5
[17 empty] | 20240206_HepCl_127 | Media T30 C | Analyte | 3.82 | 61411.234 | 1934272 |  | 61411.234 | 100 | 477.47 | 377.5
[17 empty] | 20240206_HepCl_128 | Media T60 A | Analyte | 3.82 | 61999.914 | 1957041 |  | 61999.914 | 100 | 482.05 | 382.1
[17 empty] | 20240206_HepCl_129 | Media T60 B | Analyte | 3.82 | 50020.82 | 1529021 |  | 50020.82 | 100 | 388.91 | 288.9
[17 empty] | 20240206_HepCl_130 | Media T60 C | Analyte | 3.82 | 62273.973 | 1903030 |  | 62273.973 | 100 | 484.18 | 384.2
[17 empty] | 20240206_HepCl_134 | Media T120 A | Analyte | 3.82 | 67364.336 | 2137155 |  | 67364.336 | 100 | 523.76 | 423.8
[17 empty] | 20240206_HepCl_135 | Media T120 B | Analyte | 3.82 | 74916.945 | 2339204 |  | 74916.945 | 100 | 582.48 | 482.5
[17 empty] | 20240206_HepCl_136 | Media T120 C | Analyte | 3.82 | 65483.613 | 2024615 |  | 65483.613 | 100 | 509.14 | 409.1
[17 empty] | 20240206_HepCl_137 | Media T240 A | Analyte | 3.82 | 57828.598 | 1748022 |  | 57828.598 | 100 | 449.62 | 349.6
[17 empty] | 20240206_HepCl_138 | Media T240 B | Analyte | 3.82 | 60564.098 | 1891748 |  | 60564.098 | 100 | 470.89 | 370.9
[17 empty] | 20240206_HepCl_139 | Media T240 C | Analyte | 3.82 | 61393.453 | 1952005 |  | 61393.453 | 100 | 477.34 | 377.3
[17 empty] | 20240206_HepCl_143 | Inactive T0 A | Analyte | 3.82 | 51183.547 | 1550672 |  | 51183.547 | 100 | 397.95 | 298
[17 empty] | 20240206_HepCl_144 | Inactive T0 B | Analyte | 3.82 | 60897.484 | 1924150 |  | 60897.484 | 100 | 473.48 | 373.5
[17 empty] | 20240206_HepCl_145 | Inactive T0 C | Analyte | 3.82 | 57956.52 | 1805876 |  | 57956.52 | 100 | 450.61 | 350.6
[17 empty] | 20240206_HepCl_146 | Inactive T240 A | Analyte | 3.82 | 54418.652 | 1682014 |  | 54418.652 | 100 | 423.11 | 323.1
[17 empty] | 20240206_HepCl_147 | Inactive T240 B | Analyte | 3.82 | 53906.863 | 1653033 |  | 53906.863 | 100 | 419.13 | 319.1
[17 empty] | 20240206_HepCl_148 | Inactive T240 C | Analyte | 3.82 | 59785.922 | 1831872 |  | 59785.922 | 100 | 464.84 | 364.8
[17 empty] | 20240206_HepCl_004 | Crash Blank | Blank | 3.75 | 129.917 | 978 |  | 129.917 | 100 | 1.01 | -99
[17 empty] | 20240206_HepCl_003 | Matrix Blank | Blank | 3.77 | 160.252 | 829 |  | 160.252 | 100 | 1.25 | -98.8
[17 empty] | 20240206_HepCl_001 | Blank | Blank | 3.78 | 81.312 | 812 |  | 81.312 | 100 | 0.63 | -99.4
[17 empty] | 20240206_HepCl_002 | Blank | Blank | 3.8 | 121.729 | 673 |  | 121.729 | 100 | 0.95 | -99.1
[17 empty] | 20240206_HepCl_005 | Blank | Blank | 3.81 | 157.713 | 840 |  | 157.713 | 100 | 1.23 | -98.8
[17 empty] | 20240206_HepCl_006 | Blank | Blank | 3.78 | 23.43 | 417 |  | 23.43 | 100 | 0.18 | -99.8
[17 empty] | 20240206_HepCl_013 | Blank | Blank | 3.76 | 119.286 | 826 |  | 119.286 | 100 | 0.93 | -99.1
[17 empty] | 20240206_HepCl_020 | Blank | Blank | 3.77 | 185.395 | 1192 |  | 185.395 | 100 | 1.44 | -98.6
[17 empty] | 20240206_HepCl_026 | Blank | Blank | 3.78 | 79.602 | 700 |  | 79.602 | 100 | 0.62 | -99.4
[17 empty] | 20240206_HepCl_027 | Blank | Blank | 3.77 | 189.619 | 1152 |  | 189.619 | 100 | 1.47 | -98.5
[17 empty] | 20240206_HepCl_032 | Blank | Blank | 3.73 | 0.622 | 35 |  | 0.622 | 100 | 0 | -100
[17 empty] | 20240206_HepCl_033 | Blank | Blank | 3.77 | 148.775 | 910 |  | 148.775 | 100 | 1.16 | -98.8
[17 empty] | 20240206_HepCl_038 | Blank | Blank | 3.78 | 193.399 | 1174 |  | 193.399 | 100 | 1.5 | -98.5
[17 empty] | 20240206_HepCl_043 | Blank | Blank | 3.77 | 91.56 | 786 |  | 91.56 | 100 | 0.71 | -99.3
[17 empty] | 20240206_HepCl_048 | Blank | Blank | 3.77 | 45.974 | 527 |  | 45.974 | 100 | 0.36 | -99.6
[17 empty] | 20240206_HepCl_049 | Blank | Blank | 3.79 | 30.281 | 527 |  | 30.281 | 100 | 0.24 | -99.8
[17 empty] | 20240206_HepCl_056 | Blank | Blank | 3.78 | 41.153 | 652 |  | 41.153 | 100 | 0.32 | -99.7
[17 empty] | 20240206_HepCl_063 | Blank | Blank | 3.8 | 128.157 | 907 |  | 128.157 | 100 | 1 | -99
[17 empty] | 20240206_HepCl_069 | Blank | Blank | 3.79 | 147.617 | 971 |  | 147.617 | 100 | 1.15 | -98.9
[17 empty] | 20240206_HepCl_070 | Blank | Blank | 3.78 | 171.114 | 972 |  | 171.114 | 100 | 1.33 | -98.7
[17 empty] | 20240206_HepCl_075 | Blank | Blank | 3.78 | 212.657 | 1293 |  | 212.657 | 100 | 1.65 | -98.3
[17 empty] | 20240206_HepCl_076 | Blank | Blank | 3.77 | 110.334 | 1025 |  | 110.334 | 100 | 0.86 | -99.1
[17 empty] | 20240206_HepCl_081 | Blank | Blank | 3.8 | 184.305 | 1099 |  | 184.305 | 100 | 1.43 | -98.6
[17 empty] | 20240206_HepCl_086 | Blank | Blank | 3.8 | 48.337 | 697 |  | 48.337 | 100 | 0.38 | -99.6
[17 empty] | 20240206_HepCl_091 | Blank | Blank | 3.79 | 141.691 | 640 |  | 141.691 | 100 | 1.1 | -98.9
[17 empty] | 20240206_HepCl_092 | Blank | Blank | 3.78 | 193.432 | 872 |  | 193.432 | 100 | 1.5 | -98.5
[17 empty] | 20240206_HepCl_099 | Blank | Blank | 3.78 | 74.173 | 512 |  | 74.173 | 100 | 0.58 | -99.4
[17 empty] | 20240206_HepCl_106 | Blank | Blank | 3.77 | 178.652 | 1024 |  | 178.652 | 100 | 1.39 | -98.6
[17 empty] | 20240206_HepCl_108 | Blank | Blank | 3.66 | 145.36 | 734 |  | 145.36 | 100 | 1.13 | -98.9
[17 empty] | 20240206_HepCl_115 | Blank | Blank | 3.81 | 192.55 | 724 |  | 192.55 | 100 | 1.5 | -98.5
[17 empty] | 20240206_HepCl_117 | Blank | Blank | 3.66 | 124.869 | 828 |  | 124.869 | 100 | 0.97 | -99
[17 empty] | 20240206_HepCl_124 | Blank | Blank | 3.77 | 173.451 | 797 |  | 173.451 | 100 | 1.35 | -98.7
[17 empty] | 20240206_HepCl_131 | Blank | Blank | 3.75 | 91.41 | 723 |  | 91.41 | 100 | 0.71 | -99.3
[17 empty] | 20240206_HepCl_133 | Blank | Blank | 3.78 | 139.391 | 641 |  | 139.391 | 100 | 1.08 | -98.9
[17 empty] | 20240206_HepCl_140 | Blank | Blank | 3.78 | 54.732 | 682 |  | 54.732 | 100 | 0.43 | -99.6
[17 empty] | 20240206_HepCl_142 | Blank | Blank | 3.8 | 122.332 | 558 |  | 122.332 | 100 | 0.95 | -99
[17 empty] | 20240206_HepCl_149 | Blank | Blank | 3.77 | 228.384 | 845 |  | 228.384 | 100 | 1.78 | -98.2
[17 empty] | 20240206_HepCl_150 | Blank | Blank | 3.79 | 26.952 | 415 |  | 26.952 | 100 | 0.21 | -99.8
[17 empty] | 20240206_HepCl_151 | Blank | Blank | 3.79 | 176.011 | 908 |  | 176.011 | 100 | 1.37 | -98.6

## Na PFOA
Human Active Cells | [8 empty] | Media Only | [7 empty] | Compound 3: PFOA
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20240206_HepCl_007 | CC01 | Standard | 3.79 | 14689.672 | 349695 | 134058.391 | 10.958 | 0.175 | 0.34 | 93.2
0 | Active T0 A | 11.78 | 6.0499 | 424.08 | 402.96 | 21.241 | 5.271% |  | 0 | Media T0 A | 10.94 | 5.9759 | 393.84 | 355.8 | 34.1987 | 9.612% | 20240206_HepCl_050 | CC01 | Standard | 3.79 | 15336.969 | 336924 | 146270.547 | 10.485 | 0.175 | 0.3 | 72
| Active T0 B | 11.2 | 5.9994 | 403.2 | [5 empty] | Media T0 B | 9.61 | 5.8463 | 345.96 |  |  |  | 20240206_HepCl_008 | CC02 | Standard | 3.79 | 14954.547 | 335498 | 118457.406 | 12.624 | 0.281 | 0.47 | 66.7
| Active T0 C | 10.6 | 5.9444 | 381.6 | [5 empty] | Media T0 C | 9.1 | 5.7918 | 327.6 |  |  |  | 20240206_HepCl_051 | CC02 | Standard | 3.79 | 16400.055 | 407683 | 153555.141 | 10.68 | 0.281 | 0.32 | 12.5
15 | Active T15 A | 13.89 | 6.2147 | 500.04 | 454.56 | 43.5322 | 9.577% |  | 15 | Media T15 A | 13.97 | 6.2204 | 502.92 | 455.4 | 73.1536 | 16.06% | 20240206_HepCl_009 | CC03 | Standard | 3.79 | 16010.505 | 399784 | 146184.828 | 10.952 | 0.444 | 0.34 | -23.9
| Active T15 B | 12.51 | 6.11 | 450.36 | [5 empty] | Media T15 B | 13.67 | 6.1987 | 492.12 |  |  |  | 20240206_HepCl_052 | CC03 | Standard | 3.79 | 14273.191 | 344687 | 133897.094 | 10.66 | 0.444 | 0.31 | -29.1
| Active T15 C | 11.48 | 6.0241 | 413.28 | [5 empty] | Media T15 C | 10.31 | 5.9166 | 371.16 |  |  |  | 20240206_HepCl_010 | CC04 | Standard | 3.79 | 18814.031 | 485555 | 134612.875 | 13.976 | 0.713 | 0.58 | -19.3
30 | Active T30 A | 11.18 | 5.9976 | 402.48 | 382.08 | 18.6969 | 4.893% |  | 30 | Media T30 A | 9.37 | 5.821 | 337.32 | 302.28 | 41.3441 | 13.68% | 20240206_HepCl_034 | CC04 | Standard | 3.78 | 18807.299 | 455470 | 136284.063 | 13.8 | 0.713 | 0.56 | -21.2
| Active T30 B | 10.16 | 5.902 | 365.76 | [5 empty] | Media T30 B | 7.13 | 5.5478 | 256.68 |  |  |  | 20240206_HepCl_039 | CC04 | Standard | 3.79 | 20343.488 | 512426 | 149225.266 | 13.633 | 0.713 | 0.55 | -23.1
| Active T30 C | 10.5 | 5.9349 | 378 | [5 empty] | Media T30 C | 8.69 | 5.7457 | 312.84 |  |  |  | 20240206_HepCl_044 | CC04 | Standard | 3.79 | 17677.65 | 434824 | 146873.625 | 12.036 | 0.713 | 0.42 | -40.7
60 | Active T60 A | 14.57 | 6.2625 | 524.52 | 441.6 | 85.8133 | 19.43% |  | 60 | Media T60 A | 12.11 | 6.0776 | 435.96 | 425.64 | 14.2976 | 3.359% | 20240206_HepCl_053 | CC04 | Standard | 3.79 | 16063.406 | 364068 | 112335.555 | 14.299 | 0.713 | 0.6 | -15.7
| Active T60 B | 9.81 | 5.8669 | 353.16 | [5 empty] | Media T60 B | 11.37 | 6.0145 | 409.32 |  |  |  | 20240206_HepCl_077 | CC04 | Standard | 3.79 | 19587.879 | 493252 | 150515.109 | 13.014 | 0.713 | 0.5 | -29.9
| Active T60 C | 12.42 | 6.1028 | 447.12 | [5 empty] | Media T60 C | 11.99 | 6.0676 | 431.64 |  |  |  | 20240206_HepCl_082 | CC04 | Standard | 3.79 | 15517.276 | 358209 | 129994.75 | 11.937 | 0.713 | 0.41 | -41.8
120 | Active T120 A | 8.32 | 5.7022 | 299.52 | 283.2 | 32.0811 | 11.33% |  | 120 | Media T120 A | 5.41 | 5.2718 | 194.76 | 223.32 | 27.7582 | 12.43% | 20240206_HepCl_087 | CC04 | Standard | 3.79 | 17412.711 | 454139 | 157955.078 | 11.024 | 0.713 | 0.34 | -51.8
| Active T120 B | 8.44 | 5.7165 | 303.84 | [5 empty] | Media T120 B | 6.25 | 5.4161 | 225 |  |  |  | 20240206_HepCl_011 | CC05 | Standard | 3.79 | 30746.496 | 862894 | 134493.047 | 22.861 | 1.138 | 1.27 | 11.9
| Active T120 C | 6.84 | 5.5063 | 246.24 | [5 empty] | Media T120 C | 6.95 | 5.5223 | 250.2 |  |  |  | 20240206_HepCl_035 | CC05 | Standard | 3.78 | 29587.498 | 837094 | 135748.172 | 21.796 | 1.138 | 1.19 | 4.6
240 | Active T240 A | 8.21 | 5.6889 | 295.56 | 302.64 | 6.1552 | 2.034% |  | 240 | Media T240 A | 6.98 | 5.5266 | 251.28 | 282.96 | 30.0422 | 10.62% | 20240206_HepCl_040 | CC05 | Standard | 3.79 | 25224.449 | 708081 | 136307.844 | 18.506 | 1.138 | 0.93 | -18.2
| Active T240 B | 8.52 | 5.7259 | 306.72 | [5 empty] | Media T240 B | 7.96 | 5.6579 | 286.56 |  |  |  | 20240206_HepCl_045 | CC05 | Standard | 3.78 | 29065.047 | 801278 | 143390.891 | 20.27 | 1.138 | 1.07 | -6
| Active T240 C | 8.49 | 5.7224 | 305.64 | [5 empty] | Media T240 C | 8.64 | 5.7399 | 311.04 |  |  |  | 20240206_HepCl_054 | CC05 | Standard | 3.79 | 28396.881 | 758421 | 117635.977 | 24.14 | 1.138 | 1.37 | 20.7
[17 empty] | 20240206_HepCl_078 | CC05 | Standard | 3.79 | 25835.053 | 668365 | 137131.031 | 18.84 | 1.138 | 0.96 | -15.9
Human Inactive Cells | [16 empty] | 20240206_HepCl_083 | CC05 | Standard | 3.79 | 32774.672 | 913665 | 128168.617 | 25.572 | 1.138 | 1.49 | 30.6
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20240206_HepCl_088 | CC05 | Standard | 3.79 | 27265.17 | 740062 | 141630.219 | 19.251 | 1.138 | 0.99 | -13
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20240206_HepCl_012 | CC06 | Standard | 3.79 | 33982.055 | 948781 | 114462.625 | 29.688 | 1.825 | 1.81 | -0.9
0 | Inactive T0 A | 8.2 | 5.6877 | 295.2 | 267.48 | 25.0116 | 9.351% | [9 empty] | 20240206_HepCl_036 | CC06 | Standard | 3.79 | 35359.563 | 976728 | 130111.891 | 27.176 | 1.825 | 1.61 | -11.7
| Inactive T0 B | 7.24 | 5.5631 | 260.64 | [12 empty] | 20240206_HepCl_041 | CC06 | Standard | 3.79 | 39211.633 | 1135086 | 122278.68 | 32.067 | 1.825 | 2 | 9.4
| Inactive T0 C | 6.85 | 5.5078 | 246.6 | [12 empty] | 20240206_HepCl_046 | CC06 | Standard | 3.78 | 39930.25 | 1149416 | 135189.078 | 29.537 | 1.825 | 1.8 | -1.5
240 | Inactive T240 A | 9.13 | 5.7951 | 328.68 | 283.32 | 64.1487 | 22.64% | [9 empty] | 20240206_HepCl_055 | CC06 | Standard | 3.79 | 40665.832 | 1205194 | 142075.469 | 28.623 | 1.825 | 1.73 | -5.4
| Inactive T240 B | 6.61 | 5.4721 | 237.96 | [12 empty] | 20240206_HepCl_079 | CC06 | Standard | 3.79 | 37115.809 | 1071561 | 129202.703 | 28.727 | 1.825 | 1.73 | -5
| Inactive T240 C | 7.41 | 5.5863 | 266.76 | [12 empty] | 20240206_HepCl_084 | CC06 | Standard | 3.79 | 38892.891 | 1098905 | 135884.641 | 28.622 | 1.825 | 1.73 | -5.5
[17 empty] | 20240206_HepCl_089 | CC06 | Standard | 3.79 | 38642.188 | 1151668 | 131731.141 | 29.334 | 1.825 | 1.78 | -2.4
| Calibration Curve | [15 empty] | 20240206_HepCl_014 | CC07 | Standard | 3.79 | 54403.395 | 1644471 | 124790.297 | 43.596 | 2.919 | 2.9 | -0.6
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20240206_HepCl_057 | CC07 | Standard | 3.79 | 52478.242 | 1539972 | 117855.555 | 44.528 | 2.919 | 2.97 | 1.9
| 20240206_HepCl_011 | CC05 | 1.138 | 1.27 | 1.1587 | 0.2056 | 0.17743 | 0.64619 | 0.018234 | [7 empty] | 20240206_HepCl_015 | CC08 | Standard | 3.79 | 83895.859 | 2620377 | 129505.023 | 64.782 | 4.669 | 4.57 | -2.2
| 20240206_HepCl_035 | CC05 | 1.138 | 1.19 | [12 empty] | 20240206_HepCl_058 | CC08 | Standard | 3.79 | 87295.852 | 2678251 | 113011.789 | 77.245 | 4.669 | 5.54 | 18.8
| 20240206_HepCl_040 | CC05 | 1.138 | 0.93 | [12 empty] | 20240206_HepCl_016 | CC09 | Standard | 3.79 | 116868.016 | 3671228 | 117331.039 | 99.605 | 7.45 | 7.3 | -2
| 20240206_HepCl_045 | CC05 | 1.138 | 1.07 | [12 empty] | 20240206_HepCl_059 | CC09 | Standard | 3.79 | 110376.977 | 3406501 | 118829.922 | 92.887 | 7.45 | 6.77 | -9.1
| 20240206_HepCl_054 | CC05 | 1.138 | 1.37 | [12 empty] | 20240206_HepCl_017 | CC10 | Standard | 3.79 | 193112.781 | 6144241 | 125046.508 | 154.433 | 11.875 | 11.61 | -2.3
| 20240206_HepCl_078 | CC05 | 1.138 | 0.96 | [12 empty] | 20240206_HepCl_060 | CC10 | Standard | 3.79 | 226155.516 | 7283410 | 152427.219 | 148.37 | 11.875 | 11.13 | -6.3
| 20240206_HepCl_083 | CC05 | 1.138 | 1.49 | [12 empty] | 20240206_HepCl_018 | CC11 | Standard | 3.79 | 378762.781 | 12584447 | 119904.609 | 315.887 | 19.125 | 24.29 | 27
| 20240206_HepCl_088 | CC05 | 1.138 | 0.99 | [12 empty] | 20240206_HepCl_061 | CC11 | Standard | 3.79 | 378595.563 | 12403840 | 134633.641 | 281.204 | 19.125 | 21.56 | 12.8
[17 empty] | 20240206_HepCl_019 | CC12 | Standard | 3.79 | 543942.688 | 17636362 | 114519.438 | 474.978 | 30.5 | 36.78 | 20.6
| Blanks should be less than… | 0.32309 |  | nM | [12 empty] | 20240206_HepCl_062 | CC12 | Standard | 3.79 | 542615.563 | 17402342 | 130157.563 | 416.891 | 30.5 | 32.22 | 5.6
| eLOQ | 1.138 |  | nM | [12 empty] | 20240206_HepCl_107 | CC12 | Standard | 3.78 | 468996.781 | 14921748 | 116307.844 | 403.237 | 30.5 | 31.15 | 2.1
[17 empty] | 20240206_HepCl_116 | CC12 | Standard | 3.78 | 508351.656 | 16565745 | 136432.594 | 372.603 | 30.5 | 28.74 | -5.8
| QC FLAGS | [15 empty] | 20240206_HepCl_132 | CC12 | Standard | 3.79 | 557822.938 | 18217662 | 125093.547 | 445.925 | 30.5 | 34.5 | 13.1
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20240206_HepCl_141 | CC12 | Standard | 3.79 | 532644.188 | 17493900 | 142639.484 | 373.42 | 30.5 | 28.81 | -5.6
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20240206_HepCl_021 | CC13 | Standard | 3.79 | 655501.875 | 20798108 | 119743.57 | 547.421 | 48.75 | 42.47 | -12.9
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20240206_HepCl_037 | CC13 | Standard | 3.79 | 745636.563 | 24235634 | 135464.547 | 550.429 | 48.75 | 42.71 | -12.4
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20240206_HepCl_042 | CC13 | Standard | 3.79 | 699896.938 | 22405544 | 116004.852 | 603.334 | 48.75 | 46.87 | -3.9
| Beloq eLoQ | [15 empty] | 20240206_HepCl_047 | CC13 | Standard | 3.79 | 904464 | 29545996 | 140587.422 | 643.346 | 48.75 | 50.01 | 2.6
[17 empty] | 20240206_HepCl_064 | CC13 | Standard | 3.79 | 780961.375 | 25277208 | 130902.141 | 596.599 | 48.75 | 46.34 | -5
[17 empty] | 20240206_HepCl_080 | CC13 | Standard | 3.78 | 922489.438 | 29977964 | 138754.016 | 664.838 | 48.75 | 51.7 | 6
[17 empty] | 20240206_HepCl_085 | CC13 | Standard | 3.79 | 820857.625 | 26492756 | 142142.094 | 577.491 | 48.75 | 44.84 | -8
[17 empty] | 20240206_HepCl_090 | CC13 | Standard | 3.79 | 858363.625 | 27619504 | 125548.438 | 683.691 | 48.75 | 53.18 | 9.1
[17 empty] | 20240206_HepCl_022 | CC14 | Standard | 3.79 | 1384856.5 | 45921412 | 111885.367 | 1237.746 | 78.125 | 96.69 | 23.8
[17 empty] | 20240206_HepCl_065 | CC14 | Standard | 3.79 | 1138025.875 | 36036232 | 122410.234 | 929.682 | 78.125 | 72.5 | -7.2
[17 empty] | 20240206_HepCl_023 | CC15 | Standard | 3.79 | 1852453.625 | 59662784 | 89403.023 | 2072.026 | 125 | 162.22 | 29.8
[17 empty] | 20240206_HepCl_066 | CC15 | Standard | 3.79 | 1911083.125 | 61873040 | 116367.969 | 1642.276 | 125 | 128.47 | 2.8
[17 empty] | 20240206_HepCl_024 | CC16 | Standard | 3.79 | 2635860 | 83402032 | 133968.5 | 1967.522 | 200 | 154.01 | -23
[17 empty] | 20240206_HepCl_067 | CC16 | Standard | 3.79 | 2784664 | 89580352 | 132073.734 | 2108.416 | 200 | 165.08 | -17.5
[17 empty] | 20240206_HepCl_025 | CC17 | Standard | 3.79 | 3855763.5 | 121183752 | 90058.109 | 4281.417 | 325.625 | 335.75 | 3.1
[17 empty] | 20240206_HepCl_068 | CC17 | Standard | 3.79 | 4283628.5 | 135021648 | 98631.766 | 4343.052 | 325.625 | 340.6 | 4.6
[17 empty] | 20240206_HepCl_028 | CCQC A | QC | 3.79 | 21164.85 | 567620 | 144180.359 | 14.679 | 0.625 | 0.63 | 0.9
[17 empty] | 20240206_HepCl_071 | CCQC A | QC | 3.79 | 27920.525 | 739345 | 167490.781 | 16.67 | 0.625 | 0.79 | 25.9
[17 empty] | 20240206_HepCl_029 | CCQC B | QC | 3.79 | 77265.156 | 2400383 | 177202.609 | 43.603 | 2.506 | 2.9 | 15.8
[17 empty] | 20240206_HepCl_072 | CCQC B | QC | 3.79 | 63021.105 | 1860668 | 132315.688 | 47.629 | 2.506 | 3.22 | 28.4
[17 empty] | 20240206_HepCl_030 | CCQC C | QC | 3.79 | 150378.359 | 4879617 | 138244.625 | 108.777 | 6.269 | 8.02 | 28
[17 empty] | 20240206_HepCl_073 | CCQC C | QC | 3.79 | 162964.266 | 5223161 | 156897.484 | 103.867 | 6.269 | 7.64 | 21.8
[17 empty] | 20240206_HepCl_031 | CCQC D | QC | 3.79 | 558387.188 | 18370440 | 142977.109 | 390.543 | 25 | 30.15 | 20.6
[17 empty] | 20240206_HepCl_074 | CCQC D | QC | 3.79 | 623981.313 | 20126016 | 166386.391 | 375.019 | 25 | 28.93 | 15.7
[17 empty] | 20240206_HepCl_093 | Active T0 A | Analyte | 3.79 | 757865.313 | 23980454 | 483745.031 | 156.666 |  | 11.78
[17 empty] | 20240206_HepCl_094 | Active T0 B | Analyte | 3.79 | 685363.375 | 21845232 | 459075.188 | 149.292 |  | 11.2
[17 empty] | 20240206_HepCl_095 | Active T0 C | Analyte | 3.79 | 748576.5 | 24609110 | 528794.375 | 141.563 |  | 10.6
[17 empty] | 20240206_HepCl_096 | Active T15 A | Analyte | 3.79 | 1012456.563 | 33141090 | 551724.25 | 183.508 |  | 13.89
[17 empty] | 20240206_HepCl_097 | Active T15 B | Analyte | 3.79 | 968402.5 | 32242352 | 583719.563 | 165.902 |  | 12.51
[17 empty] | 20240206_HepCl_098 | Active T15 C | Analyte | 3.79 | 786365.438 | 25254146 | 514480.375 | 152.847 |  | 11.48
[17 empty] | 20240206_HepCl_100 | Active T30 A | Analyte | 3.79 | 726149.313 | 23132068 | 487161.469 | 149.057 |  | 11.18
[17 empty] | 20240206_HepCl_101 | Active T30 B | Analyte | 3.78 | 778169.25 | 25243274 | 572075.313 | 136.026 |  | 10.16
[17 empty] | 20240206_HepCl_102 | Active T30 C | Analyte | 3.79 | 675882.688 | 21901790 | 481638.969 | 140.33 |  | 10.5
[17 empty] | 20240206_HepCl_103 | Active T60 A | Analyte | 3.79 | 1018539.625 | 33397900 | 530218.813 | 192.098 |  | 14.57
[17 empty] | 20240206_HepCl_104 | Active T60 B | Analyte | 3.79 | 787809.75 | 25440130 | 598683.313 | 131.59 |  | 9.81
[17 empty] | 20240206_HepCl_105 | Active T60 C | Analyte | 3.79 | 989299.5 | 32508514 | 600279.688 | 164.806 |  | 12.42
[17 empty] | 20240206_HepCl_109 | Active T120 A | Analyte | 3.79 | 843010.313 | 27373468 | 748670.625 | 112.601 |  | 8.32
[17 empty] | 20240206_HepCl_110 | Active T120 B | Analyte | 3.79 | 697223.625 | 22892238 | 611171.313 | 114.08 |  | 8.44
[17 empty] | 20240206_HepCl_111 | Active T120 C | Analyte | 3.79 | 590744.375 | 18706572 | 630458.313 | 93.701 |  | 6.84
[17 empty] | 20240206_HepCl_112 | Active T240 A | Analyte | 3.79 | 589344.875 | 18359838 | 530024 | 111.192 |  | 8.21
[17 empty] | 20240206_HepCl_113 | Active T240 B | Analyte | 3.79 | 637305.75 | 20351406 | 553846.313 | 115.069 |  | 8.52
[17 empty] | 20240206_HepCl_114 | Active T240 C | Analyte | 3.79 | 594142.438 | 19279252 | 517922.813 | 114.716 |  | 8.49
[17 empty] | 20240206_HepCl_118 | Media T0 A | Analyte | 3.79 | 706007.5 | 22494514 | 483696.719 | 145.961 |  | 10.94
[17 empty] | 20240206_HepCl_119 | Media T0 B | Analyte | 3.78 | 715865.375 | 23240018 | 554689.25 | 129.057 |  | 9.61
[17 empty] | 20240206_HepCl_120 | Media T0 C | Analyte | 3.79 | 741524.938 | 24445986 | 605034.563 | 122.559 |  | 9.1
[17 empty] | 20240206_HepCl_121 | Media T15 A | Analyte | 3.79 | 804611.063 | 25870144 | 436108.25 | 184.498 |  | 13.97
[17 empty] | 20240206_HepCl_122 | Media T15 B | Analyte | 3.79 | 898459.813 | 29312212 | 497355.656 | 180.647 |  | 13.67
[17 empty] | 20240206_HepCl_123 | Media T15 C | Analyte | 3.79 | 632550.5 | 20060792 | 458635.219 | 137.92 |  | 10.31
[17 empty] | 20240206_HepCl_125 | Media T30 A | Analyte | 3.79 | 692136.75 | 22767466 | 549553.5 | 125.945 |  | 9.37
[17 empty] | 20240206_HepCl_126 | Media T30 B | Analyte | 3.79 | 528796.875 | 16542680 | 542602.625 | 97.456 |  | 7.13
[17 empty] | 20240206_HepCl_127 | Media T30 C | Analyte | 3.79 | 671984.125 | 22347616 | 572704.688 | 117.335 |  | 8.69
[17 empty] | 20240206_HepCl_128 | Media T60 A | Analyte | 3.79 | 831415.125 | 26780574 | 517056.781 | 160.798 |  | 12.11
[17 empty] | 20240206_HepCl_129 | Media T60 B | Analyte | 3.79 | 825454.438 | 26125740 | 545332.438 | 151.367 |  | 11.37
[17 empty] | 20240206_HepCl_130 | Media T60 C | Analyte | 3.79 | 899184.625 | 28970670 | 564425.188 | 159.31 |  | 11.99
[17 empty] | 20240206_HepCl_134 | Media T120 A | Analyte | 3.79 | 488209.906 | 15396140 | 646892.625 | 75.47 |  | 5.41
[17 empty] | 20240206_HepCl_135 | Media T120 B | Analyte | 3.79 | 454990.031 | 14362970 | 527762.875 | 86.211 |  | 6.25
[17 empty] | 20240206_HepCl_136 | Media T120 C | Analyte | 3.79 | 586462.188 | 19158654 | 616629.188 | 95.108 |  | 6.95
[17 empty] | 20240206_HepCl_137 | Media T240 A | Analyte | 3.79 | 529178.5 | 16877646 | 554051.313 | 95.511 |  | 6.98
[17 empty] | 20240206_HepCl_138 | Media T240 B | Analyte | 3.79 | 547773.438 | 17375820 | 507204.156 | 107.999 |  | 7.96
[17 empty] | 20240206_HepCl_139 | Media T240 C | Analyte | 3.79 | 736954 | 24175192 | 631669 | 116.668 |  | 8.64
[17 empty] | 20240206_HepCl_143 | Inactive T0 A | Analyte | 3.79 | 582762.813 | 19011156 | 525006.813 | 111.001 |  | 8.2
[17 empty] | 20240206_HepCl_144 | Inactive T0 B | Analyte | 3.79 | 544444.063 | 17852800 | 551036.875 | 98.804 |  | 7.24
[17 empty] | 20240206_HepCl_145 | Inactive T0 C | Analyte | 3.79 | 446171.031 | 14175868 | 475163.781 | 93.898 |  | 6.85
[17 empty] | 20240206_HepCl_146 | Inactive T240 A | Analyte | 3.79 | 550984.25 | 17694306 | 448450.656 | 122.864 |  | 9.13
[17 empty] | 20240206_HepCl_147 | Inactive T240 B | Analyte | 3.79 | 445698.344 | 14486281 | 490642.688 | 90.84 |  | 6.61
[17 empty] | 20240206_HepCl_148 | Inactive T240 C | Analyte | 3.79 | 624689.375 | 20648322 | 618428.688 | 101.012 |  | 7.41
[17 empty] | 20240206_HepCl_004 | Crash Blank | Blank | 3.84 | 4666.37 | 63665 | 7.834 | 59565.611 |  | 4677.96
[17 empty] | 20240206_HepCl_003 | Matrix Blank | Blank | 3.85 | 5091.157 | 66742 | 77.175 | 6596.899 |  | 517.62
[17 empty] | 20240206_HepCl_001 | Blank | Blank | 3.87 | 4602.562 | 58641 | 51.652 | 8910.714 |  | 699.36
[17 empty] | 20240206_HepCl_002 | Blank | Blank | 3.86 | 5433.033 | 70617 | 2.539 | 213983.182 |  | 16806.45
[17 empty] | 20240206_HepCl_005 | Blank | Blank | 3.84 | 5930.396 | 71987 | 12.389 | 47868.238 |  | 3759.21
[17 empty] | 20240206_HepCl_006 | Blank | Blank | 3.85 | 5161.199 | 67760 | 7.742 | 66664.932 |  | 5235.57
[17 empty] | 20240206_HepCl_013 | Blank | Blank | 3.85 | 4084.238 | 59924 | 27.52 | 14840.981 |  | 1165.14
[17 empty] | 20240206_HepCl_020 | Blank | Blank | 3.85 | 4801.735 | 63427
[17 empty] | 20240206_HepCl_026 | Blank | Blank | 3.84 | 5251.62 | 64577 | 14.134 | 37155.936 |  | 2917.83
[17 empty] | 20240206_HepCl_027 | Blank | Blank | 3.86 | 5246.035 | 71244 | 12.725 | 41226.208 |  | 3237.53
[17 empty] | 20240206_HepCl_032 | Blank | Blank | 3.84 | 5866.67 | 66857 | 66.006 | 8888.086 |  | 697.58
[17 empty] | 20240206_HepCl_033 | Blank | Blank | 3.84 | 4312.259 | 64444 | 0.973 | 443192.086 |  | 34809.31
[17 empty] | 20240206_HepCl_038 | Blank | Blank | 3.85 | 5647.38 | 69157 | 25.41 | 22225.03 |  | 1745.11
[17 empty] | 20240206_HepCl_043 | Blank | Blank | 3.86 | 5520.508 | 70644 | 12.584 | 43869.263 |  | 3445.12
[17 empty] | 20240206_HepCl_048 | Blank | Blank | 3.85 | 5790.475 | 66684 | 13.116 | 44148.178 |  | 3467.03
[17 empty] | 20240206_HepCl_049 | Blank | Blank | 3.85 | 4424.665 | 60633 | 8.044 | 55005.781 |  | 4319.82
[17 empty] | 20240206_HepCl_056 | Blank | Blank | 3.85 | 5508.639 | 79349 | 61.414 | 8969.68 |  | 703.99
[17 empty] | 20240206_HepCl_063 | Blank | Blank | 3.85 | 4456.979 | 67963
[17 empty] | 20240206_HepCl_069 | Blank | Blank | 3.85 | 4351.339 | 61376 | 70.645 | 6159.444 |  | 483.26
[17 empty] | 20240206_HepCl_070 | Blank | Blank | 3.85 | 5999.731 | 69681
[17 empty] | 20240206_HepCl_075 | Blank | Blank | 3.85 | 5932.333 | 74277 | 79.723 | 7441.181 |  | 583.93
[17 empty] | 20240206_HepCl_076 | Blank | Blank | 3.85 | 4732 | 69217 | 99.62 | 4750.05 |  | 372.56
[17 empty] | 20240206_HepCl_081 | Blank | Blank | 3.85 | 4536.292 | 61774 | 14.27 | 31789.012 |  | 2496.3
[17 empty] | 20240206_HepCl_086 | Blank | Blank | 3.85 | 4412.375 | 65235 | 10.153 | 43458.83 |  | 3412.88
[17 empty] | 20240206_HepCl_091 | Blank | Blank | 3.85 | 4792.176 | 72944 | 8.037 | 59626.428 |  | 4682.74
[17 empty] | 20240206_HepCl_092 | Blank | Blank | 3.85 | 5023.125 | 73481 | 40.381 | 12439.328 |  | 976.5
[17 empty] | 20240206_HepCl_099 | Blank | Blank | 3.85 | 4249.628 | 62655 | 5.725 | 74229.31 |  | 5829.7
[17 empty] | 20240206_HepCl_106 | Blank | Blank | 3.85 | 5541.751 | 75448
[17 empty] | 20240206_HepCl_108 | Blank | Blank | 3.85 | 4669.948 | 65573 | 12.753 | 36618.427 |  | 2875.61
[17 empty] | 20240206_HepCl_115 | Blank | Blank | 3.84 | 3708.198 | 57257 | 0.544 | 681654.044 |  | 53538.93
[17 empty] | 20240206_HepCl_117 | Blank | Blank | 3.85 | 4831.553 | 65807 | 1.556 | 310511.118 |  | 24388.09
[17 empty] | 20240206_HepCl_124 | Blank | Blank | 3.84 | 5092.387 | 64352 | 121.379 | 4195.443 |  | 329
[17 empty] | 20240206_HepCl_131 | Blank | Blank | 3.85 | 5546.071 | 72105 | 9.231 | 60080.934 |  | 4718.44
[17 empty] | 20240206_HepCl_133 | Blank | Blank | 3.85 | 4060.826 | 54872 | 101.073 | 4017.716 |  | 315.04
[17 empty] | 20240206_HepCl_140 | Blank | Blank | 3.85 | 4071.576 | 58685 | 16.476 | 24712.163 |  | 1940.46
[17 empty] | 20240206_HepCl_142 | Blank | Blank | 3.85 | 4597.399 | 64995 | 99.922 | 4600.988 |  | 360.85
[17 empty] | 20240206_HepCl_149 | Blank | Blank | 3.85 | 5672.912 | 66329 | 104.156 | 5446.553 |  | 427.27
[17 empty] | 20240206_HepCl_150 | Blank | Blank | 3.85 | 4435.503 | 61997 | 107.154 | 4139.372 |  | 324.6
[17 empty] | 20240206_HepCl_151 | Blank | Blank | 3.85 | 4759.074 | 64106 | 108.747 | 4376.281 |  | 343.21
Compound 4: M8PFOA
[17 empty] | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
[17 empty] | 20240206_HepCl_007 | CC01 | Standard | 3.79 | 134058.391 | 4381584 |  | 134058.391 | 100 | 103.53 | 3.5
[17 empty] | 20240206_HepCl_050 | CC01 | Standard | 3.79 | 146270.547 | 4814792 |  | 146270.547 | 100 | 112.96 | 13
[17 empty] | 20240206_HepCl_008 | CC02 | Standard | 3.79 | 118457.406 | 3813211 |  | 118457.406 | 100 | 91.48 | -8.5
[17 empty] | 20240206_HepCl_051 | CC02 | Standard | 3.79 | 153555.141 | 5056891 |  | 153555.141 | 100 | 118.59 | 18.6
[17 empty] | 20240206_HepCl_009 | CC03 | Standard | 3.79 | 146184.828 | 4807071 |  | 146184.828 | 100 | 112.9 | 12.9
[17 empty] | 20240206_HepCl_052 | CC03 | Standard | 3.78 | 133897.094 | 4409536 |  | 133897.094 | 100 | 103.41 | 3.4
[17 empty] | 20240206_HepCl_010 | CC04 | Standard | 3.79 | 134612.875 | 4458353 |  | 134612.875 | 100 | 103.96 | 4
[17 empty] | 20240206_HepCl_034 | CC04 | Standard | 3.78 | 136284.063 | 4536227 |  | 136284.063 | 100 | 105.25 | 5.3
[17 empty] | 20240206_HepCl_039 | CC04 | Standard | 3.79 | 149225.266 | 4943138 |  | 149225.266 | 100 | 115.24 | 15.2
[17 empty] | 20240206_HepCl_044 | CC04 | Standard | 3.78 | 146873.625 | 4790017 |  | 146873.625 | 100 | 113.43 | 13.4
[17 empty] | 20240206_HepCl_053 | CC04 | Standard | 3.78 | 112335.555 | 3608396 |  | 112335.555 | 100 | 86.76 | -13.2
[17 empty] | 20240206_HepCl_077 | CC04 | Standard | 3.78 | 150515.109 | 5014824 |  | 150515.109 | 100 | 116.24 | 16.2
[17 empty] | 20240206_HepCl_082 | CC04 | Standard | 3.78 | 129994.75 | 4272615 |  | 129994.75 | 100 | 100.39 | 0.4
[17 empty] | 20240206_HepCl_087 | CC04 | Standard | 3.79 | 157955.078 | 5231489 |  | 157955.078 | 100 | 121.99 | 22
[17 empty] | 20240206_HepCl_011 | CC05 | Standard | 3.79 | 134493.047 | 4452551 |  | 134493.047 | 100 | 103.87 | 3.9
[17 empty] | 20240206_HepCl_035 | CC05 | Standard | 3.78 | 135748.172 | 4522542 |  | 135748.172 | 100 | 104.84 | 4.8
[17 empty] | 20240206_HepCl_040 | CC05 | Standard | 3.79 | 136307.844 | 4519393 |  | 136307.844 | 100 | 105.27 | 5.3
[17 empty] | 20240206_HepCl_045 | CC05 | Standard | 3.78 | 143390.891 | 4802327 |  | 143390.891 | 100 | 110.74 | 10.7
[17 empty] | 20240206_HepCl_054 | CC05 | Standard | 3.79 | 117635.977 | 3856072 |  | 117635.977 | 100 | 90.85 | -9.2
[17 empty] | 20240206_HepCl_078 | CC05 | Standard | 3.78 | 137131.031 | 4567322 |  | 137131.031 | 100 | 105.9 | 5.9
[17 empty] | 20240206_HepCl_083 | CC05 | Standard | 3.78 | 128168.617 | 4134631 |  | 128168.617 | 100 | 98.98 | -1
[17 empty] | 20240206_HepCl_088 | CC05 | Standard | 3.79 | 141630.219 | 4617237 |  | 141630.219 | 100 | 109.38 | 9.4
[17 empty] | 20240206_HepCl_012 | CC06 | Standard | 3.79 | 114462.625 | 3698309 |  | 114462.625 | 100 | 88.4 | -11.6
[17 empty] | 20240206_HepCl_036 | CC06 | Standard | 3.78 | 130111.891 | 4285330 |  | 130111.891 | 100 | 100.48 | 0.5
[17 empty] | 20240206_HepCl_041 | CC06 | Standard | 3.79 | 122278.68 | 3964777 |  | 122278.68 | 100 | 94.43 | -5.6
[17 empty] | 20240206_HepCl_046 | CC06 | Standard | 3.78 | 135189.078 | 4472953 |  | 135189.078 | 100 | 104.4 | 4.4
[17 empty] | 20240206_HepCl_055 | CC06 | Standard | 3.79 | 142075.469 | 4662613 |  | 142075.469 | 100 | 109.72 | 9.7
[17 empty] | 20240206_HepCl_079 | CC06 | Standard | 3.78 | 129202.703 | 4249507 |  | 129202.703 | 100 | 99.78 | -0.2
[17 empty] | 20240206_HepCl_084 | CC06 | Standard | 3.78 | 135884.641 | 4425236 |  | 135884.641 | 100 | 104.94 | 4.9
[17 empty] | 20240206_HepCl_089 | CC06 | Standard | 3.79 | 131731.141 | 4373728 |  | 131731.141 | 100 | 101.73 | 1.7
[17 empty] | 20240206_HepCl_014 | CC07 | Standard | 3.79 | 124790.297 | 4079174 |  | 124790.297 | 100 | 96.37 | -3.6
[17 empty] | 20240206_HepCl_057 | CC07 | Standard | 3.78 | 117855.555 | 3814840 |  | 117855.555 | 100 | 91.02 | -9
[17 empty] | 20240206_HepCl_015 | CC08 | Standard | 3.79 | 129505.023 | 4280320 |  | 129505.023 | 100 | 100.02 | 0
[17 empty] | 20240206_HepCl_058 | CC08 | Standard | 3.78 | 113011.789 | 3641974 |  | 113011.789 | 100 | 87.28 | -12.7
[17 empty] | 20240206_HepCl_016 | CC09 | Standard | 3.79 | 117331.039 | 3835934 |  | 117331.039 | 100 | 90.61 | -9.4
[17 empty] | 20240206_HepCl_059 | CC09 | Standard | 3.79 | 118829.922 | 3856934 |  | 118829.922 | 100 | 91.77 | -8.2
[17 empty] | 20240206_HepCl_017 | CC10 | Standard | 3.79 | 125046.508 | 4131017 |  | 125046.508 | 100 | 96.57 | -3.4
[17 empty] | 20240206_HepCl_060 | CC10 | Standard | 3.78 | 152427.219 | 5080944 |  | 152427.219 | 100 | 117.72 | 17.7
[17 empty] | 20240206_HepCl_018 | CC11 | Standard | 3.79 | 119904.609 | 3949575 |  | 119904.609 | 100 | 92.6 | -7.4
[17 empty] | 20240206_HepCl_061 | CC11 | Standard | 3.79 | 134633.641 | 4370997 |  | 134633.641 | 100 | 103.98 | 4
[17 empty] | 20240206_HepCl_019 | CC12 | Standard | 3.79 | 114519.438 | 3692099 |  | 114519.438 | 100 | 88.44 | -11.6
[17 empty] | 20240206_HepCl_062 | CC12 | Standard | 3.79 | 130157.563 | 4261513 |  | 130157.563 | 100 | 100.52 | 0.5
[17 empty] | 20240206_HepCl_107 | CC12 | Standard | 3.78 | 116307.844 | 3760237 |  | 116307.844 | 100 | 89.82 | -10.2
[17 empty] | 20240206_HepCl_116 | CC12 | Standard | 3.78 | 136432.594 | 4517728 |  | 136432.594 | 100 | 105.37 | 5.4
[17 empty] | 20240206_HepCl_132 | CC12 | Standard | 3.78 | 125093.547 | 4092294 |  | 125093.547 | 100 | 96.61 | -3.4
[17 empty] | 20240206_HepCl_141 | CC12 | Standard | 3.79 | 142639.484 | 4738849 |  | 142639.484 | 100 | 110.16 | 10.2
[17 empty] | 20240206_HepCl_021 | CC13 | Standard | 3.79 | 119743.57 | 3841606 |  | 119743.57 | 100 | 92.48 | -7.5
[17 empty] | 20240206_HepCl_037 | CC13 | Standard | 3.79 | 135464.547 | 4472643 |  | 135464.547 | 100 | 104.62 | 4.6
[17 empty] | 20240206_HepCl_042 | CC13 | Standard | 3.78 | 116004.852 | 3779762 |  | 116004.852 | 100 | 89.59 | -10.4
[17 empty] | 20240206_HepCl_047 | CC13 | Standard | 3.79 | 140587.422 | 4566574 |  | 140587.422 | 100 | 108.57 | 8.6
[17 empty] | 20240206_HepCl_064 | CC13 | Standard | 3.79 | 130902.141 | 4297562 |  | 130902.141 | 100 | 101.09 | 1.1
[17 empty] | 20240206_HepCl_080 | CC13 | Standard | 3.78 | 138754.016 | 4544742 |  | 138754.016 | 100 | 107.16 | 7.2
[17 empty] | 20240206_HepCl_085 | CC13 | Standard | 3.78 | 142142.094 | 4720642 |  | 142142.094 | 100 | 109.77 | 9.8
[17 empty] | 20240206_HepCl_090 | CC13 | Standard | 3.78 | 125548.438 | 4030725 |  | 125548.438 | 100 | 96.96 | -3
[17 empty] | 20240206_HepCl_022 | CC14 | Standard | 3.79 | 111885.367 | 3616760 |  | 111885.367 | 100 | 86.41 | -13.6
[17 empty] | 20240206_HepCl_065 | CC14 | Standard | 3.79 | 122410.234 | 3944801 |  | 122410.234 | 100 | 94.54 | -5.5
[17 empty] | 20240206_HepCl_023 | CC15 | Standard | 3.79 | 89403.023 | 2825222 |  | 89403.023 | 100 | 69.04 | -31
[17 empty] | 20240206_HepCl_066 | CC15 | Standard | 3.79 | 116367.969 | 3750585 |  | 116367.969 | 100 | 89.87 | -10.1
[17 empty] | 20240206_HepCl_024 | CC16 | Standard | 3.78 | 133968.5 | 4275505 |  | 133968.5 | 100 | 103.46 | 3.5
[17 empty] | 20240206_HepCl_067 | CC16 | Standard | 3.79 | 132073.734 | 4259799 |  | 132073.734 | 100 | 102 | 2
[17 empty] | 20240206_HepCl_025 | CC17 | Standard | 3.79 | 90058.109 | 2840044 |  | 90058.109 | 100 | 69.55 | -30.4
[17 empty] | 20240206_HepCl_068 | CC17 | Standard | 3.78 | 98631.766 | 3073554 |  | 98631.766 | 100 | 76.17 | -23.8
[17 empty] | 20240206_HepCl_028 | CCQC A | QC | 3.79 | 144180.359 | 4730226 |  | 144180.359 | 100 | 111.35 | 11.3
[17 empty] | 20240206_HepCl_071 | CCQC A | QC | 3.79 | 167490.781 | 5517542 |  | 167490.781 | 100 | 129.35 | 29.4
[17 empty] | 20240206_HepCl_029 | CCQC B | QC | 3.79 | 177202.609 | 5804371 |  | 177202.609 | 100 | 136.85 | 36.9
[17 empty] | 20240206_HepCl_072 | CCQC B | QC | 3.79 | 132315.688 | 4281620 |  | 132315.688 | 100 | 102.19 | 2.2
[17 empty] | 20240206_HepCl_030 | CCQC C | QC | 3.78 | 138244.625 | 4529511 |  | 138244.625 | 100 | 106.76 | 6.8
[17 empty] | 20240206_HepCl_073 | CCQC C | QC | 3.79 | 156897.484 | 5229336 |  | 156897.484 | 100 | 121.17 | 21.2
[17 empty] | 20240206_HepCl_031 | CCQC D | QC | 3.79 | 142977.109 | 4675255 |  | 142977.109 | 100 | 110.42 | 10.4
[17 empty] | 20240206_HepCl_074 | CCQC D | QC | 3.78 | 166386.391 | 5484595 |  | 166386.391 | 100 | 128.5 | 28.5
[17 empty] | 20240206_HepCl_093 | Active T0 A | Analyte | 3.79 | 483745.031 | 15100567 |  | 483745.031 | 100 | 373.59 | 273.6
[17 empty] | 20240206_HepCl_094 | Active T0 B | Analyte | 3.78 | 459075.188 | 14633644 |  | 459075.188 | 100 | 354.54 | 254.5
[17 empty] | 20240206_HepCl_095 | Active T0 C | Analyte | 3.79 | 528794.375 | 17161112 |  | 528794.375 | 100 | 408.38 | 308.4
[17 empty] | 20240206_HepCl_096 | Active T15 A | Analyte | 3.79 | 551724.25 | 17942246 |  | 551724.25 | 100 | 426.09 | 326.1
[17 empty] | 20240206_HepCl_097 | Active T15 B | Analyte | 3.79 | 583719.563 | 19246358 |  | 583719.563 | 100 | 450.8 | 350.8
[17 empty] | 20240206_HepCl_098 | Active T15 C | Analyte | 3.78 | 514480.375 | 16539207 |  | 514480.375 | 100 | 397.33 | 297.3
[17 empty] | 20240206_HepCl_100 | Active T30 A | Analyte | 3.78 | 487161.469 | 16129338 |  | 487161.469 | 100 | 376.23 | 276.2
[17 empty] | 20240206_HepCl_101 | Active T30 B | Analyte | 3.78 | 572075.313 | 19031012 |  | 572075.313 | 100 | 441.81 | 341.8
[17 empty] | 20240206_HepCl_102 | Active T30 C | Analyte | 3.79 | 481638.969 | 15452899 |  | 481638.969 | 100 | 371.96 | 272
[17 empty] | 20240206_HepCl_103 | Active T60 A | Analyte | 3.79 | 530218.813 | 17092980 |  | 530218.813 | 100 | 409.48 | 309.5
[17 empty] | 20240206_HepCl_104 | Active T60 B | Analyte | 3.78 | 598683.313 | 19767818 |  | 598683.313 | 100 | 462.36 | 362.4
[17 empty] | 20240206_HepCl_105 | Active T60 C | Analyte | 3.79 | 600279.688 | 19534292 |  | 600279.688 | 100 | 463.59 | 363.6
[17 empty] | 20240206_HepCl_109 | Active T120 A | Analyte | 3.78 | 748670.625 | 24375130 |  | 748670.625 | 100 | 578.19 | 478.2
[17 empty] | 20240206_HepCl_110 | Active T120 B | Analyte | 3.78 | 611171.313 | 19885652 |  | 611171.313 | 100 | 472 | 372
[17 empty] | 20240206_HepCl_111 | Active T120 C | Analyte | 3.78 | 630458.313 | 20525438 |  | 630458.313 | 100 | 486.9 | 386.9
[17 empty] | 20240206_HepCl_112 | Active T240 A | Analyte | 3.78 | 530024 | 16970868 |  | 530024 | 100 | 409.33 | 309.3
[17 empty] | 20240206_HepCl_113 | Active T240 B | Analyte | 3.78 | 553846.313 | 17926820 |  | 553846.313 | 100 | 427.73 | 327.7
[17 empty] | 20240206_HepCl_114 | Active T240 C | Analyte | 3.78 | 517922.813 | 16951756 |  | 517922.813 | 100 | 399.99 | 300
[17 empty] | 20240206_HepCl_118 | Media T0 A | Analyte | 3.78 | 483696.719 | 15596621 |  | 483696.719 | 100 | 373.55 | 273.6
[17 empty] | 20240206_HepCl_119 | Media T0 B | Analyte | 3.78 | 554689.25 | 18180086 |  | 554689.25 | 100 | 428.38 | 328.4
[17 empty] | 20240206_HepCl_120 | Media T0 C | Analyte | 3.78 | 605034.563 | 20058984 |  | 605034.563 | 100 | 467.26 | 367.3
[17 empty] | 20240206_HepCl_121 | Media T15 A | Analyte | 3.79 | 436108.25 | 14171177 |  | 436108.25 | 100 | 336.8 | 236.8
[17 empty] | 20240206_HepCl_122 | Media T15 B | Analyte | 3.78 | 497355.656 | 16241408 |  | 497355.656 | 100 | 384.1 | 284.1
[17 empty] | 20240206_HepCl_123 | Media T15 C | Analyte | 3.79 | 458635.219 | 14618891 |  | 458635.219 | 100 | 354.2 | 254.2
[17 empty] | 20240206_HepCl_125 | Media T30 A | Analyte | 3.79 | 549553.5 | 18123876 |  | 549553.5 | 100 | 424.41 | 324.4
[17 empty] | 20240206_HepCl_126 | Media T30 B | Analyte | 3.79 | 542602.625 | 17445314 |  | 542602.625 | 100 | 419.05 | 319
[17 empty] | 20240206_HepCl_127 | Media T30 C | Analyte | 3.79 | 572704.688 | 18743662 |  | 572704.688 | 100 | 442.29 | 342.3
[17 empty] | 20240206_HepCl_128 | Media T60 A | Analyte | 3.78 | 517056.781 | 16824996 |  | 517056.781 | 100 | 399.32 | 299.3
[17 empty] | 20240206_HepCl_129 | Media T60 B | Analyte | 3.78 | 545332.438 | 17639016 |  | 545332.438 | 100 | 421.15 | 321.2
[17 empty] | 20240206_HepCl_130 | Media T60 C | Analyte | 3.79 | 564425.188 | 17941832 |  | 564425.188 | 100 | 435.9 | 335.9
[17 empty] | 20240206_HepCl_134 | Media T120 A | Analyte | 3.79 | 646892.625 | 21053214 |  | 646892.625 | 100 | 499.59 | 399.6
[17 empty] | 20240206_HepCl_135 | Media T120 B | Analyte | 3.78 | 527762.875 | 16439441 |  | 527762.875 | 100 | 407.58 | 307.6
[17 empty] | 20240206_HepCl_136 | Media T120 C | Analyte | 3.79 | 616629.188 | 19846670 |  | 616629.188 | 100 | 476.22 | 376.2
[17 empty] | 20240206_HepCl_137 | Media T240 A | Analyte | 3.79 | 554051.313 | 17937966 |  | 554051.313 | 100 | 427.89 | 327.9
[17 empty] | 20240206_HepCl_138 | Media T240 B | Analyte | 3.78 | 507204.156 | 16314303 |  | 507204.156 | 100 | 391.71 | 291.7
[17 empty] | 20240206_HepCl_139 | Media T240 C | Analyte | 3.78 | 631669 | 20682628 |  | 631669 | 100 | 487.83 | 387.8
[17 empty] | 20240206_HepCl_143 | Inactive T0 A | Analyte | 3.78 | 525006.813 | 16898870 |  | 525006.813 | 100 | 405.46 | 305.5
[17 empty] | 20240206_HepCl_144 | Inactive T0 B | Analyte | 3.79 | 551036.875 | 18172490 |  | 551036.875 | 100 | 425.56 | 325.6
[17 empty] | 20240206_HepCl_145 | Inactive T0 C | Analyte | 3.79 | 475163.781 | 15171547 |  | 475163.781 | 100 | 366.96 | 267
[17 empty] | 20240206_HepCl_146 | Inactive T240 A | Analyte | 3.79 | 448450.656 | 14187672 |  | 448450.656 | 100 | 346.33 | 246.3
[17 empty] | 20240206_HepCl_147 | Inactive T240 B | Analyte | 3.79 | 490642.688 | 15946417 |  | 490642.688 | 100 | 378.92 | 278.9
[17 empty] | 20240206_HepCl_148 | Inactive T240 C | Analyte | 3.79 | 618428.688 | 20617388 |  | 618428.688 | 100 | 477.61 | 377.6
[17 empty] | 20240206_HepCl_004 | Crash Blank | Blank | 3.9 | 7.834 | 292 |  | 7.834 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_003 | Matrix Blank | Blank | 3.88 | 77.175 | 526 |  | 77.175 | 100 | 0.06 | -99.9
[17 empty] | 20240206_HepCl_001 | Blank | Blank | 3.77 | 51.652 | 463 |  | 51.652 | 100 | 0.04 | -100
[17 empty] | 20240206_HepCl_002 | Blank | Blank | 3.86 | 2.539 | 144 |  | 2.539 | 100 | 0 | -100
[17 empty] | 20240206_HepCl_005 | Blank | Blank | 3.86 | 12.389 | 310 |  | 12.389 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_006 | Blank | Blank | 3.8 | 7.742 | 287 |  | 7.742 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_013 | Blank | Blank | 3.64 | 27.52 | 389 |  | 27.52 | 100 | 0.02 | -100
[17 empty] | 20240206_HepCl_020 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepCl_026 | Blank | Blank | 3.77 | 14.134 | 366 |  | 14.134 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_027 | Blank | Blank | 3.68 | 12.725 | 262 |  | 12.725 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_032 | Blank | Blank | 3.8 | 66.006 | 444 |  | 66.006 | 100 | 0.05 | -99.9
[17 empty] | 20240206_HepCl_033 | Blank | Blank | 3.69 | 0.973 | 51 |  | 0.973 | 100 | 0 | -100
[17 empty] | 20240206_HepCl_038 | Blank | Blank | 3.84 | 25.41 | 427 |  | 25.41 | 100 | 0.02 | -100
[17 empty] | 20240206_HepCl_043 | Blank | Blank | 3.96 | 12.584 | 294 |  | 12.584 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_048 | Blank | Blank | 3.83 | 13.116 | 293 |  | 13.116 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_049 | Blank | Blank | 3.72 | 8.044 | 271 |  | 8.044 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_056 | Blank | Blank | 3.86 | 61.414 | 389 |  | 61.414 | 100 | 0.05 | -100
[17 empty] | 20240206_HepCl_063 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepCl_069 | Blank | Blank | 3.89 | 70.645 | 511 |  | 70.645 | 100 | 0.05 | -99.9
[17 empty] | 20240206_HepCl_070 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepCl_075 | Blank | Blank | 3.79 | 79.723 | 480 |  | 79.723 | 100 | 0.06 | -99.9
[17 empty] | 20240206_HepCl_076 | Blank | Blank | 3.83 | 99.62 | 430 |  | 99.62 | 100 | 0.08 | -99.9
[17 empty] | 20240206_HepCl_081 | Blank | Blank | 3.8 | 14.27 | 304 |  | 14.27 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_086 | Blank | Blank | 3.75 | 10.153 | 372 |  | 10.153 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_091 | Blank | Blank | 3.78 | 8.037 | 275 |  | 8.037 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_092 | Blank | Blank | 3.83 | 40.381 | 290 |  | 40.381 | 100 | 0.03 | -100
[17 empty] | 20240206_HepCl_099 | Blank | Blank | 3.66 | 5.725 | 214 |  | 5.725 | 100 | 0 | -100
[17 empty] | 20240206_HepCl_106 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepCl_108 | Blank | Blank | 3.84 | 12.753 | 406 |  | 12.753 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_115 | Blank | Blank | 3.86 | 0.544 | 62 |  | 0.544 | 100 | 0 | -100
[17 empty] | 20240206_HepCl_117 | Blank | Blank | 3.78 | 1.556 | 170 |  | 1.556 | 100 | 0 | -100
[17 empty] | 20240206_HepCl_124 | Blank | Blank | 3.83 | 121.379 | 853 |  | 121.379 | 100 | 0.09 | -99.9
[17 empty] | 20240206_HepCl_131 | Blank | Blank | 3.78 | 9.231 | 356 |  | 9.231 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_133 | Blank | Blank | 3.86 | 101.073 | 831 |  | 101.073 | 100 | 0.08 | -99.9
[17 empty] | 20240206_HepCl_140 | Blank | Blank | 3.86 | 16.476 | 597 |  | 16.476 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_142 | Blank | Blank | 3.81 | 99.922 | 601 |  | 99.922 | 100 | 0.08 | -99.9
[17 empty] | 20240206_HepCl_149 | Blank | Blank | 3.84 | 104.156 | 560 |  | 104.156 | 100 | 0.08 | -99.9
[17 empty] | 20240206_HepCl_150 | Blank | Blank | 3.78 | 107.154 | 666 |  | 107.154 | 100 | 0.08 | -99.9
[17 empty] | 20240206_HepCl_151 | Blank | Blank | 3.87 | 108.747 | 553 |  | 108.747 | 100 | 0.08 | -99.9

## PFOSA
Human Active Cells | [8 empty] | Media Only | [7 empty] | Compound 5: PFOSA
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20240206_HepCl_007 | CC01 | Standard | 5.16 | 1261.853 | 31157 | 28015.162 | 4.504 | 0.175 | 0.3 | 72.1
0 | Active T0 A | 53.68 | 6.8734 | 966.24 | 973.2 | 64.3629 | 6.614% |  | 0 | Media T0 A | 52.65 | 6.854 | 947.7 | 949.26 | 19.4869 | 2.053% | 20240206_HepCl_050 | CC01 | Standard | 5.17 | 1132.14 | 28317 | 24507.627 | 4.62 | 0.175 | 0.31 | 78.6
| Active T0 B | 57.82 | 6.9477 | 1040.76 | [5 empty] | Media T0 B | 51.7 | 6.8358 | 930.6 |  |  |  | 20240206_HepCl_008 | CC02 | Standard | 5.16 | 1216.324 | 29260 | 29313.949 | 4.149 | 0.281 | 0.27 | -5.3
| Active T0 C | 50.7 | 6.8163 | 912.6 | [5 empty] | Media T0 C | 53.86 | 6.8768 | 969.48 |  |  |  | 20240206_HepCl_051 | CC02 | Standard | 5.17 | 989.433 | 24612 | 23632.779 | 4.187 | 0.281 | 0.27 | -4
15 | Active T15 A | 46.55 | 6.7309 | 837.9 | 865.62 | 40.7596 | 4.709% |  | 15 | Media T15 A | 65.48 | 7.0721 | 1178.64 | 987.54 | 226.7953 | 22.97% | 20240206_HepCl_009 | CC03 | Standard | 5.17 | 1417.638 | 35882 | 29950.836 | 4.733 | 0.444 | 0.32 | -27.1
| Active T15 B | 47.03 | 6.7412 | 846.54 | [5 empty] | Media T15 B | 58.17 | 6.9537 | 1047.06 |  |  |  | 20240206_HepCl_052 | CC03 | Standard | 5.17 | 1100.143 | 26431 | 23528.35 | 4.676 | 0.444 | 0.32 | -28.3
| Active T15 C | 50.69 | 6.8161 | 912.42 | [5 empty] | Media T15 C | 40.94 | 6.6025 | 736.92 |  |  |  | 20240206_HepCl_010 | CC04 | Standard | 5.17 | 2318.271 | 56932 | 28339.805 | 8.18 | 0.713 | 0.66 | -7.1
30 | Active T30 A | 48.4 | 6.7699 | 871.2 | 766.14 | 96.6116 | 12.61% |  | 30 | Media T30 A | 39.11 | 6.5567 | 703.98 | 625.98 | 70.3416 | 11.24% | 20240206_HepCl_034 | CC04 | Standard | 5.16 | 2133.638 | 50852 | 24535.77 | 8.696 | 0.713 | 0.71 | 0
| Active T30 B | 41.45 | 6.6149 | 746.1 | [5 empty] | Media T30 B | 33.7 | 6.4079 | 606.6 |  |  |  | 20240206_HepCl_039 | CC04 | Standard | 5.16 | 2245.164 | 54945 | 26556.42 | 8.454 | 0.713 | 0.69 | -3.4
| Active T30 C | 37.84 | 6.5237 | 681.12 | [5 empty] | Media T30 C | 31.52 | 6.341 | 567.36 |  |  |  | 20240206_HepCl_044 | CC04 | Standard | 5.16 | 2043.82 | 51768 | 25185.447 | 8.115 | 0.713 | 0.66 | -8
60 | Active T60 A | 36.89 | 6.4983 | 664.02 | 691.08 | 23.4638 | 3.395% |  | 60 | Media T60 A | 53.56 | 6.8712 | 964.08 | 945.36 | 52.6007 | 5.564% | 20240206_HepCl_053 | CC04 | Standard | 5.17 | 1953.565 | 45869 | 23608.605 | 8.275 | 0.713 | 0.67 | -5.8
| Active T60 B | 39.08 | 6.556 | 703.44 | [5 empty] | Media T60 B | 54.78 | 6.8937 | 986.04 |  |  |  | 20240206_HepCl_077 | CC04 | Standard | 5.18 | 1691.695 | 39901 | 20045.84 | 8.439 | 0.713 | 0.69 | -3.6
| Active T60 C | 39.21 | 6.5593 | 705.78 | [5 empty] | Media T60 C | 49.22 | 6.7867 | 885.96 |  |  |  | 20240206_HepCl_082 | CC04 | Standard | 5.19 | 1678.866 | 41565 | 20375.537 | 8.24 | 0.713 | 0.67 | -6.3
120 | Active T120 A | 30.42 | 6.3055 | 547.56 | 538.02 | 8.2624 | 1.536% |  | 120 | Media T120 A | 45.34 | 6.7046 | 816.12 | 790.26 | 36.5539 | 4.626% | 20240206_HepCl_087 | CC04 | Standard | 5.19 | 1601.046 | 38244 | 19531.541 | 8.197 | 0.713 | 0.66 | -6.9
| Active T120 B | 29.62 | 6.2788 | 533.16 | [5 empty] | Media T120 B | 44.79 | 6.6924 | 806.22 |  |  |  | 20240206_HepCl_011 | CC05 | Standard | 5.17 | 3038.386 | 74959 | 25832.488 | 11.762 | 1.138 | 1.01 | -11
| Active T120 C | 29.63 | 6.2792 | 533.34 | [5 empty] | Media T120 C | 41.58 | 6.618 | 748.44 |  |  |  | 20240206_HepCl_035 | CC05 | Standard | 5.16 | 2891.436 | 68256 | 23397 | 12.358 | 1.138 | 1.07 | -5.8
240 | Active T240 A | 17.58 | 5.7571 | 316.44 | 320.04 | 31.4748 | 9.835% |  | 240 | Media T240 A | 47.41 | 6.7492 | 853.38 | 843.78 | 30.322 | 3.594% | 20240206_HepCl_040 | CC05 | Standard | 5.16 | 2739.347 | 66732 | 23445.711 | 11.684 | 1.138 | 1.01 | -11.6
| Active T240 B | 16.14 | 5.6717 | 290.52 | [5 empty] | Media T240 B | 48.23 | 6.7664 | 868.14 |  |  |  | 20240206_HepCl_045 | CC05 | Standard | 5.16 | 2716.965 | 65917 | 22468.701 | 12.092 | 1.138 | 1.05 | -8.1
| Active T240 C | 19.62 | 5.8669 | 353.16 | [5 empty] | Media T240 C | 44.99 | 6.6968 | 809.82 |  |  |  | 20240206_HepCl_054 | CC05 | Standard | 5.17 | 2599.149 | 62042 | 22096.361 | 11.763 | 1.138 | 1.01 | -11
[17 empty] | 20240206_HepCl_078 | CC05 | Standard | 5.18 | 2125.962 | 50142 | 17964.746 | 11.834 | 1.138 | 1.02 | -10.3
Human Inactive Cells | [16 empty] | 20240206_HepCl_083 | CC05 | Standard | 5.19 | 1908.929 | 45048 | 16582.211 | 11.512 | 1.138 | 0.99 | -13.1
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20240206_HepCl_088 | CC05 | Standard | 5.2 | 2017.531 | 48224 | 17208.271 | 11.724 | 1.138 | 1.01 | -11.3
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20240206_HepCl_012 | CC06 | Standard | 5.17 | 5194.874 | 130056 | 26238.695 | 19.799 | 1.825 | 1.8 | -1.3
0 | Inactive T0 A | 43.8 | 6.67 | 788.4 | 720.06 | 60.6844 | 8.428% | [9 empty] | 20240206_HepCl_036 | CC06 | Standard | 5.16 | 4643.119 | 112331 | 23039.303 | 20.153 | 1.825 | 1.84 | 0.6
| Inactive T0 B | 38.85 | 6.5501 | 699.3 | [12 empty] | 20240206_HepCl_041 | CC06 | Standard | 5.16 | 4788.455 | 119334 | 24264.488 | 19.734 | 1.825 | 1.79 | -1.7
| Inactive T0 C | 37.36 | 6.511 | 672.48 | [12 empty] | 20240206_HepCl_046 | CC06 | Standard | 5.16 | 4706.088 | 116252 | 22817.016 | 20.625 | 1.825 | 1.88 | 3.1
240 | Inactive T240 A | 37.43 | 6.5128 | 673.74 | 644.76 | 40.9839 | 6.356% | [9 empty] | 20240206_HepCl_055 | CC06 | Standard | 5.17 | 4287.854 | 104165 | 21123.363 | 20.299 | 1.825 | 1.85 | 1.4
| Inactive T240 B | 34.21 | 6.4229 | 615.78 | [12 empty] | 20240206_HepCl_079 | CC06 | Standard | 5.19 | 3641.675 | 90102 | 18868.227 | 19.301 | 1.825 | 1.75 | -4
| Inactive T240 C | 32.02 | 6.3567 | 576.36 | [12 empty] | 20240206_HepCl_084 | CC06 | Standard | 5.19 | 3367.372 | 79964 | 16807.939 | 20.034 | 1.825 | 1.82 | -0.1
[17 empty] | 20240206_HepCl_089 | CC06 | Standard | 5.2 | 3380.709 | 82340 | 16733.023 | 20.204 | 1.825 | 1.84 | 0.8
| Calibration Curve | [15 empty] | 20240206_HepCl_014 | CC07 | Standard | 5.17 | 7565.646 | 184579 | 27822.641 | 27.192 | 2.919 | 2.53 | -13.5
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20240206_HepCl_057 | CC07 | Standard | 5.17 | 5888.268 | 141624 | 22068.023 | 26.682 | 2.919 | 2.48 | -15.2
| 20240206_HepCl_012 | CC06 | 1.825 | 1.8 | 1.8213 | 0.04051 | 0.022243 | 0.12732 | -0.0020548 | [7 empty] | 20240206_HepCl_015 | CC08 | Standard | 5.17 | 12126.696 | 291778 | 27143.525 | 44.676 | 4.669 | 4.24 | -9.2
| 20240206_HepCl_036 | CC06 | 1.825 | 1.84 | [12 empty] | 20240206_HepCl_058 | CC08 | Standard | 5.17 | 9765.783 | 237113 | 22590.502 | 43.23 | 4.669 | 4.1 | -12.2
| 20240206_HepCl_041 | CC06 | 1.825 | 1.79 | [12 empty] | 20240206_HepCl_016 | CC09 | Standard | 5.17 | 13285.316 | 318537 | 17637.496 | 75.324 | 7.45 | 7.24 | -2.8
| 20240206_HepCl_046 | CC06 | 1.825 | 1.88 | [12 empty] | 20240206_HepCl_059 | CC09 | Standard | 5.17 | 10607.894 | 255942 | 14474.795 | 73.285 | 7.45 | 7.04 | -5.4
| 20240206_HepCl_055 | CC06 | 1.825 | 1.85 | [12 empty] | 20240206_HepCl_017 | CC10 | Standard | 5.17 | 32425.984 | 802510 | 26663.059 | 121.614 | 11.875 | 11.78 | -0.8
| 20240206_HepCl_079 | CC06 | 1.825 | 1.75 | [12 empty] | 20240206_HepCl_060 | CC10 | Standard | 5.17 | 27365.682 | 683191 | 21487.777 | 127.355 | 11.875 | 12.35 | 4
| 20240206_HepCl_084 | CC06 | 1.825 | 1.82 | [12 empty] | 20240206_HepCl_018 | CC11 | Standard | 5.17 | 61522.891 | 1515608 | 27305.922 | 225.31 | 19.125 | 21.95 | 14.8
| 20240206_HepCl_089 | CC06 | 1.825 | 1.84 | [12 empty] | 20240206_HepCl_061 | CC11 | Standard | 5.17 | 52900.586 | 1319640 | 22796.053 | 232.06 | 19.125 | 22.61 | 18.2
[17 empty] | 20240206_HepCl_019 | CC12 | Standard | 5.17 | 87626.531 | 2146827 | 27131.459 | 322.97 | 30.5 | 31.52 | 3.4
| Blanks should be less than… | 0.063662 |  | nM | [12 empty] | 20240206_HepCl_062 | CC12 | Standard | 5.18 | 67234.352 | 1651028 | 20643.158 | 325.698 | 30.5 | 31.79 | 4.2
| eLOQ | 1.825 |  | nM | [12 empty] | 20240206_HepCl_107 | CC12 | Standard | 5.2 | 45782.332 | 1091899 | 14109.446 | 324.48 | 30.5 | 31.67 | 3.8
[17 empty] | 20240206_HepCl_116 | CC12 | Standard | 5.2 | 40330.492 | 982018 | 12328.883 | 327.122 | 30.5 | 31.93 | 4.7
| QC FLAGS | [15 empty] | 20240206_HepCl_132 | CC12 | Standard | 5.21 | 38098.008 | 909254 | 11270.036 | 338.047 | 30.5 | 33 | 8.2
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20240206_HepCl_141 | CC12 | Standard | 5.22 | 33885.363 | 798328 | 10386.876 | 326.232 | 30.5 | 31.84 | 4.4
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20240206_HepCl_021 | CC13 | Standard | 5.16 | 141763.813 | 3414605 | 27395.945 | 517.463 | 48.75 | 50.59 | 3.8
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20240206_HepCl_037 | CC13 | Standard | 5.16 | 133907.734 | 3295336 | 26302.416 | 509.108 | 48.75 | 49.77 | 2.1
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20240206_HepCl_042 | CC13 | Standard | 5.16 | 134816.469 | 3379367 | 25713.629 | 524.3 | 48.75 | 51.26 | 5.1
| Beloq eLoQ | [15 empty] | 20240206_HepCl_047 | CC13 | Standard | 5.17 | 126875.133 | 3033352 | 24279.836 | 522.554 | 48.75 | 51.09 | 4.8
[17 empty] | 20240206_HepCl_064 | CC13 | Standard | 5.18 | 116863.461 | 2871577 | 21099.6 | 553.866 | 48.75 | 54.16 | 11.1
[17 empty] | 20240206_HepCl_080 | CC13 | Standard | 5.19 | 104348.242 | 2556895 | 17471.68 | 597.242 | 48.75 | 58.41 | 19.8
[17 empty] | 20240206_HepCl_085 | CC13 | Standard | 5.19 | 97814.875 | 2321435 | 19143.832 | 510.947 | 48.75 | 49.95 | 2.5
[17 empty] | 20240206_HepCl_090 | CC13 | Standard | 5.2 | 95417.172 | 2251454 | 17418.449 | 547.794 | 48.75 | 53.56 | 9.9
[17 empty] | 20240206_HepCl_022 | CC14 | Standard | 5.17 | 205663.422 | 4891867 | 26067.559 | 788.963 | 78.125 | 77.21 | -1.2
[17 empty] | 20240206_HepCl_065 | CC14 | Standard | 5.18 | 160091.453 | 3896586 | 19912.988 | 803.955 | 78.125 | 78.68 | 0.7
[17 empty] | 20240206_HepCl_023 | CC15 | Standard | 5.16 | 309681.5 | 7288266 | 23888.445 | 1296.365 | 125 | 126.95 | 1.6
[17 empty] | 20240206_HepCl_066 | CC15 | Standard | 5.18 | 236571.453 | 5677405 | 18836.156 | 1255.943 | 125 | 122.99 | -1.6
[17 empty] | 20240206_HepCl_024 | CC16 | Standard | 5.16 | 749450.375 | 18072108 | 45147.977 | 1659.987 | 200 | 162.6 | -18.7
[17 empty] | 20240206_HepCl_067 | CC16 | Standard | 5.18 | 634856.063 | 14937564 | 37734.23 | 1682.441 | 200 | 164.8 | -17.6
[17 empty] | 20240206_HepCl_025 | CC17 | Standard | 5.17 | 1270558.75 | 29727522 | 35427.227 | 3586.391 | 325.625 | 351.46 | 7.9
[17 empty] | 20240206_HepCl_068 | CC17 | Standard | 5.18 | 1065313.625 | 25544994 | 31416.877 | 3390.896 | 325.625 | 332.29 | 2
[17 empty] | 20240206_HepCl_028 | CCQC A | QC | 5.17 | 3432.701 | 82059 | 37207.656 | 9.226 | 0.625 | 0.76 | 22.3
[17 empty] | 20240206_HepCl_071 | CCQC A | QC | 5.18 | 2971.267 | 73145 | 32357.793 | 9.183 | 0.625 | 0.76 | 21.6
[17 empty] | 20240206_HepCl_029 | CCQC B | QC | 5.16 | 11178.361 | 261624 | 34276.102 | 32.613 | 2.506 | 3.06 | 22
[17 empty] | 20240206_HepCl_072 | CCQC B | QC | 5.18 | 9997.105 | 241535 | 30769.189 | 32.491 | 2.506 | 3.04 | 21.5
[17 empty] | 20240206_HepCl_030 | CCQC C | QC | 5.16 | 21085.359 | 512127 | 27290.953 | 77.261 | 6.269 | 7.43 | 18.6
[17 empty] | 20240206_HepCl_073 | CCQC C | QC | 5.18 | 15434.472 | 363637 | 18361.693 | 84.058 | 6.269 | 8.1 | 29.2
[17 empty] | 20240206_HepCl_031 | CCQC D | QC | 5.17 | 130586.773 | 3169813 | 38428.234 | 339.82 | 25 | 33.17 | 32.7
[17 empty] | 20240206_HepCl_074 | CCQC D | QC | 5.18 | 113470.266 | 2831497 | 33615.02 | 337.558 | 25 | 32.95 | 31.8
[17 empty] | 20240206_HepCl_093 | Active T0 A | Analyte | 5.2 | 422083.375 | 10185595 | 76879.266 | 549.021 |  | 53.68
[17 empty] | 20240206_HepCl_094 | Active T0 B | Analyte | 5.2 | 440704.281 | 10793917 | 74541.352 | 591.221 |  | 57.82
[17 empty] | 20240206_HepCl_095 | Active T0 C | Analyte | 5.2 | 380902.75 | 9238476 | 73446.242 | 518.614 |  | 50.7
[17 empty] | 20240206_HepCl_096 | Active T15 A | Analyte | 5.2 | 341947.656 | 8135609 | 71795.5 | 476.28 |  | 46.55
[17 empty] | 20240206_HepCl_097 | Active T15 B | Analyte | 5.2 | 323529.906 | 7916220 | 67235.516 | 481.189 |  | 47.03
[17 empty] | 20240206_HepCl_098 | Active T15 C | Analyte | 5.2 | 307835.531 | 7611555 | 59377.094 | 518.442 |  | 50.69
[17 empty] | 20240206_HepCl_100 | Active T30 A | Analyte | 5.2 | 296008.063 | 7315709 | 59786.223 | 495.111 |  | 48.4
[17 empty] | 20240206_HepCl_101 | Active T30 B | Analyte | 5.2 | 257113.063 | 6303558 | 60600.289 | 424.277 |  | 41.45
[17 empty] | 20240206_HepCl_102 | Active T30 C | Analyte | 5.2 | 250668 | 6286585 | 64703.051 | 387.413 |  | 37.84
[17 empty] | 20240206_HepCl_103 | Active T60 A | Analyte | 5.2 | 270694.688 | 6658858 | 71666.102 | 377.716 |  | 36.89
[17 empty] | 20240206_HepCl_104 | Active T60 B | Analyte | 5.2 | 272790.063 | 6748657 | 68193.617 | 400.023 |  | 39.08
[17 empty] | 20240206_HepCl_105 | Active T60 C | Analyte | 5.2 | 268667.219 | 6495771 | 66932.391 | 401.401 |  | 39.21
[17 empty] | 20240206_HepCl_109 | Active T120 A | Analyte | 5.2 | 338326.813 | 8293125 | 67789.539 | 499.084 |  | 48.79
[17 empty] | 20240206_HepCl_110 | Active T120 B | Analyte | 5.2 | 309013.594 | 7511372 | 70369.023 | 439.133 |  | 42.91
[17 empty] | 20240206_HepCl_111 | Active T120 C | Analyte | 5.2 | 304656.094 | 7518917 | 72949.461 | 417.626 |  | 40.8
[17 empty] | 20240206_HepCl_112 | Active T240 A | Analyte | 5.2 | 324357.844 | 7814818 | 69046.656 | 469.766 |  | 45.91
[17 empty] | 20240206_HepCl_113 | Active T240 B | Analyte | 5.2 | 311989.313 | 7394863 | 69858.797 | 446.6 |  | 43.64
[17 empty] | 20240206_HepCl_114 | Active T240 C | Analyte | 5.2 | 318822.563 | 7843473 | 72495.445 | 439.783 |  | 42.97
[17 empty] | 20240206_HepCl_118 | Media T0 A | Analyte | 5.2 | 338138.125 | 8250365 | 62798.77 | 538.447 |  | 52.65
[17 empty] | 20240206_HepCl_119 | Media T0 B | Analyte | 5.2 | 322311.656 | 7793794 | 60950.219 | 528.811 |  | 51.7
[17 empty] | 20240206_HepCl_120 | Media T0 C | Analyte | 5.2 | 336100.875 | 8085665 | 61018.313 | 550.82 |  | 53.86
[17 empty] | 20240206_HepCl_121 | Media T15 A | Analyte | 5.21 | 357696.781 | 8728060 | 53439.043 | 669.355 |  | 65.48
[17 empty] | 20240206_HepCl_122 | Media T15 B | Analyte | 5.21 | 331523.156 | 8058599 | 55735.617 | 594.814 |  | 58.17
[17 empty] | 20240206_HepCl_123 | Media T15 C | Analyte | 5.21 | 245906.297 | 5946874 | 58680.23 | 419.062 |  | 40.94
[17 empty] | 20240206_HepCl_125 | Media T30 A | Analyte | 5.21 | 193000.906 | 4462446 | 48207.898 | 400.351 |  | 39.11
[17 empty] | 20240206_HepCl_126 | Media T30 B | Analyte | 5.21 | 188581.766 | 4653999 | 54628.371 | 345.208 |  | 33.7
[17 empty] | 20240206_HepCl_127 | Media T30 C | Analyte | 5.21 | 186013.891 | 4614897 | 57590.84 | 322.992 |  | 31.52
[17 empty] | 20240206_HepCl_128 | Media T60 A | Analyte | 5.21 | 342255.781 | 8209658 | 62483.848 | 547.751 |  | 53.56
[17 empty] | 20240206_HepCl_129 | Media T60 B | Analyte | 5.21 | 303342.594 | 7345818 | 54148.367 | 560.206 |  | 54.78
[17 empty] | 20240206_HepCl_130 | Media T60 C | Analyte | 5.21 | 281791.906 | 6948520 | 55963.391 | 503.529 |  | 49.22
[17 empty] | 20240206_HepCl_134 | Media T120 A | Analyte | 5.21 | 324197.281 | 8006442 | 69880.008 | 463.934 |  | 45.34
[17 empty] | 20240206_HepCl_135 | Media T120 B | Analyte | 5.21 | 308013.719 | 7359169 | 67209.852 | 458.287 |  | 44.79
[17 empty] | 20240206_HepCl_136 | Media T120 C | Analyte | 5.21 | 296929.656 | 7207137 | 69773.344 | 425.563 |  | 41.58
[17 empty] | 20240206_HepCl_137 | Media T240 A | Analyte | 5.22 | 322356.75 | 7625099 | 66457.195 | 485.059 |  | 47.41
[17 empty] | 20240206_HepCl_138 | Media T240 B | Analyte | 5.21 | 305873.188 | 7348542 | 61993.867 | 493.393 |  | 48.23
[17 empty] | 20240206_HepCl_139 | Media T240 C | Analyte | 5.21 | 310387.938 | 7491211 | 67426.344 | 460.336 |  | 44.99
[17 empty] | 20240206_HepCl_143 | Inactive T0 A | Analyte | 5.21 | 342590.844 | 8160378 | 76440.156 | 448.182 |  | 43.8
[17 empty] | 20240206_HepCl_144 | Inactive T0 B | Analyte | 5.22 | 312290.625 | 7718516 | 78526.922 | 397.686 |  | 38.85
[17 empty] | 20240206_HepCl_145 | Inactive T0 C | Analyte | 5.22 | 284597.656 | 7016211 | 74402.867 | 382.509 |  | 37.36
[17 empty] | 20240206_HepCl_146 | Inactive T240 A | Analyte | 5.22 | 267962 | 6510039 | 69915.516 | 383.265 |  | 37.43
[17 empty] | 20240206_HepCl_147 | Inactive T240 B | Analyte | 5.22 | 266627.656 | 6609631 | 76102.867 | 350.352 |  | 34.21
[17 empty] | 20240206_HepCl_148 | Inactive T240 C | Analyte | 5.22 | 237742.609 | 5822674 | 72467.32 | 328.069 |  | 32.02
[17 empty] | 20240206_HepCl_004 | Crash Blank | Blank | 5.17 | 64.09 | 314 | 25.148 | 254.851 |  | 24.84
[17 empty] | 20240206_HepCl_003 | Matrix Blank | Blank |  |  |  | 16.108
[17 empty] | 20240206_HepCl_001 | Blank | Blank | 5.17 | 109.911 | 1041 | 42.896 | 256.227 |  | 24.98
[17 empty] | 20240206_HepCl_002 | Blank | Blank | 5.16 | 24.497 | 556 | 17.555 | 139.544 |  | 13.54
[17 empty] | 20240206_HepCl_005 | Blank | Blank | 5.34 | 20.382 | 174 | 5.246 | 388.525 |  | 37.95
[17 empty] | 20240206_HepCl_006 | Blank | Blank | 5.05 | 3.732 | 137 | 4.082 | 91.426 |  | 8.82
[17 empty] | 20240206_HepCl_013 | Blank | Blank | 5.19 | 15.857 | 175 | 17.19 | 92.245 |  | 8.9
[17 empty] | 20240206_HepCl_020 | Blank | Blank | 5.17 | 23.177 | 373 | 7.525 | 308 |  | 30.05
[17 empty] | 20240206_HepCl_026 | Blank | Blank | 5.16 | 246.718 | 3720 | 6.313 | 3908.094 |  | 383
[17 empty] | 20240206_HepCl_027 | Blank | Blank | 5.16 | 106.462 | 944 | 36.603 | 290.856 |  | 28.37
[17 empty] | 20240206_HepCl_032 | Blank | Blank | 5.16 | 21.903 | 545 | 10.67 | 205.276 |  | 19.98
[17 empty] | 20240206_HepCl_033 | Blank | Blank | 5.16 | 66.152 | 425 | 14.221 | 465.171 |  | 45.46
[17 empty] | 20240206_HepCl_038 | Blank | Blank | 5.16 | 56.02 | 487 | 14.408 | 388.812 |  | 37.98
[17 empty] | 20240206_HepCl_043 | Blank | Blank | 5.16 | 28.789 | 384 | 11.28 | 255.222 |  | 24.88
[17 empty] | 20240206_HepCl_048 | Blank | Blank | 5.16 | 14.72 | 369 | 12.853 | 114.526 |  | 11.09
[17 empty] | 20240206_HepCl_049 | Blank | Blank | 5.09 | 5.489 | 174 | 2.151 | 255.184 |  | 24.88
[17 empty] | 20240206_HepCl_056 | Blank | Blank | 5.18 | 9.435 | 176 | 22.688 | 41.586 |  | 3.94
[17 empty] | 20240206_HepCl_063 | Blank | Blank | 5.18 | 16.734 | 442 | 0.246 | 6802.439 |  | 666.75
[17 empty] | 20240206_HepCl_069 | Blank | Blank | 5.18 | 180.023 | 3059 | 41.587 | 432.883 |  | 42.3
[17 empty] | 20240206_HepCl_070 | Blank | Blank | 5.18 | 118.174 | 1030 | 18.371 | 643.264 |  | 62.92
[17 empty] | 20240206_HepCl_075 | Blank | Blank | 5.19 | 9.891 | 317 | 11.861 | 83.391 |  | 8.03
[17 empty] | 20240206_HepCl_076 | Blank | Blank | 5.17 | 32.667 | 298 | 12.105 | 269.864 |  | 26.32
[17 empty] | 20240206_HepCl_081 | Blank | Blank | 5.19 | 16.401 | 424 | 23.058 | 71.129 |  | 6.83
[17 empty] | 20240206_HepCl_086 | Blank | Blank | 5.19 | 56.057 | 655 | 6.612 | 847.807 |  | 82.98
[17 empty] | 20240206_HepCl_091 | Blank | Blank | 5.2 | 60.385 | 584 | 33.975 | 177.734 |  | 17.28
[17 empty] | 20240206_HepCl_092 | Blank | Blank | 5.2 | 4.567 | 152 | 27.48 | 16.619 |  | 1.49
[17 empty] | 20240206_HepCl_099 | Blank | Blank | 5.2 | 64.883 | 768 | 11.866 | 546.798 |  | 53.47
[17 empty] | 20240206_HepCl_106 | Blank | Blank | 5.2 | 63.298 | 885 | 14.916 | 424.363 |  | 41.46
[17 empty] | 20240206_HepCl_108 | Blank | Blank | 5.19 | 34.506 | 428 | 24.083 | 143.279 |  | 13.91
[17 empty] | 20240206_HepCl_115 | Blank | Blank | 5.2 | 54.096 | 305 | 8.215 | 658.503 |  | 64.42
[17 empty] | 20240206_HepCl_117 | Blank | Blank | 5.2 | 35.661 | 504 | 6.233 | 572.132 |  | 55.95
[17 empty] | 20240206_HepCl_124 | Blank | Blank | 5.21 | 96.891 | 1086 | 39.24 | 246.919 |  | 24.07
[17 empty] | 20240206_HepCl_131 | Blank | Blank | 5.22 | 92.428 | 994 | 30.685 | 301.216 |  | 29.39
[17 empty] | 20240206_HepCl_133 | Blank | Blank | 5.21 | 16.904 | 414 | 4.249 | 397.835 |  | 38.86
[17 empty] | 20240206_HepCl_140 | Blank | Blank | 5.22 | 52.428 | 1022 | 31.17 | 168.2 |  | 16.35
[17 empty] | 20240206_HepCl_142 | Blank | Blank | 5.2 | 21.74 | 410 | 4.231 | 513.827 |  | 50.23
[17 empty] | 20240206_HepCl_149 | Blank | Blank | 5.21 | 101.697 | 1052 | 12.896 | 788.593 |  | 77.17
[17 empty] | 20240206_HepCl_150 | Blank | Blank | 5.23 | 27.967 | 472 | 23.086 | 121.143 |  | 11.74
[17 empty] | 20240206_HepCl_151 | Blank | Blank | 5.16 | 1.968 | 118 | 62.664 | 3.141 |  | 0.17
Compound 6: M8FOSA
[17 empty] | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
[17 empty] | 20240206_HepCl_007 | CC01 | Standard | 5.17 | 28015.162 | 764846 |  | 28015.162 | 100 | 121.54 | 21.5
[17 empty] | 20240206_HepCl_050 | CC01 | Standard | 5.17 | 24507.627 | 671672 |  | 24507.627 | 100 | 106.32 | 6.3
[17 empty] | 20240206_HepCl_008 | CC02 | Standard | 5.17 | 29313.949 | 810743 |  | 29313.949 | 100 | 127.17 | 27.2
[17 empty] | 20240206_HepCl_051 | CC02 | Standard | 5.17 | 23632.779 | 648234 |  | 23632.779 | 100 | 102.53 | 2.5
[17 empty] | 20240206_HepCl_009 | CC03 | Standard | 5.17 | 29950.836 | 830589 |  | 29950.836 | 100 | 129.94 | 29.9
[17 empty] | 20240206_HepCl_052 | CC03 | Standard | 5.17 | 23528.35 | 650772 |  | 23528.35 | 100 | 102.07 | 2.1
[17 empty] | 20240206_HepCl_010 | CC04 | Standard | 5.17 | 28339.805 | 775903 |  | 28339.805 | 100 | 122.95 | 22.9
[17 empty] | 20240206_HepCl_034 | CC04 | Standard | 5.16 | 24535.77 | 665024 |  | 24535.77 | 100 | 106.44 | 6.4
[17 empty] | 20240206_HepCl_039 | CC04 | Standard | 5.17 | 26556.42 | 744566 |  | 26556.42 | 100 | 115.21 | 15.2
[17 empty] | 20240206_HepCl_044 | CC04 | Standard | 5.17 | 25185.447 | 693300 |  | 25185.447 | 100 | 109.26 | 9.3
[17 empty] | 20240206_HepCl_053 | CC04 | Standard | 5.17 | 23608.605 | 645217 |  | 23608.605 | 100 | 102.42 | 2.4
[17 empty] | 20240206_HepCl_077 | CC04 | Standard | 5.18 | 20045.84 | 558513 |  | 20045.84 | 100 | 86.96 | -13
[17 empty] | 20240206_HepCl_082 | CC04 | Standard | 5.19 | 20375.537 | 572269 |  | 20375.537 | 100 | 88.4 | -11.6
[17 empty] | 20240206_HepCl_087 | CC04 | Standard | 5.19 | 19531.541 | 545490 |  | 19531.541 | 100 | 84.73 | -15.3
[17 empty] | 20240206_HepCl_011 | CC05 | Standard | 5.17 | 25832.488 | 724632 |  | 25832.488 | 100 | 112.07 | 12.1
[17 empty] | 20240206_HepCl_035 | CC05 | Standard | 5.16 | 23397 | 644247 |  | 23397 | 100 | 101.5 | 1.5
[17 empty] | 20240206_HepCl_040 | CC05 | Standard | 5.17 | 23445.711 | 645545 |  | 23445.711 | 100 | 101.71 | 1.7
[17 empty] | 20240206_HepCl_045 | CC05 | Standard | 5.16 | 22468.701 | 616974 |  | 22468.701 | 100 | 97.48 | -2.5
[17 empty] | 20240206_HepCl_054 | CC05 | Standard | 5.17 | 22096.361 | 611518 |  | 22096.361 | 100 | 95.86 | -4.1
[17 empty] | 20240206_HepCl_078 | CC05 | Standard | 5.18 | 17964.746 | 493692 |  | 17964.746 | 100 | 77.94 | -22.1
[17 empty] | 20240206_HepCl_083 | CC05 | Standard | 5.19 | 16582.211 | 447530 |  | 16582.211 | 100 | 71.94 | -28.1
[17 empty] | 20240206_HepCl_088 | CC05 | Standard | 5.19 | 17208.271 | 474056 |  | 17208.271 | 100 | 74.65 | -25.3
[17 empty] | 20240206_HepCl_012 | CC06 | Standard | 5.17 | 26238.695 | 724166 |  | 26238.695 | 100 | 113.83 | 13.8
[17 empty] | 20240206_HepCl_036 | CC06 | Standard | 5.16 | 23039.303 | 625832 |  | 23039.303 | 100 | 99.95 | 0
[17 empty] | 20240206_HepCl_041 | CC06 | Standard | 5.17 | 24264.488 | 682297 |  | 24264.488 | 100 | 105.27 | 5.3
[17 empty] | 20240206_HepCl_046 | CC06 | Standard | 5.17 | 22817.016 | 625104 |  | 22817.016 | 100 | 98.99 | -1
[17 empty] | 20240206_HepCl_055 | CC06 | Standard | 5.17 | 21123.363 | 582288 |  | 21123.363 | 100 | 91.64 | -8.4
[17 empty] | 20240206_HepCl_079 | CC06 | Standard | 5.19 | 18868.227 | 529578 |  | 18868.227 | 100 | 81.86 | -18.1
[17 empty] | 20240206_HepCl_084 | CC06 | Standard | 5.19 | 16807.939 | 454619 |  | 16807.939 | 100 | 72.92 | -27.1
[17 empty] | 20240206_HepCl_089 | CC06 | Standard | 5.19 | 16733.023 | 463597 |  | 16733.023 | 100 | 72.59 | -27.4
[17 empty] | 20240206_HepCl_014 | CC07 | Standard | 5.17 | 27822.641 | 750505 |  | 27822.641 | 100 | 120.7 | 20.7
[17 empty] | 20240206_HepCl_057 | CC07 | Standard | 5.17 | 22068.023 | 603914 |  | 22068.023 | 100 | 95.74 | -4.3
[17 empty] | 20240206_HepCl_015 | CC08 | Standard | 5.17 | 27143.525 | 718182 |  | 27143.525 | 100 | 117.76 | 17.8
[17 empty] | 20240206_HepCl_058 | CC08 | Standard | 5.17 | 22590.502 | 634597 |  | 22590.502 | 100 | 98 | -2
[17 empty] | 20240206_HepCl_016 | CC09 | Standard | 5.17 | 17637.496 | 472292 |  | 17637.496 | 100 | 76.52 | -23.5
[17 empty] | 20240206_HepCl_059 | CC09 | Standard | 5.17 | 14474.795 | 396937 |  | 14474.795 | 100 | 62.8 | -37.2
[17 empty] | 20240206_HepCl_017 | CC10 | Standard | 5.17 | 26663.059 | 733176 |  | 26663.059 | 100 | 115.67 | 15.7
[17 empty] | 20240206_HepCl_060 | CC10 | Standard | 5.17 | 21487.777 | 594106 |  | 21487.777 | 100 | 93.22 | -6.8
[17 empty] | 20240206_HepCl_018 | CC11 | Standard | 5.17 | 27305.922 | 747337 |  | 27305.922 | 100 | 118.46 | 18.5
[17 empty] | 20240206_HepCl_061 | CC11 | Standard | 5.18 | 22796.053 | 633107 |  | 22796.053 | 100 | 98.9 | -1.1
[17 empty] | 20240206_HepCl_019 | CC12 | Standard | 5.17 | 27131.459 | 745667 |  | 27131.459 | 100 | 117.7 | 17.7
[17 empty] | 20240206_HepCl_062 | CC12 | Standard | 5.18 | 20643.158 | 562092 |  | 20643.158 | 100 | 89.56 | -10.4
[17 empty] | 20240206_HepCl_107 | CC12 | Standard | 5.2 | 14109.446 | 384397 |  | 14109.446 | 100 | 61.21 | -38.8
[17 empty] | 20240206_HepCl_116 | CC12 | Standard | 5.2 | 12328.883 | 348261 |  | 12328.883 | 100 | 53.49 | -46.5
[17 empty] | 20240206_HepCl_132 | CC12 | Standard | 5.21 | 11270.036 | 304960 |  | 11270.036 | 100 | 48.89 | -51.1
[17 empty] | 20240206_HepCl_141 | CC12 | Standard | 5.21 | 10386.876 | 280167 |  | 10386.876 | 100 | 45.06 | -54.9
[17 empty] | 20240206_HepCl_021 | CC13 | Standard | 5.17 | 27395.945 | 753648 |  | 27395.945 | 100 | 118.85 | 18.9
[17 empty] | 20240206_HepCl_037 | CC13 | Standard | 5.16 | 26302.416 | 722371 |  | 26302.416 | 100 | 114.11 | 14.1
[17 empty] | 20240206_HepCl_042 | CC13 | Standard | 5.16 | 25713.629 | 715980 |  | 25713.629 | 100 | 111.55 | 11.6
[17 empty] | 20240206_HepCl_047 | CC13 | Standard | 5.17 | 24279.836 | 662569 |  | 24279.836 | 100 | 105.33 | 5.3
[17 empty] | 20240206_HepCl_064 | CC13 | Standard | 5.18 | 21099.6 | 583234 |  | 21099.6 | 100 | 91.54 | -8.5
[17 empty] | 20240206_HepCl_080 | CC13 | Standard | 5.19 | 17471.68 | 464999 |  | 17471.68 | 100 | 75.8 | -24.2
[17 empty] | 20240206_HepCl_085 | CC13 | Standard | 5.19 | 19143.832 | 518020 |  | 19143.832 | 100 | 83.05 | -16.9
[17 empty] | 20240206_HepCl_090 | CC13 | Standard | 5.19 | 17418.449 | 459441 |  | 17418.449 | 100 | 75.57 | -24.4
[17 empty] | 20240206_HepCl_022 | CC14 | Standard | 5.17 | 26067.559 | 716640 |  | 26067.559 | 100 | 113.09 | 13.1
[17 empty] | 20240206_HepCl_065 | CC14 | Standard | 5.18 | 19912.988 | 553677 |  | 19912.988 | 100 | 86.39 | -13.6
[17 empty] | 20240206_HepCl_023 | CC15 | Standard | 5.17 | 23888.445 | 647062 |  | 23888.445 | 100 | 103.64 | 3.6
[17 empty] | 20240206_HepCl_066 | CC15 | Standard | 5.18 | 18836.156 | 503821 |  | 18836.156 | 100 | 81.72 | -18.3
[17 empty] | 20240206_HepCl_024 | CC16 | Standard | 5.17 | 45147.977 | 1246717 |  | 45147.977 | 100 | 195.87 | 95.9
[17 empty] | 20240206_HepCl_067 | CC16 | Standard | 5.18 | 37734.23 | 1023848 |  | 37734.23 | 100 | 163.7 | 63.7
[17 empty] | 20240206_HepCl_025 | CC17 | Standard | 5.17 | 35427.227 | 936213 |  | 35427.227 | 100 | 153.69 | 53.7
[17 empty] | 20240206_HepCl_068 | CC17 | Standard | 5.18 | 31416.877 | 849521 |  | 31416.877 | 100 | 136.3 | 36.3
[17 empty] | 20240206_HepCl_028 | CCQC A | QC | 5.17 | 37207.656 | 1028814 |  | 37207.656 | 100 | 161.42 | 61.4
[17 empty] | 20240206_HepCl_071 | CCQC A | QC | 5.18 | 32357.793 | 885092 |  | 32357.793 | 100 | 140.38 | 40.4
[17 empty] | 20240206_HepCl_029 | CCQC B | QC | 5.17 | 34276.102 | 919331 |  | 34276.102 | 100 | 148.7 | 48.7
[17 empty] | 20240206_HepCl_072 | CCQC B | QC | 5.18 | 30769.189 | 828347 |  | 30769.189 | 100 | 133.49 | 33.5
[17 empty] | 20240206_HepCl_030 | CCQC C | QC | 5.17 | 27290.953 | 756072 |  | 27290.953 | 100 | 118.4 | 18.4
[17 empty] | 20240206_HepCl_073 | CCQC C | QC | 5.18 | 18361.693 | 481360 |  | 18361.693 | 100 | 79.66 | -20.3
[17 empty] | 20240206_HepCl_031 | CCQC D | QC | 5.17 | 38428.234 | 1065680 |  | 38428.234 | 100 | 166.71 | 66.7
[17 empty] | 20240206_HepCl_074 | CCQC D | QC | 5.18 | 33615.02 | 927726 |  | 33615.02 | 100 | 145.83 | 45.8
[17 empty] | 20240206_HepCl_093 | Active T0 A | Analyte | 5.19 | 76879.266 | 2077377 |  | 76879.266 | 100 | 333.53 | 233.5
[17 empty] | 20240206_HepCl_094 | Active T0 B | Analyte | 5.19 | 74541.352 | 2025422 |  | 74541.352 | 100 | 323.38 | 223.4
[17 empty] | 20240206_HepCl_095 | Active T0 C | Analyte | 5.19 | 73446.242 | 1989905 |  | 73446.242 | 100 | 318.63 | 218.6
[17 empty] | 20240206_HepCl_096 | Active T15 A | Analyte | 5.2 | 71795.5 | 1924504 |  | 71795.5 | 100 | 311.47 | 211.5
[17 empty] | 20240206_HepCl_097 | Active T15 B | Analyte | 5.19 | 67235.516 | 1770558 |  | 67235.516 | 100 | 291.69 | 191.7
[17 empty] | 20240206_HepCl_098 | Active T15 C | Analyte | 5.19 | 59377.094 | 1611156 |  | 59377.094 | 100 | 257.6 | 157.6
[17 empty] | 20240206_HepCl_100 | Active T30 A | Analyte | 5.2 | 59786.223 | 1646308 |  | 59786.223 | 100 | 259.37 | 159.4
[17 empty] | 20240206_HepCl_101 | Active T30 B | Analyte | 5.2 | 60600.289 | 1649459 |  | 60600.289 | 100 | 262.9 | 162.9
[17 empty] | 20240206_HepCl_102 | Active T30 C | Analyte | 5.2 | 64703.051 | 1782430 |  | 64703.051 | 100 | 280.7 | 180.7
[17 empty] | 20240206_HepCl_103 | Active T60 A | Analyte | 5.2 | 71666.102 | 1965729 |  | 71666.102 | 100 | 310.91 | 210.9
[17 empty] | 20240206_HepCl_104 | Active T60 B | Analyte | 5.2 | 68193.617 | 1893662 |  | 68193.617 | 100 | 295.84 | 195.8
[17 empty] | 20240206_HepCl_105 | Active T60 C | Analyte | 5.2 | 66932.391 | 1831529 |  | 66932.391 | 100 | 290.37 | 190.4
[17 empty] | 20240206_HepCl_109 | Active T120 A | Analyte | 5.2 | 67789.539 | 1909254 |  | 67789.539 | 100 | 294.09 | 194.1
[17 empty] | 20240206_HepCl_110 | Active T120 B | Analyte | 5.2 | 70369.023 | 1907301 |  | 70369.023 | 100 | 305.28 | 205.3
[17 empty] | 20240206_HepCl_111 | Active T120 C | Analyte | 5.2 | 72949.461 | 2025578 |  | 72949.461 | 100 | 316.48 | 216.5
[17 empty] | 20240206_HepCl_112 | Active T240 A | Analyte | 5.2 | 69046.656 | 1919022 |  | 69046.656 | 100 | 299.55 | 199.5
[17 empty] | 20240206_HepCl_113 | Active T240 B | Analyte | 5.2 | 69858.797 | 1869751 |  | 69858.797 | 100 | 303.07 | 203.1
[17 empty] | 20240206_HepCl_114 | Active T240 C | Analyte | 5.2 | 72495.445 | 2027736 |  | 72495.445 | 100 | 314.51 | 214.5
[17 empty] | 20240206_HepCl_118 | Media T0 A | Analyte | 5.2 | 62798.77 | 1682330 |  | 62798.77 | 100 | 272.44 | 172.4
[17 empty] | 20240206_HepCl_119 | Media T0 B | Analyte | 5.2 | 60950.219 | 1672878 |  | 60950.219 | 100 | 264.42 | 164.4
[17 empty] | 20240206_HepCl_120 | Media T0 C | Analyte | 5.2 | 61018.313 | 1648730 |  | 61018.313 | 100 | 264.72 | 164.7
[17 empty] | 20240206_HepCl_121 | Media T15 A | Analyte | 5.2 | 53439.043 | 1492113 |  | 53439.043 | 100 | 231.83 | 131.8
[17 empty] | 20240206_HepCl_122 | Media T15 B | Analyte | 5.2 | 55735.617 | 1534875 |  | 55735.617 | 100 | 241.8 | 141.8
[17 empty] | 20240206_HepCl_123 | Media T15 C | Analyte | 5.21 | 58680.23 | 1601727 |  | 58680.23 | 100 | 254.57 | 154.6
[17 empty] | 20240206_HepCl_125 | Media T30 A | Analyte | 5.21 | 48207.898 | 1287356 |  | 48207.898 | 100 | 209.14 | 109.1
[17 empty] | 20240206_HepCl_126 | Media T30 B | Analyte | 5.21 | 54628.371 | 1511146 |  | 54628.371 | 100 | 236.99 | 137
[17 empty] | 20240206_HepCl_127 | Media T30 C | Analyte | 5.21 | 57590.84 | 1610994 |  | 57590.84 | 100 | 249.85 | 149.8
[17 empty] | 20240206_HepCl_128 | Media T60 A | Analyte | 5.21 | 62483.848 | 1706733 |  | 62483.848 | 100 | 271.07 | 171.1
[17 empty] | 20240206_HepCl_129 | Media T60 B | Analyte | 5.21 | 54148.367 | 1483369 |  | 54148.367 | 100 | 234.91 | 134.9
[17 empty] | 20240206_HepCl_130 | Media T60 C | Analyte | 5.21 | 55963.391 | 1539156 |  | 55963.391 | 100 | 242.79 | 142.8
[17 empty] | 20240206_HepCl_134 | Media T120 A | Analyte | 5.21 | 69880.008 | 1912559 |  | 69880.008 | 100 | 303.16 | 203.2
[17 empty] | 20240206_HepCl_135 | Media T120 B | Analyte | 5.21 | 67209.852 | 1777711 |  | 67209.852 | 100 | 291.58 | 191.6
[17 empty] | 20240206_HepCl_136 | Media T120 C | Analyte | 5.21 | 69773.344 | 1931253 |  | 69773.344 | 100 | 302.7 | 202.7
[17 empty] | 20240206_HepCl_137 | Media T240 A | Analyte | 5.21 | 66457.195 | 1793467 |  | 66457.195 | 100 | 288.31 | 188.3
[17 empty] | 20240206_HepCl_138 | Media T240 B | Analyte | 5.21 | 61993.867 | 1675997 |  | 61993.867 | 100 | 268.95 | 168.9
[17 empty] | 20240206_HepCl_139 | Media T240 C | Analyte | 5.21 | 67426.344 | 1841450 |  | 67426.344 | 100 | 292.52 | 192.5
[17 empty] | 20240206_HepCl_143 | Inactive T0 A | Analyte | 5.21 | 76440.156 | 2082636 |  | 76440.156 | 100 | 331.62 | 231.6
[17 empty] | 20240206_HepCl_144 | Inactive T0 B | Analyte | 5.22 | 78526.922 | 2211083 |  | 78526.922 | 100 | 340.67 | 240.7
[17 empty] | 20240206_HepCl_145 | Inactive T0 C | Analyte | 5.22 | 74402.867 | 2043815 |  | 74402.867 | 100 | 322.78 | 222.8
[17 empty] | 20240206_HepCl_146 | Inactive T240 A | Analyte | 5.22 | 69915.516 | 1874804 |  | 69915.516 | 100 | 303.31 | 203.3
[17 empty] | 20240206_HepCl_147 | Inactive T240 B | Analyte | 5.22 | 76102.867 | 2106583 |  | 76102.867 | 100 | 330.16 | 230.2
[17 empty] | 20240206_HepCl_148 | Inactive T240 C | Analyte | 5.22 | 72467.32 | 2037564 |  | 72467.32 | 100 | 314.38 | 214.4
[17 empty] | 20240206_HepCl_004 | Crash Blank | Blank | 5.16 | 25.148 | 81 |  | 25.148 | 100 | 0.11 | -99.9
[17 empty] | 20240206_HepCl_003 | Matrix Blank | Blank | 5.34 | 16.108 | 131 |  | 16.108 | 100 | 0.07 | -99.9
[17 empty] | 20240206_HepCl_001 | Blank | Blank | 5.17 | 42.896 | 169 |  | 42.896 | 100 | 0.19 | -99.8
[17 empty] | 20240206_HepCl_002 | Blank | Blank | 5.28 | 17.555 | 100 |  | 17.555 | 100 | 0.08 | -99.9
[17 empty] | 20240206_HepCl_005 | Blank | Blank | 5.16 | 5.246 | 76 |  | 5.246 | 100 | 0.02 | -100
[17 empty] | 20240206_HepCl_006 | Blank | Blank | 5.34 | 4.082 | 87 |  | 4.082 | 100 | 0.02 | -100
[17 empty] | 20240206_HepCl_013 | Blank | Blank | 5.16 | 17.19 | 145 |  | 17.19 | 100 | 0.07 | -99.9
[17 empty] | 20240206_HepCl_020 | Blank | Blank | 5.17 | 7.525 | 265 |  | 7.525 | 100 | 0.03 | -100
[17 empty] | 20240206_HepCl_026 | Blank | Blank | 5.17 | 6.313 | 212 |  | 6.313 | 100 | 0.03 | -100
[17 empty] | 20240206_HepCl_027 | Blank | Blank | 5.08 | 36.603 | 157 |  | 36.603 | 100 | 0.16 | -99.8
[17 empty] | 20240206_HepCl_032 | Blank | Blank | 5.17 | 10.67 | 222 |  | 10.67 | 100 | 0.05 | -100
[17 empty] | 20240206_HepCl_033 | Blank | Blank | 5.05 | 14.221 | 106 |  | 14.221 | 100 | 0.06 | -99.9
[17 empty] | 20240206_HepCl_038 | Blank | Blank | 5.17 | 14.408 | 148 |  | 14.408 | 100 | 0.06 | -99.9
[17 empty] | 20240206_HepCl_043 | Blank | Blank | 5.17 | 11.28 | 138 |  | 11.28 | 100 | 0.05 | -100
[17 empty] | 20240206_HepCl_048 | Blank | Blank | 5.17 | 12.853 | 206 |  | 12.853 | 100 | 0.06 | -99.9
[17 empty] | 20240206_HepCl_049 | Blank | Blank | 5.07 | 2.151 | 58 |  | 2.151 | 100 | 0.01 | -100
[17 empty] | 20240206_HepCl_056 | Blank | Blank | 5.17 | 22.688 | 234 |  | 22.688 | 100 | 0.1 | -99.9
[17 empty] | 20240206_HepCl_063 | Blank | Blank | 5.16 | 0.246 | 22 |  | 0.246 | 100 | 0 | -100
[17 empty] | 20240206_HepCl_069 | Blank | Blank | 5.19 | 41.587 | 250 |  | 41.587 | 100 | 0.18 | -99.8
[17 empty] | 20240206_HepCl_070 | Blank | Blank | 5.18 | 18.371 | 188 |  | 18.371 | 100 | 0.08 | -99.9
[17 empty] | 20240206_HepCl_075 | Blank | Blank | 5.18 | 11.861 | 221 |  | 11.861 | 100 | 0.05 | -99.9
[17 empty] | 20240206_HepCl_076 | Blank | Blank | 5.35 | 12.105 | 107 |  | 12.105 | 100 | 0.05 | -99.9
[17 empty] | 20240206_HepCl_081 | Blank | Blank | 5.18 | 23.058 | 176 |  | 23.058 | 100 | 0.1 | -99.9
[17 empty] | 20240206_HepCl_086 | Blank | Blank | 5.18 | 6.612 | 162 |  | 6.612 | 100 | 0.03 | -100
[17 empty] | 20240206_HepCl_091 | Blank | Blank | 5.19 | 33.975 | 209 |  | 33.975 | 100 | 0.15 | -99.9
[17 empty] | 20240206_HepCl_092 | Blank | Blank | 5.2 | 27.48 | 172 |  | 27.48 | 100 | 0.12 | -99.9
[17 empty] | 20240206_HepCl_099 | Blank | Blank | 5.2 | 11.866 | 286 |  | 11.866 | 100 | 0.05 | -99.9
[17 empty] | 20240206_HepCl_106 | Blank | Blank | 5.19 | 14.916 | 270 |  | 14.916 | 100 | 0.06 | -99.9
[17 empty] | 20240206_HepCl_108 | Blank | Blank | 5.19 | 24.083 | 244 |  | 24.083 | 100 | 0.1 | -99.9
[17 empty] | 20240206_HepCl_115 | Blank | Blank | 5.2 | 8.215 | 206 |  | 8.215 | 100 | 0.04 | -100
[17 empty] | 20240206_HepCl_117 | Blank | Blank | 5.19 | 6.233 | 137 |  | 6.233 | 100 | 0.03 | -100
[17 empty] | 20240206_HepCl_124 | Blank | Blank | 5.21 | 39.24 | 274 |  | 39.24 | 100 | 0.17 | -99.8
[17 empty] | 20240206_HepCl_131 | Blank | Blank | 5.2 | 30.685 | 278 |  | 30.685 | 100 | 0.13 | -99.9
[17 empty] | 20240206_HepCl_133 | Blank | Blank | 5.21 | 4.249 | 113 |  | 4.249 | 100 | 0.02 | -100
[17 empty] | 20240206_HepCl_140 | Blank | Blank | 5.23 | 31.17 | 335 |  | 31.17 | 100 | 0.14 | -99.9
[17 empty] | 20240206_HepCl_142 | Blank | Blank | 5.21 | 4.231 | 124 |  | 4.231 | 100 | 0.02 | -100
[17 empty] | 20240206_HepCl_149 | Blank | Blank | 5.22 | 12.896 | 235 |  | 12.896 | 100 | 0.06 | -99.9
[17 empty] | 20240206_HepCl_150 | Blank | Blank | 5.05 | 23.086 | 196 |  | 23.086 | 100 | 0.1 | -99.9
[17 empty] | 20240206_HepCl_151 | Blank | Blank | 5.21 | 62.664 | 263 |  | 62.664 | 100 | 0.27 | -99.7

## Phenacetin
Human Active Cells | [8 empty] | Media Only | [7 empty] | Compound 1: phenacetin
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20240206_HepClReRun_Ref_007 | CC01 | Analyte | 3.8 | 1171.234 | 41010 | 99147.227 | 1.181 | 0.175
0 | Active T0 A | 58.02 | 7.2388 | 1392.48 | 1457.12 | 96.3995 | 6.616% |  | 0 | Media T0 A | 113.33 | 7.9084 | 2719.92 | 2207.68 | 479.2383 | 21.71% | 20240206_HepClReRun_Ref_050 | CC01 | Analyte | 3.8 | 1288.76 | 44314 | 91640.219 | 1.406 | 0.175
| Active T0 B | 58.79 | 7.252 | 1410.96 | [5 empty] | Media T0 B | 88.87 | 7.6652 | 2132.88 |  |  |  | 20240206_HepClReRun_Ref_008 | CC02 | Analyte | 3.8 | 2217.497 | 81953 | 110029.922 | 2.015 | 0.281
| Active T0 C | 65.33 | 7.3575 | 1567.92 | [5 empty] | Media T0 C | 73.76 | 7.4789 | 1770.24 |  |  |  | 20240206_HepClReRun_Ref_051 | CC02 | Analyte | 3.8 | 2885.166 | 107134 | 118718.242 | 2.43 | 0.281
15 | Active T15 A | 52.96 | 7.1476 | 1271.04 | 1023.28 | 215.107 | 21.02% |  | 15 | Media T15 A | 44.59 | 6.9756 | 1070.16 | 1186.64 | 299.7251 | 25.26% | 20240206_HepClReRun_Ref_009 | CC03 | Analyte | 3.8 | 3997.633 | 141888 | 97847.359 | 4.086 | 0.444 | 0.08 | -82.7
| Active T15 B | 36.84 | 6.7846 | 884.16 | [5 empty] | Media T15 B | 40.11 | 6.8697 | 962.64 |  |  |  | 20240206_HepClReRun_Ref_052 | CC03 | Analyte | 3.8 | 4851.163 | 177036 | 104661.281 | 4.635 | 0.444 | 0.15 | -66.4
| Active T15 C | 38.11 | 6.8185 | 914.64 | [5 empty] | Media T15 C | 63.63 | 7.3311 | 1527.12 |  |  |  | 20240206_HepClReRun_Ref_010 | CC04 | Analyte | 3.8 | 3947.641 | 141588 | 56906.484 | 6.937 | 0.713 | 0.45 | -36.6
30 | Active T30 A | 42.11 | 6.9183 | 1010.64 | 953.84 | 157.8597 | 16.55% |  | 30 | Media T30 A | 83.75 | 7.6059 | 2010 | 1932.56 | 278.4374 | 14.41% | 20240206_HepClReRun_Ref_034 | CC04 | Analyte | 3.8 | 5296.986 | 190711 | 50823.348 | 10.422 | 0.713 | 0.91 | 27.8
| Active T30 B | 44.81 | 6.9805 | 1075.44 | [5 empty] | Media T30 B | 90.17 | 7.6798 | 2164.08 |  |  |  | 20240206_HepClReRun_Ref_039 | CC04 | Analyte | 3.8 | 4708.708 | 168591 | 45127.563 | 10.434 | 0.713 | 0.91 | 28
| Active T30 C | 32.31 | 6.6534 | 775.44 | [5 empty] | Media T30 C | 67.65 | 7.3924 | 1623.6 |  |  |  | 20240206_HepClReRun_Ref_044 | CC04 | Analyte | 3.8 | 4851.796 | 178003 | 50215.895 | 9.662 | 0.713 | 0.81 | 13.8
60 | Active T60 A | 33.12 | 6.6782 | 794.88 | 794.8 | 96.36 | 12.12% |  | 60 | Media T60 A | 84.63 | 7.6163 | 2031.12 | 2060.32 | 359.3309 | 17.44% | 20240206_HepClReRun_Ref_053 | CC04 | Analyte | 3.8 | 5083.494 | 184147 | 51603.223 | 9.851 | 0.713 | 0.84 | 17.3
| Active T60 B | 29.1 | 6.5488 | 698.4 | [5 empty] | Media T60 B | 71.52 | 7.448 | 1716.48 |  |  |  | 20240206_HepClReRun_Ref_077 | CC04 | Analyte | 3.8 | 5318.31 | 191308 | 55185.578 | 9.637 | 0.713 | 0.81 | 13.3
| Active T60 C | 37.13 | 6.7925 | 891.12 | [5 empty] | Media T60 C | 101.39 | 7.797 | 2433.36 |  |  |  | 20240206_HepClReRun_Ref_082 | CC04 | Analyte | 3.8 | 5480.817 | 197308 | 57444.809 | 9.541 | 0.713 | 0.79 | 11.5
120 | Active T120 A | 42.82 | 6.9351 | 1027.68 | 840.32 | 162.2841 | 19.31% |  | 120 | Media T120 A | 73.95 | 7.4814 | 1774.8 | 1421.28 | 309.6594 | 21.79% | 20240206_HepClReRun_Ref_087 | CC04 | Analyte | 3.8 | 4596.38 | 170385 | 46731.051 | 9.836 | 0.713 | 0.83 | 17
| Active T120 B | 30.99 | 6.6117 | 743.76 | [5 empty] | Media T120 B | 53.79 | 7.1631 | 1290.96 |  |  |  | 20240206_HepClReRun_Ref_011 | CC05 | Standard | 3.8 | 7222.345 | 258879 | 62519.641 | 11.552 | 1.138 | 1.06 | -6.9
| Active T120 C | 31.23 | 6.6194 | 749.52 | [5 empty] | Media T120 C | 49.92 | 7.0885 | 1198.08 |  |  |  | 20240206_HepClReRun_Ref_035 | CC05 | Standard | 3.8 | 9176.8 | 320384 | 81373.328 | 11.277 | 1.138 | 1.02 | -10.1
240 | Active T240 A | 18.61 | 6.1018 | 446.64 | 422.64 | 50.9886 | 12.06% |  | 240 | Media T240 A | 52.01 | 7.1295 | 1248.24 | 1373.44 | 217.6847 | 15.85% | 20240206_HepClReRun_Ref_040 | CC05 | Standard | 3.8 | 8076.035 | 289767 | 66556.109 | 12.134 | 1.138 | 1.14 | -0.2
| Active T240 B | 19.05 | 6.1251 | 457.2 | [5 empty] | Media T240 B | 51.97 | 7.1287 | 1247.28 |  |  |  | 20240206_HepClReRun_Ref_045 | CC05 | Standard | 3.8 | 8571.604 | 301937 | 75406.516 | 11.367 | 1.138 | 1.03 | -9
| Active T240 C | 15.17 | 5.8974 | 364.08 | [5 empty] | Media T240 C | 67.7 | 7.3931 | 1624.8 |  |  |  | 20240206_HepClReRun_Ref_054 | CC05 | Standard | 3.8 | 8625.105 | 303739 | 76128.5 | 11.33 | 1.138 | 1.03 | -9.5
[17 empty] | 20240206_HepClReRun_Ref_078 | CC05 | Standard | 3.8 | 8629.438 | 304052 | 80799.367 | 10.68 | 1.138 | 0.94 | -17
Human Inactive Cells | [16 empty] | 20240206_HepClReRun_Ref_083 | CC05 | Standard | 3.8 | 9571.905 | 336232 | 73697.18 | 12.988 | 1.138 | 1.25 | 9.7
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20240206_HepClReRun_Ref_088 | CC05 | Standard | 3.8 | 8446.699 | 304756 | 71098.961 | 11.88 | 1.138 | 1.1 | -3.1
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20240206_HepClReRun_Ref_012 | CC06 | Standard | 3.8 | 12805.168 | 459773 | 88497.523 | 14.47 | 1.825 | 1.44 | -20.9
0 | Inactive T0 A | 74.06 | 7.4829 | 1777.44 | 1798.8 | 294.4616 | 16.37% | [9 empty] | 20240206_HepClReRun_Ref_036 | CC06 | Standard | 3.8 | 15669.537 | 564795 | 96763.203 | 16.194 | 1.825 | 1.67 | -8.5
| Inactive T0 B | 63.15 | 7.3236 | 1515.6 | [12 empty] | 20240206_HepClReRun_Ref_041 | CC06 | Standard | 3.8 | 14739.754 | 527682 | 86041.914 | 17.131 | 1.825 | 1.79 | -1.7
| Inactive T0 C | 87.64 | 7.6513 | 2103.36 | [12 empty] | 20240206_HepClReRun_Ref_046 | CC06 | Standard | 3.8 | 16642.393 | 599708 | 89999.438 | 18.492 | 1.825 | 1.97 | 8.1
240 | Inactive T240 A | 66.86 | 7.3807 | 1604.64 | 1918.44 | 443.7802 | 23.13% | [9 empty] | 20240206_HepClReRun_Ref_055 | CC06 | Standard | 3.8 | 17092.779 | 607234 | 86843.813 | 19.682 | 1.825 | 2.13 | 16.6
| Inactive T240 B | 93.01 | 7.7108 | 2232.24 | [12 empty] | 20240206_HepClReRun_Ref_079 | CC06 | Standard | 3.8 | 17672.369 | 624694 | 93015.078 | 18.999 | 1.825 | 2.04 | 11.7
| Inactive T240 C | 75.46 | 7.5017 | 1811.04 | [12 empty] | 20240206_HepClReRun_Ref_084 | CC06 | Standard | 3.8 | 16975.592 | 607750 | 102521.625 | 16.558 | 1.825 | 1.72 | -5.9
[17 empty] | 20240206_HepClReRun_Ref_089 | CC06 | Standard | 3.8 | 16159.386 | 582782 | 93787.023 | 17.23 | 1.825 | 1.81 | -1
| Calibration Curve | [15 empty] | 20240206_HepClReRun_Ref_014 | CC07 | Standard | 3.8 | 16686.445 | 592203 | 76204.688 | 21.897 | 2.919 | 2.42 | -17.1
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20240206_HepClReRun_Ref_057 | CC07 | Standard | 3.8 | 19873.146 | 703580 | 83129.859 | 23.906 | 2.919 | 2.68 | -8
| 20240206_HepClReRun_Ref_011 | CC05 | 1.138 | 1.06 | 1.0713 | 0.093264 | 0.087061 | 0.29313 | -0.058656 | [7 empty] | 20240206_HepClReRun_Ref_015 | CC08 | Standard | 3.8 | 29356.393 | 1028975 | 76172.734 | 38.539 | 4.669 | 4.61 | -1.3
| 20240206_HepClReRun_Ref_035 | CC05 | 1.138 | 1.02 | [12 empty] | 20240206_HepClReRun_Ref_058 | CC08 | Standard | 3.8 | 34638.613 | 1236244 | 95656.969 | 36.211 | 4.669 | 4.3 | -7.8
| 20240206_HepClReRun_Ref_040 | CC05 | 1.138 | 1.14 | [12 empty] | 20240206_HepClReRun_Ref_016 | CC09 | Standard | 3.8 | 55292.457 | 1962972 | 79696.414 | 69.379 | 7.45 | 8.67 | 16.3
| 20240206_HepClReRun_Ref_045 | CC05 | 1.138 | 1.03 | [12 empty] | 20240206_HepClReRun_Ref_059 | CC09 | Standard | 3.8 | 64771.605 | 2311056 | 92235.273 | 70.224 | 7.45 | 8.78 | 17.8
| 20240206_HepClReRun_Ref_054 | CC05 | 1.138 | 1.03 | [12 empty] | 20240206_HepClReRun_Ref_017 | CC10 | Standard | 3.8 | 83447.602 | 2943994 | 81798.641 | 102.016 | 11.875 | 12.96 | 9.1
| 20240206_HepClReRun_Ref_078 | CC05 | 1.138 | 0.94 | [12 empty] | 20240206_HepClReRun_Ref_060 | CC10 | Standard | 3.8 | 104594.359 | 3677755 | 97453.102 | 107.328 | 11.875 | 13.66 | 15
| 20240206_HepClReRun_Ref_083 | CC05 | 1.138 | 1.25 | [12 empty] | 20240206_HepClReRun_Ref_018 | CC11 | Standard | 3.8 | 116151.75 | 4092842 | 72615.586 | 159.954 | 19.125 | 20.58 | 7.6
| 20240206_HepClReRun_Ref_088 | CC05 | 1.138 | 1.1 | [12 empty] | 20240206_HepClReRun_Ref_061 | CC11 | Standard | 3.8 | 148204.063 | 5226313 | 86978.148 | 170.392 | 19.125 | 21.96 | 14.8
[17 empty] | 20240206_HepClReRun_Ref_019 | CC12 | Standard | 3.8 | 189893.594 | 6711653 | 76026 | 249.775 | 30.5 | 32.4 | 6.2
| Blanks should be less than… | 0.14656 |  | nM | [12 empty] | 20240206_HepClReRun_Ref_062 | CC12 | Standard | 3.8 | 224008.953 | 7929918 | 85374.516 | 262.384 | 30.5 | 34.06 | 11.7
| eLOQ | 1.138 |  | nM | [12 empty] | 20240206_HepClReRun_Ref_107 | CC12 | Standard | 3.8 | 212640.453 | 6274637 | 105078.148 | 202.364 | 30.5 | 26.16 | -14.2
[17 empty] | 20240206_HepClReRun_Ref_116 | CC12 | Standard | 3.79 | 172536.125 | 5998684 | 92012.313 | 187.514 | 30.5 | 24.21 | -20.6
| QC FLAGS | [15 empty] | 20240206_HepClReRun_Ref_132 | CC12 | Standard | 3.79 | 190385.391 | 6703584 | 99651.063 | 191.052 | 30.5 | 24.67 | -19.1
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20240206_HepClReRun_Ref_141 | CC12 | Standard | 3.8 | 179850.547 | 6251511 | 87360.313 | 205.872 | 30.5 | 26.62 | -12.7
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20240206_HepClReRun_Ref_021 | CC13 | Standard | 3.8 | 188722.281 | 6669372 | 46317.648 | 407.452 | 48.75 | 53.14 | 9
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20240206_HepClReRun_Ref_037 | CC13 | Analyte | 3.8 | 203285.75 | 7171843 | 39609.715 | 513.222 | 48.75 | 67.06 | 37.6
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20240206_HepClReRun_Ref_042 | CC13 | Standard | 3.8 | 195398.391 | 7026393 | 43471.086 | 449.49 | 48.75 | 58.67 | 20.4
| Beloq eLoQ | [15 empty] | 20240206_HepClReRun_Ref_047 | CC13 | Analyte | 3.8 | 209094.203 | 7393666 | 38000.164 | 550.246 | 48.75 | 71.93 | 47.5
[17 empty] | 20240206_HepClReRun_Ref_064 | CC13 | Standard | 3.8 | 226042 | 8009873 | 50045.719 | 451.671 | 48.75 | 58.96 | 20.9
[17 empty] | 20240206_HepClReRun_Ref_080 | CC13 | Analyte | 3.8 | 215148.844 | 7610073 | 46678.375 | 460.918 | 48.75 | 60.18 | 23.4
[17 empty] | 20240206_HepClReRun_Ref_085 | CC13 | Analyte | 3.8 | 208243.859 | 7354952 | 39008.566 | 533.841 | 48.75 | 69.77 | 43.1
[17 empty] | 20240206_HepClReRun_Ref_090 | CC13 | Analyte | 3.8 | 202572.359 | 7073163 | 29047.742 | 697.377 | 48.75 | 91.29 | 87.3
[17 empty] | 20240206_HepClReRun_Ref_022 | CC14 | Standard | 3.8 | 622536.063 | 21863364 | 99196.797 | 627.577 | 78.125 | 82.1 | 5.1
[17 empty] | 20240206_HepClReRun_Ref_065 | CC14 | Standard | 3.8 | 683015 | 23939198 | 116990.578 | 583.821 | 78.125 | 76.35 | -2.3
[17 empty] | 20240206_HepClReRun_Ref_023 | CC15 | Standard | 3.8 | 797437.438 | 28023148 | 75629.461 | 1054.401 | 125 | 138.25 | 10.6
[17 empty] | 20240206_HepClReRun_Ref_066 | CC15 | Standard | 3.8 | 948660.063 | 33312734 | 93066.141 | 1019.34 | 125 | 133.64 | 6.9
[17 empty] | 20240206_HepClReRun_Ref_024 | CC16 | Standard | 3.8 | 1173109.75 | 40990896 | 82667.789 | 1419.065 | 200 | 186.23 | -6.9
[17 empty] | 20240206_HepClReRun_Ref_067 | CC16 | Standard | 3.8 | 1267215.75 | 44322560 | 96650.32 | 1311.135 | 200 | 172.03 | -14
[17 empty] | 20240206_HepClReRun_Ref_025 | CC17 | Analyte | 3.8 | 1816943 | 63451664 | 103689.242 | 1752.297 | 325.625 | 230.07 | -29.3
[17 empty] | 20240206_HepClReRun_Ref_068 | CC17 | Analyte | 3.8 | 1956425.625 | 67798808 | 93195.711 | 2099.266 | 325.625 | 275.72 | -15.3
[17 empty] | 20240206_HepClReRun_Ref_028 | CCQC A | QC | 3.8 | 6443.611 | 229278 | 87326.75 | 7.379 | 0.625 | 0.51 | -18.4
[17 empty] | 20240206_HepClReRun_Ref_071 | CCQC A | QC | 3.8 | 6273.182 | 229047 | 95330.953 | 6.58 | 0.625 | 0.41 | -35.2
[17 empty] | 20240206_HepClReRun_Ref_029 | CCQC B | QC | 3.8 | 21407.742 | 760359 | 81890.242 | 26.142 | 2.506 | 2.98 | 18.8
[17 empty] | 20240206_HepClReRun_Ref_072 | CCQC B | QC | 3.8 | 22465.502 | 797937 | 79754.047 | 28.168 | 2.506 | 3.25 | 29.5
[17 empty] | 20240206_HepClReRun_Ref_030 | CCQC C | QC | 3.8 | 60443.637 | 2140083 | 103744.516 | 58.262 | 6.269 | 7.2 | 14.9
[17 empty] | 20240206_HepClReRun_Ref_073 | CCQC C | QC | 3.8 | 61658.652 | 2182974 | 116253.453 | 53.038 | 6.269 | 6.52 | 4
[17 empty] | 20240206_HepClReRun_Ref_031 | CCQC D | QC | 3.8 | 205287.438 | 7265895 | 72114.469 | 284.669 | 25 | 36.99 | 48
[17 empty] | 20240206_HepClReRun_Ref_074 | CCQC D | QC | 3.8 | 206486.453 | 7287286 | 62696.5 | 329.343 | 25 | 42.87 | 71.5
[17 empty] | 20240206_HepClReRun_Ref_093 | Active T0 A | Analyte | 3.79 | 708850.375 | 24625148 | 159468.484 | 444.508 |  | 58.02
[17 empty] | 20240206_HepClReRun_Ref_094 | Active T0 B | Analyte | 3.79 | 620444.563 | 21533276 | 137772.844 | 450.339 |  | 58.79
[17 empty] | 20240206_HepClReRun_Ref_095 | Active T0 C | Analyte | 3.79 | 613315.25 | 21407144 | 122639.313 | 500.097 |  | 65.33
[17 empty] | 20240206_HepClReRun_Ref_096 | Active T15 A | Analyte | 3.79 | 779389.25 | 27266958 | 191937.047 | 406.065 |  | 52.96
[17 empty] | 20240206_HepClReRun_Ref_097 | Active T15 B | Analyte | 3.79 | 766846.688 | 26843304 | 270446.25 | 283.549 |  | 36.84
[17 empty] | 20240206_HepClReRun_Ref_098 | Active T15 C | Analyte | 3.79 | 671016.5 | 23370706 | 228889.453 | 293.162 |  | 38.11
[17 empty] | 20240206_HepClReRun_Ref_100 | Active T30 A | Analyte | 3.79 | 489140.531 | 17043044 | 151167.922 | 323.574 |  | 42.11
[17 empty] | 20240206_HepClReRun_Ref_101 | Active T30 B | Analyte | 3.79 | 472730.625 | 16506397 | 137365.531 | 344.141 |  | 44.81
[17 empty] | 20240206_HepClReRun_Ref_102 | Active T30 C | Analyte | 3.79 | 468518.094 | 16473934 | 188103.422 | 249.075 |  | 32.31
[17 empty] | 20240206_HepClReRun_Ref_103 | Active T60 A | Analyte | 3.79 | 359900.875 | 12575928 | 140990.125 | 255.267 |  | 33.12
[17 empty] | 20240206_HepClReRun_Ref_104 | Active T60 B | Analyte | 3.79 | 435406.719 | 15333224 | 193763.313 | 224.711 |  | 29.1
[17 empty] | 20240206_HepClReRun_Ref_105 | Active T60 C | Analyte | 3.79 | 385295.563 | 13508390 | 134827.359 | 285.77 |  | 37.13
[17 empty] | 20240206_HepClReRun_Ref_109 | Active T120 A | Analyte | 3.8 | 218175.859 | 7586293 | 66314.43 | 329.002 |  | 42.82
[17 empty] | 20240206_HepClReRun_Ref_110 | Active T120 B | Analyte | 3.79 | 168940.172 | 5890616 | 70664.453 | 239.074 |  | 30.99
[17 empty] | 20240206_HepClReRun_Ref_111 | Active T120 C | Analyte | 3.8 | 218543.875 | 7734972 | 90734.82 | 240.86 |  | 31.23
[17 empty] | 20240206_HepClReRun_Ref_112 | Active T240 A | Analyte | 3.8 | 217568.984 | 7617007 | 150100.688 | 144.949 |  | 18.61
[17 empty] | 20240206_HepClReRun_Ref_113 | Active T240 B | Analyte | 3.8 | 180226.625 | 6264853 | 121538.523 | 148.288 |  | 19.05
[17 empty] | 20240206_HepClReRun_Ref_114 | Active T240 C | Analyte | 3.79 | 167913.188 | 5850443 | 141360.891 | 118.783 |  | 15.17
[17 empty] | 20240206_HepClReRun_Ref_118 | Media T0 A | Analyte | 3.8 | 514244.094 | 17834656 | 59452.156 | 864.971 |  | 113.33
[17 empty] | 20240206_HepClReRun_Ref_119 | Media T0 B | Analyte | 3.79 | 496787.813 | 17352444 | 73160.086 | 679.042 |  | 88.87
[17 empty] | 20240206_HepClReRun_Ref_120 | Media T0 C | Analyte | 3.8 | 458327.563 | 16039667 | 81241.563 | 564.154 |  | 73.76
[17 empty] | 20240206_HepClReRun_Ref_121 | Media T15 A | Analyte | 3.79 | 698198.75 | 24233036 | 203907.672 | 342.409 |  | 44.59
[17 empty] | 20240206_HepClReRun_Ref_122 | Media T15 B | Analyte | 3.8 | 635170.063 | 22181076 | 205969.297 | 308.381 |  | 40.11
[17 empty] | 20240206_HepClReRun_Ref_123 | Media T15 C | Analyte | 3.79 | 601084.375 | 20824372 | 123378.438 | 487.188 |  | 63.63
[17 empty] | 20240206_HepClReRun_Ref_125 | Media T30 A | Analyte | 3.79 | 678412.188 | 23676540 | 105990.336 | 640.07 |  | 83.75
[17 empty] | 20240206_HepClReRun_Ref_126 | Media T30 B | Analyte | 3.79 | 638405.5 | 22125304 | 92672.016 | 688.887 |  | 90.17
[17 empty] | 20240206_HepClReRun_Ref_127 | Media T30 C | Analyte | 3.8 | 635660.25 | 22192772 | 122775.945 | 517.74 |  | 67.65
[17 empty] | 20240206_HepClReRun_Ref_128 | Media T60 A | Analyte | 3.79 | 705459.438 | 24660066 | 109065.461 | 646.822 |  | 84.63
[17 empty] | 20240206_HepClReRun_Ref_129 | Media T60 B | Analyte | 3.8 | 608528.625 | 21181664 | 111220.977 | 547.135 |  | 71.52
[17 empty] | 20240206_HepClReRun_Ref_130 | Media T60 C | Analyte | 3.79 | 686196.813 | 23924096 | 88635.961 | 774.174 |  | 101.39
[17 empty] | 20240206_HepClReRun_Ref_134 | Media T120 A | Analyte | 3.79 | 560979.875 | 19509776 | 99175.742 | 565.642 |  | 73.95
[17 empty] | 20240206_HepClReRun_Ref_135 | Media T120 B | Analyte | 3.79 | 632460.938 | 22164536 | 153366.016 | 412.387 |  | 53.79
[17 empty] | 20240206_HepClReRun_Ref_136 | Media T120 C | Analyte | 3.79 | 600658.625 | 20927208 | 156841.313 | 382.972 |  | 49.92
[17 empty] | 20240206_HepClReRun_Ref_137 | Media T240 A | Analyte | 3.79 | 502612.063 | 17573236 | 126010.648 | 398.865 |  | 52.01
[17 empty] | 20240206_HepClReRun_Ref_138 | Media T240 B | Analyte | 3.79 | 533611.875 | 18618588 | 133898.75 | 398.519 |  | 51.97
[17 empty] | 20240206_HepClReRun_Ref_139 | Media T240 C | Analyte | 3.79 | 533907.75 | 18771786 | 103058.133 | 518.065 |  | 67.7
[17 empty] | 20240206_HepClReRun_Ref_143 | Inactive T0 A | Analyte | 3.79 | 384702.844 | 13394410 | 67915.203 | 566.446 |  | 74.06
[17 empty] | 20240206_HepClReRun_Ref_144 | Inactive T0 B | Analyte | 3.79 | 398442.688 | 13978169 | 82405.945 | 483.512 |  | 63.15
[17 empty] | 20240206_HepClReRun_Ref_145 | Inactive T0 C | Analyte | 3.79 | 386388.594 | 13546750 | 57695.191 | 669.707 |  | 87.64
[17 empty] | 20240206_HepClReRun_Ref_146 | Inactive T240 A | Analyte | 3.79 | 565748.938 | 19768738 | 110553.828 | 511.741 |  | 66.86
[17 empty] | 20240206_HepClReRun_Ref_147 | Inactive T240 B | Analyte | 3.79 | 567725.813 | 19909150 | 79907.016 | 710.483 |  | 93.01
[17 empty] | 20240206_HepClReRun_Ref_148 | Inactive T240 C | Analyte | 3.79 | 596142.5 | 20803386 | 103302.727 | 577.083 |  | 75.46
[17 empty] | 20240206_HepClReRun_Ref_003 | Matrix Blank | Blank |  |  |  | 15.171
[17 empty] | 20240206_HepClReRun_Ref_004 | Crash Blank | Blank |  |  |  | 1.89
[17 empty] | 20240206_HepClReRun_Ref_001 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_002 | Blank | Blank |  |  |  | 11.192
[17 empty] | 20240206_HepClReRun_Ref_005 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_006 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_013 | Blank | Blank |  |  |  | 14.119
[17 empty] | 20240206_HepClReRun_Ref_020 | Blank | Blank |  |  |  | 11.739
[17 empty] | 20240206_HepClReRun_Ref_026 | Blank | Blank |  |  |  | 7.415
[17 empty] | 20240206_HepClReRun_Ref_027 | Blank | Blank |  |  |  | 18.043
[17 empty] | 20240206_HepClReRun_Ref_032 | Blank | Blank |  |  |  | 1.047
[17 empty] | 20240206_HepClReRun_Ref_033 | Blank | Blank |  |  |  | 12.055
[17 empty] | 20240206_HepClReRun_Ref_038 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_043 | Blank | Blank |  |  |  | 9.778
[17 empty] | 20240206_HepClReRun_Ref_048 | Blank | Blank |  |  |  | 16.982
[17 empty] | 20240206_HepClReRun_Ref_049 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_056 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_063 | Blank | Blank |  |  |  | 1.61
[17 empty] | 20240206_HepClReRun_Ref_069 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_070 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_075 | Blank | Blank |  |  |  | 9.576
[17 empty] | 20240206_HepClReRun_Ref_076 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_081 | Blank | Blank |  |  |  | 8.46
[17 empty] | 20240206_HepClReRun_Ref_086 | Blank | Blank |  |  |  | 1.97
[17 empty] | 20240206_HepClReRun_Ref_091 | Blank | Blank |  |  |  | 9.14
[17 empty] | 20240206_HepClReRun_Ref_092 | Blank | Blank |  |  |  | 0.86
[17 empty] | 20240206_HepClReRun_Ref_099 | Blank | Blank |  |  |  | 10.602
[17 empty] | 20240206_HepClReRun_Ref_106 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_108 | Blank | Blank |  |  |  | 3.136
[17 empty] | 20240206_HepClReRun_Ref_115 | Blank | Blank |  |  |  | 7.786
[17 empty] | 20240206_HepClReRun_Ref_117 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_124 | Blank | Blank |  |  |  | 2.987
[17 empty] | 20240206_HepClReRun_Ref_131 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_133 | Blank | Blank |  |  |  | 7.218
[17 empty] | 20240206_HepClReRun_Ref_140 | Blank | Blank |  |  |  | 11.756
[17 empty] | 20240206_HepClReRun_Ref_142 | Blank | Blank |  |  |  | 2.465
[17 empty] | 20240206_HepClReRun_Ref_149 | Blank | Blank |  |  |  | 6.103
[17 empty] | 20240206_HepClReRun_Ref_150 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_151 | Blank | Blank |  |  |  | 11.433
Compound 3: M8PFOA
[17 empty] | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
[17 empty] | 20240206_HepClReRun_Ref_007 | CC01 | Analyte | 3.79 | 99147.227 | 3690400 |  | 99147.227 | 100 | 118.98 | 19
[17 empty] | 20240206_HepClReRun_Ref_050 | CC01 | Analyte | 3.78 | 91640.219 | 3479469 |  | 91640.219 | 100 | 109.97 | 10
[17 empty] | 20240206_HepClReRun_Ref_008 | CC02 | Analyte | 3.79 | 110029.922 | 4115721 |  | 110029.922 | 100 | 132.04 | 32
[17 empty] | 20240206_HepClReRun_Ref_051 | CC02 | Analyte | 3.78 | 118718.242 | 4508226 |  | 118718.242 | 100 | 142.47 | 42.5
[17 empty] | 20240206_HepClReRun_Ref_009 | CC03 | Analyte | 3.79 | 97847.359 | 3658686 |  | 97847.359 | 100 | 117.42 | 17.4
[17 empty] | 20240206_HepClReRun_Ref_052 | CC03 | Analyte | 3.79 | 104661.281 | 3932631 |  | 104661.281 | 100 | 125.6 | 25.6
[17 empty] | 20240206_HepClReRun_Ref_010 | CC04 | Analyte | 3.79 | 56906.484 | 2091590 |  | 56906.484 | 100 | 68.29 | -31.7
[17 empty] | 20240206_HepClReRun_Ref_034 | CC04 | Analyte | 3.79 | 50823.348 | 1909275 |  | 50823.348 | 100 | 60.99 | -39
[17 empty] | 20240206_HepClReRun_Ref_039 | CC04 | Analyte | 3.79 | 45127.563 | 1700804 |  | 45127.563 | 100 | 54.16 | -45.8
[17 empty] | 20240206_HepClReRun_Ref_044 | CC04 | Analyte | 3.79 | 50215.895 | 1900730 |  | 50215.895 | 100 | 60.26 | -39.7
[17 empty] | 20240206_HepClReRun_Ref_053 | CC04 | Analyte | 3.79 | 51603.223 | 1962689 |  | 51603.223 | 100 | 61.93 | -38.1
[17 empty] | 20240206_HepClReRun_Ref_077 | CC04 | Analyte | 3.79 | 55185.578 | 2096160 |  | 55185.578 | 100 | 66.23 | -33.8
[17 empty] | 20240206_HepClReRun_Ref_082 | CC04 | Analyte | 3.78 | 57444.809 | 2196922 |  | 57444.809 | 100 | 68.94 | -31.1
[17 empty] | 20240206_HepClReRun_Ref_087 | CC04 | Analyte | 3.79 | 46731.051 | 1784513 |  | 46731.051 | 100 | 56.08 | -43.9
[17 empty] | 20240206_HepClReRun_Ref_011 | CC05 | Standard | 3.79 | 62519.641 | 2550542 |  | 62519.641 | 100 | 75.03 | -25
[17 empty] | 20240206_HepClReRun_Ref_035 | CC05 | Standard | 3.79 | 81373.328 | 3072050 |  | 81373.328 | 100 | 97.65 | -2.3
[17 empty] | 20240206_HepClReRun_Ref_040 | CC05 | Standard | 3.78 | 66556.109 | 2338083 |  | 66556.109 | 100 | 79.87 | -20.1
[17 empty] | 20240206_HepClReRun_Ref_045 | CC05 | Standard | 3.79 | 75406.516 | 2832145 |  | 75406.516 | 100 | 90.49 | -9.5
[17 empty] | 20240206_HepClReRun_Ref_054 | CC05 | Standard | 3.79 | 76128.5 | 2872571 |  | 76128.5 | 100 | 91.36 | -8.6
[17 empty] | 20240206_HepClReRun_Ref_078 | CC05 | Standard | 3.79 | 80799.367 | 3053570 |  | 80799.367 | 100 | 96.96 | -3
[17 empty] | 20240206_HepClReRun_Ref_083 | CC05 | Standard | 3.78 | 73697.18 | 2806167 |  | 73697.18 | 100 | 88.44 | -11.6
[17 empty] | 20240206_HepClReRun_Ref_088 | CC05 | Standard | 3.78 | 71098.961 | 2720564 |  | 71098.961 | 100 | 85.32 | -14.7
[17 empty] | 20240206_HepClReRun_Ref_012 | CC06 | Standard | 3.79 | 88497.523 | 3262129 |  | 88497.523 | 100 | 106.2 | 6.2
[17 empty] | 20240206_HepClReRun_Ref_036 | CC06 | Standard | 3.79 | 96763.203 | 3627207 |  | 96763.203 | 100 | 116.12 | 16.1
[17 empty] | 20240206_HepClReRun_Ref_041 | CC06 | Standard | 3.79 | 86041.914 | 2949560 |  | 86041.914 | 100 | 103.25 | 3.3
[17 empty] | 20240206_HepClReRun_Ref_046 | CC06 | Standard | 3.79 | 89999.438 | 3457415 |  | 89999.438 | 100 | 108 | 8
[17 empty] | 20240206_HepClReRun_Ref_055 | CC06 | Standard | 3.78 | 86843.813 | 3293688 |  | 86843.813 | 100 | 104.22 | 4.2
[17 empty] | 20240206_HepClReRun_Ref_079 | CC06 | Standard | 3.78 | 93015.078 | 3374432 |  | 93015.078 | 100 | 111.62 | 11.6
[17 empty] | 20240206_HepClReRun_Ref_084 | CC06 | Standard | 3.78 | 102521.625 | 3753753 |  | 102521.625 | 100 | 123.03 | 23
[17 empty] | 20240206_HepClReRun_Ref_089 | CC06 | Standard | 3.79 | 93787.023 | 3541642 |  | 93787.023 | 100 | 112.55 | 12.5
[17 empty] | 20240206_HepClReRun_Ref_014 | CC07 | Standard | 3.79 | 76204.688 | 2781862 |  | 76204.688 | 100 | 91.45 | -8.6
[17 empty] | 20240206_HepClReRun_Ref_057 | CC07 | Standard | 3.78 | 83129.859 | 3025645 |  | 83129.859 | 100 | 99.76 | -0.2
[17 empty] | 20240206_HepClReRun_Ref_015 | CC08 | Standard | 3.79 | 76172.734 | 2958551 |  | 76172.734 | 100 | 91.41 | -8.6
[17 empty] | 20240206_HepClReRun_Ref_058 | CC08 | Standard | 3.79 | 95656.969 | 3566732 |  | 95656.969 | 100 | 114.79 | 14.8
[17 empty] | 20240206_HepClReRun_Ref_016 | CC09 | Standard | 3.79 | 79696.414 | 3734757 |  | 79696.414 | 100 | 95.64 | -4.4
[17 empty] | 20240206_HepClReRun_Ref_059 | CC09 | Standard | 3.79 | 92235.273 | 4208975 |  | 92235.273 | 100 | 110.69 | 10.7
[17 empty] | 20240206_HepClReRun_Ref_017 | CC10 | Standard | 3.79 | 81798.641 | 3844104 |  | 81798.641 | 100 | 98.16 | -1.8
[17 empty] | 20240206_HepClReRun_Ref_060 | CC10 | Standard | 3.79 | 97453.102 | 4459191 |  | 97453.102 | 100 | 116.95 | 16.9
[17 empty] | 20240206_HepClReRun_Ref_018 | CC11 | Standard | 3.79 | 72615.586 | 2704232 |  | 72615.586 | 100 | 87.14 | -12.9
[17 empty] | 20240206_HepClReRun_Ref_061 | CC11 | Standard | 3.79 | 86978.148 | 3261943 |  | 86978.148 | 100 | 104.38 | 4.4
[17 empty] | 20240206_HepClReRun_Ref_019 | CC12 | Standard | 3.79 | 76026 | 2870378 |  | 76026 | 100 | 91.23 | -8.8
[17 empty] | 20240206_HepClReRun_Ref_062 | CC12 | Standard | 3.78 | 85374.516 | 3281335 |  | 85374.516 | 100 | 102.45 | 2.5
[17 empty] | 20240206_HepClReRun_Ref_107 | CC12 | Standard | 3.78 | 105078.148 | 3696228 |  | 105078.148 | 100 | 126.1 | 26.1
[17 empty] | 20240206_HepClReRun_Ref_116 | CC12 | Standard | 3.78 | 92012.313 | 3411119 |  | 92012.313 | 100 | 110.42 | 10.4
[17 empty] | 20240206_HepClReRun_Ref_132 | CC12 | Standard | 3.78 | 99651.063 | 3638737 |  | 99651.063 | 100 | 119.59 | 19.6
[17 empty] | 20240206_HepClReRun_Ref_141 | CC12 | Standard | 3.78 | 87360.313 | 3242247 |  | 87360.313 | 100 | 104.84 | 4.8
[17 empty] | 20240206_HepClReRun_Ref_021 | CC13 | Standard | 3.79 | 46317.648 | 1555216 |  | 46317.648 | 100 | 55.58 | -44.4
[17 empty] | 20240206_HepClReRun_Ref_037 | CC13 | Analyte | 3.79 | 39609.715 | 1478978 |  | 39609.715 | 100 | 47.53 | -52.5
[17 empty] | 20240206_HepClReRun_Ref_042 | CC13 | Standard | 3.79 | 43471.086 | 1413638 |  | 43471.086 | 100 | 52.17 | -47.8
[17 empty] | 20240206_HepClReRun_Ref_047 | CC13 | Analyte | 3.79 | 38000.164 | 1436974 |  | 38000.164 | 100 | 45.6 | -54.4
[17 empty] | 20240206_HepClReRun_Ref_064 | CC13 | Standard | 3.79 | 50045.719 | 1651115 |  | 50045.719 | 100 | 60.06 | -39.9
[17 empty] | 20240206_HepClReRun_Ref_080 | CC13 | Analyte | 3.79 | 46678.375 | 1515470 |  | 46678.375 | 100 | 56.02 | -44
[17 empty] | 20240206_HepClReRun_Ref_085 | CC13 | Analyte | 3.78 | 39008.566 | 1492758 |  | 39008.566 | 100 | 46.81 | -53.2
[17 empty] | 20240206_HepClReRun_Ref_090 | CC13 | Analyte | 3.78 | 29047.742 | 1175002 |  | 29047.742 | 100 | 34.86 | -65.1
[17 empty] | 20240206_HepClReRun_Ref_022 | CC14 | Standard | 3.79 | 99196.797 | 4421671 |  | 99196.797 | 100 | 119.04 | 19
[17 empty] | 20240206_HepClReRun_Ref_065 | CC14 | Standard | 3.78 | 116990.578 | 5138411 |  | 116990.578 | 100 | 140.39 | 40.4
[17 empty] | 20240206_HepClReRun_Ref_023 | CC15 | Standard | 3.79 | 75629.461 | 2825389 |  | 75629.461 | 100 | 90.76 | -9.2
[17 empty] | 20240206_HepClReRun_Ref_066 | CC15 | Standard | 3.78 | 93066.141 | 3368788 |  | 93066.141 | 100 | 111.68 | 11.7
[17 empty] | 20240206_HepClReRun_Ref_024 | CC16 | Standard | 3.79 | 82667.789 | 3326529 |  | 82667.789 | 100 | 99.21 | -0.8
[17 empty] | 20240206_HepClReRun_Ref_067 | CC16 | Standard | 3.79 | 96650.32 | 3640127 |  | 96650.32 | 100 | 115.99 | 16
[17 empty] | 20240206_HepClReRun_Ref_025 | CC17 | Analyte | 3.79 | 103689.242 | 3858924 |  | 103689.242 | 100 | 124.43 | 24.4
[17 empty] | 20240206_HepClReRun_Ref_068 | CC17 | Analyte | 3.78 | 93195.711 | 3550052 |  | 93195.711 | 100 | 111.84 | 11.8
[17 empty] | 20240206_HepClReRun_Ref_028 | CCQC A | QC | 3.79 | 87326.75 | 3300575 |  | 87326.75 | 100 | 104.8 | 4.8
[17 empty] | 20240206_HepClReRun_Ref_071 | CCQC A | QC | 3.79 | 95330.953 | 3615751 |  | 95330.953 | 100 | 114.4 | 14.4
[17 empty] | 20240206_HepClReRun_Ref_029 | CCQC B | QC | 3.79 | 81890.242 | 3061537 |  | 81890.242 | 100 | 98.27 | -1.7
[17 empty] | 20240206_HepClReRun_Ref_072 | CCQC B | QC | 3.79 | 79754.047 | 3041261 |  | 79754.047 | 100 | 95.71 | -4.3
[17 empty] | 20240206_HepClReRun_Ref_030 | CCQC C | QC | 3.79 | 103744.516 | 3924788 |  | 103744.516 | 100 | 124.5 | 24.5
[17 empty] | 20240206_HepClReRun_Ref_073 | CCQC C | QC | 3.79 | 116253.453 | 4423421 |  | 116253.453 | 100 | 139.51 | 39.5
[17 empty] | 20240206_HepClReRun_Ref_031 | CCQC D | QC | 3.79 | 72114.469 | 2688628 |  | 72114.469 | 100 | 86.54 | -13.5
[17 empty] | 20240206_HepClReRun_Ref_074 | CCQC D | QC | 3.79 | 62696.5 | 2378496 |  | 62696.5 | 100 | 75.24 | -24.8
[17 empty] | 20240206_HepClReRun_Ref_093 | Active T0 A | Analyte | 3.78 | 159468.484 | 6234953 |  | 159468.484 | 100 | 191.37 | 91.4
[17 empty] | 20240206_HepClReRun_Ref_094 | Active T0 B | Analyte | 3.78 | 137772.844 | 5556907 |  | 137772.844 | 100 | 165.33 | 65.3
[17 empty] | 20240206_HepClReRun_Ref_095 | Active T0 C | Analyte | 3.78 | 122639.313 | 4941604 |  | 122639.313 | 100 | 147.17 | 47.2
[17 empty] | 20240206_HepClReRun_Ref_096 | Active T15 A | Analyte | 3.78 | 191937.047 | 7737548 |  | 191937.047 | 100 | 230.33 | 130.3
[17 empty] | 20240206_HepClReRun_Ref_097 | Active T15 B | Analyte | 3.78 | 270446.25 | 10975762 |  | 270446.25 | 100 | 324.55 | 224.5
[17 empty] | 20240206_HepClReRun_Ref_098 | Active T15 C | Analyte | 3.78 | 228889.453 | 9083325 |  | 228889.453 | 100 | 274.68 | 174.7
[17 empty] | 20240206_HepClReRun_Ref_100 | Active T30 A | Analyte | 3.78 | 151167.922 | 5595652 |  | 151167.922 | 100 | 181.41 | 81.4
[17 empty] | 20240206_HepClReRun_Ref_101 | Active T30 B | Analyte | 3.78 | 137365.531 | 4915725 |  | 137365.531 | 100 | 164.85 | 64.8
[17 empty] | 20240206_HepClReRun_Ref_102 | Active T30 C | Analyte | 3.78 | 188103.422 | 6732630 |  | 188103.422 | 100 | 225.73 | 125.7
[17 empty] | 20240206_HepClReRun_Ref_103 | Active T60 A | Analyte | 3.78 | 140990.125 | 4797147 |  | 140990.125 | 100 | 169.19 | 69.2
[17 empty] | 20240206_HepClReRun_Ref_104 | Active T60 B | Analyte | 3.78 | 193763.313 | 6576127 |  | 193763.313 | 100 | 232.53 | 132.5
[17 empty] | 20240206_HepClReRun_Ref_105 | Active T60 C | Analyte | 3.78 | 134827.359 | 4495299 |  | 134827.359 | 100 | 161.8 | 61.8
[17 empty] | 20240206_HepClReRun_Ref_109 | Active T120 A | Analyte | 3.78 | 66314.43 | 2659309 |  | 66314.43 | 100 | 79.58 | -20.4
[17 empty] | 20240206_HepClReRun_Ref_110 | Active T120 B | Analyte | 3.78 | 70664.453 | 2822867 |  | 70664.453 | 100 | 84.8 | -15.2
[17 empty] | 20240206_HepClReRun_Ref_111 | Active T120 C | Analyte | 3.79 | 90734.82 | 3446046 |  | 90734.82 | 100 | 108.89 | 8.9
[17 empty] | 20240206_HepClReRun_Ref_112 | Active T240 A | Analyte | 3.78 | 150100.688 | 5841353 |  | 150100.688 | 100 | 180.13 | 80.1
[17 empty] | 20240206_HepClReRun_Ref_113 | Active T240 B | Analyte | 3.78 | 121538.523 | 4873387 |  | 121538.523 | 100 | 145.85 | 45.9
[17 empty] | 20240206_HepClReRun_Ref_114 | Active T240 C | Analyte | 3.78 | 141360.891 | 5586628 |  | 141360.891 | 100 | 169.64 | 69.6
[17 empty] | 20240206_HepClReRun_Ref_118 | Media T0 A | Analyte | 3.78 | 59452.156 | 2364865 |  | 59452.156 | 100 | 71.35 | -28.7
[17 empty] | 20240206_HepClReRun_Ref_119 | Media T0 B | Analyte | 3.78 | 73160.086 | 2981668 |  | 73160.086 | 100 | 87.8 | -12.2
[17 empty] | 20240206_HepClReRun_Ref_120 | Media T0 C | Analyte | 3.78 | 81241.563 | 3158303 |  | 81241.563 | 100 | 97.49 | -2.5
[17 empty] | 20240206_HepClReRun_Ref_121 | Media T15 A | Analyte | 3.78 | 203907.672 | 8179395 |  | 203907.672 | 100 | 244.7 | 144.7
[17 empty] | 20240206_HepClReRun_Ref_122 | Media T15 B | Analyte | 3.78 | 205969.297 | 8093554 |  | 205969.297 | 100 | 247.17 | 147.2
[17 empty] | 20240206_HepClReRun_Ref_123 | Media T15 C | Analyte | 3.78 | 123378.438 | 5218319 |  | 123378.438 | 100 | 148.06 | 48.1
[17 empty] | 20240206_HepClReRun_Ref_125 | Media T30 A | Analyte | 3.78 | 105990.336 | 4481414 |  | 105990.336 | 100 | 127.19 | 27.2
[17 empty] | 20240206_HepClReRun_Ref_126 | Media T30 B | Analyte | 3.78 | 92672.016 | 3787446 |  | 92672.016 | 100 | 111.21 | 11.2
[17 empty] | 20240206_HepClReRun_Ref_127 | Media T30 C | Analyte | 3.78 | 122775.945 | 4730254 |  | 122775.945 | 100 | 147.34 | 47.3
[17 empty] | 20240206_HepClReRun_Ref_128 | Media T60 A | Analyte | 3.78 | 109065.461 | 4490368 |  | 109065.461 | 100 | 130.88 | 30.9
[17 empty] | 20240206_HepClReRun_Ref_129 | Media T60 B | Analyte | 3.78 | 111220.977 | 4359841 |  | 111220.977 | 100 | 133.47 | 33.5
[17 empty] | 20240206_HepClReRun_Ref_130 | Media T60 C | Analyte | 3.78 | 88635.961 | 3907847 |  | 88635.961 | 100 | 106.37 | 6.4
[17 empty] | 20240206_HepClReRun_Ref_134 | Media T120 A | Analyte | 3.78 | 99175.742 | 4210011 |  | 99175.742 | 100 | 119.02 | 19
[17 empty] | 20240206_HepClReRun_Ref_135 | Media T120 B | Analyte | 3.78 | 153366.016 | 6537087 |  | 153366.016 | 100 | 184.05 | 84
[17 empty] | 20240206_HepClReRun_Ref_136 | Media T120 C | Analyte | 3.78 | 156841.313 | 6584501 |  | 156841.313 | 100 | 188.22 | 88.2
[17 empty] | 20240206_HepClReRun_Ref_137 | Media T240 A | Analyte | 3.78 | 126010.648 | 5389455 |  | 126010.648 | 100 | 151.22 | 51.2
[17 empty] | 20240206_HepClReRun_Ref_138 | Media T240 B | Analyte | 3.78 | 133898.75 | 5696388 |  | 133898.75 | 100 | 160.68 | 60.7
[17 empty] | 20240206_HepClReRun_Ref_139 | Media T240 C | Analyte | 3.78 | 103058.133 | 4536677 |  | 103058.133 | 100 | 123.67 | 23.7
[17 empty] | 20240206_HepClReRun_Ref_143 | Inactive T0 A | Analyte | 3.78 | 67915.203 | 2974119 |  | 67915.203 | 100 | 81.5 | -18.5
[17 empty] | 20240206_HepClReRun_Ref_144 | Inactive T0 B | Analyte | 3.78 | 82405.945 | 3620762 |  | 82405.945 | 100 | 98.89 | -1.1
[17 empty] | 20240206_HepClReRun_Ref_145 | Inactive T0 C | Analyte | 3.78 | 57695.191 | 2476753 |  | 57695.191 | 100 | 69.24 | -30.8
[17 empty] | 20240206_HepClReRun_Ref_146 | Inactive T240 A | Analyte | 3.78 | 110553.828 | 4797423 |  | 110553.828 | 100 | 132.67 | 32.7
[17 empty] | 20240206_HepClReRun_Ref_147 | Inactive T240 B | Analyte | 3.78 | 79907.016 | 3404979 |  | 79907.016 | 100 | 95.89 | -4.1
[17 empty] | 20240206_HepClReRun_Ref_148 | Inactive T240 C | Analyte | 3.78 | 103302.727 | 4376327 |  | 103302.727 | 100 | 123.97 | 24
[17 empty] | 20240206_HepClReRun_Ref_003 | Matrix Blank | Blank | 3.82 | 15.171 | 367 |  | 15.171 | 100 | 0.02 | -100
[17 empty] | 20240206_HepClReRun_Ref_004 | Crash Blank | Blank | 3.78 | 1.89 | 77 |  | 1.89 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_001 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_002 | Blank | Blank | 3.84 | 11.192 | 512 |  | 11.192 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_005 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_006 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_013 | Blank | Blank | 3.79 | 14.119 | 559 |  | 14.119 | 100 | 0.02 | -100
[17 empty] | 20240206_HepClReRun_Ref_020 | Blank | Blank | 3.78 | 11.739 | 323 |  | 11.739 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_026 | Blank | Blank | 3.78 | 7.415 | 304 |  | 7.415 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_027 | Blank | Blank | 3.8 | 18.043 | 493 |  | 18.043 | 100 | 0.02 | -100
[17 empty] | 20240206_HepClReRun_Ref_032 | Blank | Blank | 3.79 | 1.047 | 123 |  | 1.047 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_033 | Blank | Blank | 3.84 | 12.055 | 251 |  | 12.055 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_038 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_043 | Blank | Blank | 3.79 | 9.778 | 359 |  | 9.778 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_048 | Blank | Blank | 3.81 | 16.982 | 459 |  | 16.982 | 100 | 0.02 | -100
[17 empty] | 20240206_HepClReRun_Ref_049 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_056 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_063 | Blank | Blank | 3.83 | 1.61 | 79 |  | 1.61 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_069 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_070 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_075 | Blank | Blank | 3.78 | 9.576 | 462 |  | 9.576 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_076 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_081 | Blank | Blank | 3.84 | 8.46 | 296 |  | 8.46 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_086 | Blank | Blank | 3.78 | 1.97 | 59 |  | 1.97 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_091 | Blank | Blank | 3.8 | 9.14 | 264 |  | 9.14 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_092 | Blank | Blank | 3.78 | 0.86 | 57 |  | 0.86 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_099 | Blank | Blank | 3.78 | 10.602 | 443 |  | 10.602 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_106 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_108 | Blank | Blank | 3.82 | 3.136 | 179 |  | 3.136 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_115 | Blank | Blank | 3.82 | 7.786 | 316 |  | 7.786 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_117 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_124 | Blank | Blank | 3.82 | 2.987 | 122 |  | 2.987 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_131 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_133 | Blank | Blank | 3.78 | 7.218 | 368 |  | 7.218 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_140 | Blank | Blank | 3.79 | 11.756 | 472 |  | 11.756 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_142 | Blank | Blank | 3.83 | 2.465 | 96 |  | 2.465 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_149 | Blank | Blank | 3.84 | 6.103 | 264 |  | 6.103 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_150 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_151 | Blank | Blank | 3.83 | 11.433 | 297 |  | 11.433 | 100 | 0.01 | -100

## Propranolol
Human Active Cells | [8 empty] | Media Only | [7 empty] | Compound 2: propranolol
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV |  | Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
|  | (nM) |  | (nM) | (nM) | (nM) |  |  |  |  | (nM) |  | (nM) | (nM) | (nM) |  | 20240206_HepClReRun_Ref_007 | CC01 | Analyte | 4.33 | 5689.459 | 100067 | 99147.227 | 5.738 | 0.175
0 | Active T0 A | 43.19 | 6.9437 | 1036.56 | 1159.36 | 107.0917 | 9.237% |  | 0 | Media T0 A | 102.68 | 7.8097 | 2464.32 | 1936.48 | 495.5294 | 25.59% | 20240206_HepClReRun_Ref_050 | CC01 | Analyte | 4.33 | 6429.335 | 117375 | 91640.219 | 7.016 | 0.175
| Active T0 B | 51.39 | 7.1175 | 1233.36 | [5 empty] | Media T0 B | 77.66 | 7.5304 | 1863.84 |  |  |  | 20240206_HepClReRun_Ref_008 | CC02 | Analyte | 4.33 | 8380.075 | 161491 | 110029.922 | 7.616 | 0.281
| Active T0 C | 50.34 | 7.0969 | 1208.16 | [5 empty] | Media T0 C | 61.72 | 7.3007 | 1481.28 |  |  |  | 20240206_HepClReRun_Ref_051 | CC02 | Analyte | 4.34 | 9527.736 | 173582 | 118718.242 | 8.026 | 0.281
15 | Active T15 A | 42.55 | 6.9287 | 1021.2 | 926.88 | 88.4617 | 9.544% |  | 15 | Media T15 A | 49.25 | 7.075 | 1182 | 1270.32 | 342.0222 | 26.92% | 20240206_HepClReRun_Ref_009 | CC03 | Analyte | 4.33 | 19523.49 | 399061 | 97847.359 | 19.953 | 0.444 | 0.09 | -80.3
| Active T15 B | 35.24 | 6.7402 | 845.76 | [5 empty] | Media T15 B | 40.88 | 6.8887 | 981.12 |  |  |  | 20240206_HepClReRun_Ref_052 | CC03 | Analyte | 4.34 | 22009.053 | 436498 | 104661.281 | 21.029 | 0.444 | 0.12 | -73
| Active T15 C | 38.07 | 6.8175 | 913.68 | [5 empty] | Media T15 C | 68.66 | 7.4072 | 1647.84 |  |  |  | 20240206_HepClReRun_Ref_010 | CC04 | Analyte | 4.33 | 27509.355 | 557778 | 56906.484 | 48.341 | 0.713 | 0.95 | 33.3
30 | Active T30 A | 40.53 | 6.8801 | 972.72 | 980.16 | 24.0396 | 2.453% |  | 30 | Media T30 A | 90.8 | 7.6867 | 2179.2 | 2084.48 | 145.5561 | 6.983% | 20240206_HepClReRun_Ref_034 | CC04 | Analyte | 4.34 | 27683.043 | 577809 | 50823.348 | 54.469 | 0.713 | 1.14 | 59.4
| Active T30 B | 41.96 | 6.9148 | 1007.04 | [5 empty] | Media T30 B | 89.89 | 7.6766 | 2157.36 |  |  |  | 20240206_HepClReRun_Ref_039 | CC04 | Analyte | 4.34 | 24520.639 | 468125 | 45127.563 | 54.336 | 0.713 | 1.13 | 58.8
| Active T30 C | 40.03 | 6.8677 | 960.72 | [5 empty] | Media T30 C | 79.87 | 7.5585 | 1916.88 |  |  |  | 20240206_HepClReRun_Ref_044 | CC04 | Analyte | 4.34 | 25043.141 | 432078 | 50215.895 | 49.871 | 0.713 | 1 | 39.8
60 | Active T60 A | 37.87 | 6.8122 | 908.88 | 818.08 | 172.6653 | 21.11% |  | 60 | Media T60 A | 77.05 | 7.5225 | 1849.2 | 2241.44 | 340.7235 | 15.2% | 20240206_HepClReRun_Ref_053 | CC04 | Analyte | 4.34 | 27264.674 | 495187 | 51603.223 | 52.835 | 0.713 | 1.09 | 52.4
| Active T60 B | 25.79 | 6.428 | 618.96 | [5 empty] | Media T60 B | 102.67 | 7.8096 | 2464.08 |  |  |  | 20240206_HepClReRun_Ref_077 | CC04 | Analyte | 4.34 | 31107.773 | 647261 | 55185.578 | 56.369 | 0.713 | 1.19 | 67.5
| Active T60 C | 38.6 | 6.8313 | 926.4 | [5 empty] | Media T60 C | 100.46 | 7.7878 | 2411.04 |  |  |  | 20240206_HepClReRun_Ref_082 | CC04 | Analyte | 4.34 | 28283.32 | 513277 | 57444.809 | 49.236 | 0.713 | 0.98 | 37.1
120 | Active T120 A | 23.89 | 6.3515 | 573.36 | 685.76 | 117.004 | 17.06% |  | 120 | Media T120 A | 75.32 | 7.4998 | 1807.68 | 1453.6 | 323.3135 | 22.24% | 20240206_HepClReRun_Ref_087 | CC04 | Analyte | 4.34 | 22541.699 | 376288 | 46731.051 | 48.237 | 0.713 | 0.95 | 32.8
| Active T120 B | 33.62 | 6.6932 | 806.88 | [5 empty] | Media T120 B | 57.46 | 7.2291 | 1379.04 |  |  |  | 20240206_HepClReRun_Ref_011 | CC05 | Standard | 4.33 | 38174.992 | 757409 | 62519.641 | 61.061 | 1.138 | 1.34 | 17.5
| Active T120 C | 28.21 | 6.5177 | 677.04 | [5 empty] | Media T120 C | 48.92 | 7.0682 | 1174.08 |  |  |  | 20240206_HepClReRun_Ref_035 | CC05 | Standard | 4.33 | 40706.043 | 803446 | 81373.328 | 50.024 | 1.138 | 1 | -12
240 | Active T240 A | 10.09 | 5.4896 | 242.16 | 242.4 | 2.5286 | 1.043% |  | 240 | Media T240 A | 51.28 | 7.1154 | 1230.72 | 1394.08 | 197.8008 | 14.19% | 20240206_HepClReRun_Ref_040 | CC05 | Standard | 4.34 | 36367.973 | 702032 | 66556.109 | 54.643 | 1.138 | 1.14 | 0.3
| Active T240 B | 10 | 5.4806 | 240 | [5 empty] | Media T240 B | 55.73 | 7.1986 | 1337.52 |  |  |  | 20240206_HepClReRun_Ref_045 | CC05 | Standard | 4.34 | 36971.234 | 738878 | 75406.516 | 49.029 | 1.138 | 0.97 | -14.7
| Active T240 C | 10.21 | 5.5014 | 245.04 | [5 empty] | Media T240 C | 67.25 | 7.3865 | 1614 |  |  |  | 20240206_HepClReRun_Ref_054 | CC05 | Standard | 4.34 | 40511.148 | 799950 | 76128.5 | 53.214 | 1.138 | 1.1 | -3.5
[17 empty] | 20240206_HepClReRun_Ref_078 | CC05 | Standard | 4.34 | 39960.922 | 796852 | 80799.367 | 49.457 | 1.138 | 0.98 | -13.5
Human Inactive Cells | [16 empty] | 20240206_HepClReRun_Ref_083 | CC05 | Standard | 4.34 | 36939.109 | 709023 | 73697.18 | 50.123 | 1.138 | 1 | -11.8
Time | Sample Text | Instrument Conc. | ln (Corrected Conc) | Corrected Conc. | Average | Std Dev | CV | [9 empty] | 20240206_HepClReRun_Ref_088 | CC05 | Standard | 4.34 | 34056.609 | 664162 | 71098.961 | 47.9 | 1.138 | 0.94 | -17.7
|  | (nM) |  | (nM) | (nM) | (nM) | [10 empty] | 20240206_HepClReRun_Ref_012 | CC06 | Standard | 4.33 | 69427.516 | 1317779 | 88497.523 | 78.451 | 1.825 | 1.86 | 2.2
0 | Inactive T0 A | 74.63 | 7.4906 | 1791.12 | 1785.76 | 208.2517 | 11.66% | [9 empty] | 20240206_HepClReRun_Ref_036 | CC06 | Standard | 4.34 | 68328.227 | 1308721 | 96763.203 | 70.614 | 1.825 | 1.63 | -10.9
| Inactive T0 B | 65.62 | 7.3619 | 1574.88 | [12 empty] | 20240206_HepClReRun_Ref_041 | CC06 | Standard | 4.33 | 70868.727 | 1373456 | 86041.914 | 82.365 | 1.825 | 1.98 | 8.7
| Inactive T0 C | 82.97 | 7.5965 | 1991.28 | [12 empty] | 20240206_HepClReRun_Ref_046 | CC06 | Standard | 4.33 | 60776.066 | 1089761 | 89999.438 | 67.529 | 1.825 | 1.53 | -16
240 | Inactive T240 A | 67.14 | 7.3848 | 1611.36 | 1884 | 385.5712 | 20.47% | [9 empty] | 20240206_HepClReRun_Ref_055 | CC06 | Standard | 4.34 | 67995.023 | 1427447 | 86843.813 | 78.296 | 1.825 | 1.86 | 1.9
| Inactive T240 B | 89.86 | 7.6763 | 2156.64 | [12 empty] | 20240206_HepClReRun_Ref_079 | CC06 | Standard | 4.34 | 62070.469 | 1214556 | 93015.078 | 66.732 | 1.825 | 1.51 | -17.4
| Inactive T240 C | 75.39 | 7.5007 | 1809.36 | [12 empty] | 20240206_HepClReRun_Ref_084 | CC06 | Standard | 4.33 | 75094.086 | 1490398 | 102521.625 | 73.247 | 1.825 | 1.71 | -6.5
[17 empty] | 20240206_HepClReRun_Ref_089 | CC06 | Standard | 4.34 | 74154.109 | 1437870 | 93787.023 | 79.066 | 1.825 | 1.88 | 3.2
| Calibration Curve | [15 empty] | 20240206_HepClReRun_Ref_014 | CC07 | Standard | 4.33 | 85135.219 | 1578636 | 76204.688 | 111.719 | 2.919 | 2.87 | -1.5
| Sample Name | Cal Curve Point | Std. Conc (nM) | Actual Conc (nM) | Average Conc (nM) | StDev (nM) | CV | MDL (nM) | % Dev (avg conc) | [7 empty] | 20240206_HepClReRun_Ref_057 | CC07 | Standard | 4.34 | 102638.148 | 1975266 | 83129.859 | 123.467 | 2.919 | 3.23 | 10.7
| 20240206_HepClReRun_Ref_011 | CC05 | 1.138 | 1.34 | 1.0587 | 0.13228 | 0.12494 | 0.41576 | -0.06964 | [7 empty] | 20240206_HepClReRun_Ref_015 | CC08 | Standard | 4.33 | 121911.633 | 2212706 | 76172.734 | 160.046 | 4.669 | 4.34 | -7
| 20240206_HepClReRun_Ref_035 | CC05 | 1.138 | 1 | [12 empty] | 20240206_HepClReRun_Ref_058 | CC08 | Standard | 4.34 | 162788.688 | 2998259 | 95656.969 | 170.18 | 4.669 | 4.65 | -0.4
| 20240206_HepClReRun_Ref_040 | CC05 | 1.138 | 1.14 | [12 empty] | 20240206_HepClReRun_Ref_016 | CC09 | Standard | 4.33 | 231229.125 | 4163824 | 79696.414 | 290.137 | 7.45 | 8.29 | 11.3
| 20240206_HepClReRun_Ref_045 | CC05 | 1.138 | 0.97 | [12 empty] | 20240206_HepClReRun_Ref_059 | CC09 | Standard | 4.34 | 244243.359 | 4911281 | 92235.273 | 264.805 | 7.45 | 7.53 | 1
| 20240206_HepClReRun_Ref_054 | CC05 | 1.138 | 1.1 | [12 empty] | 20240206_HepClReRun_Ref_017 | CC10 | Standard | 4.33 | 334437.344 | 6538882 | 81798.641 | 408.854 | 11.875 | 11.9 | 0.2
| 20240206_HepClReRun_Ref_078 | CC05 | 1.138 | 0.98 | [12 empty] | 20240206_HepClReRun_Ref_060 | CC10 | Standard | 4.34 | 372020.063 | 7294902 | 97453.102 | 381.743 | 11.875 | 11.08 | -6.7
| 20240206_HepClReRun_Ref_083 | CC05 | 1.138 | 1 | [12 empty] | 20240206_HepClReRun_Ref_018 | CC11 | Standard | 4.33 | 544046.563 | 10097876 | 72615.586 | 749.215 | 19.125 | 22.24 | 16.3
| 20240206_HepClReRun_Ref_088 | CC05 | 1.138 | 0.94 | [12 empty] | 20240206_HepClReRun_Ref_061 | CC11 | Standard | 4.34 | 663656.313 | 12885014 | 86978.148 | 763.015 | 19.125 | 22.66 | 18.5
[17 empty] | 20240206_HepClReRun_Ref_019 | CC12 | Standard | 4.33 | 860768.125 | 16740276 | 76026 | 1132.202 | 30.5 | 33.87 | 11.1
| Blanks should be less than… | 0.20788 |  | nM | [12 empty] | 20240206_HepClReRun_Ref_062 | CC12 | Standard | 4.34 | 887874.5 | 16817142 | 85374.516 | 1039.976 | 30.5 | 31.07 | 1.9
| eLOQ | 1.138 |  | nM | [12 empty] | 20240206_HepClReRun_Ref_107 | CC12 | Standard | 4.33 | 1076495.125 | 20344942 | 105078.148 | 1024.471 | 30.5 | 30.6 | 0.3
[17 empty] | 20240206_HepClReRun_Ref_116 | CC12 | Standard | 4.33 | 986586.875 | 18464340 | 92012.313 | 1072.234 | 30.5 | 32.05 | 5.1
| QC FLAGS | [15 empty] | 20240206_HepClReRun_Ref_132 | CC12 | Standard | 4.33 | 994039.063 | 18405526 | 99651.063 | 997.52 | 30.5 | 29.78 | -2.3
| Calibration Curve Linearity |  |  | <7-point curve min and/or r2 < 0.98 | [12 empty] | 20240206_HepClReRun_Ref_141 | CC12 | Standard | 4.33 | 975624.938 | 19203378 | 87360.313 | 1116.783 | 30.5 | 33.41 | 9.5
| Curve Precision Checks |  |  | <75 or >125% of expected | [12 empty] | 20240206_HepClReRun_Ref_021 | CC13 | Standard | 4.33 | 870822.563 | 19199102 | 46317.648 | 1880.11 | 48.75 | 56.59 | 16.1
| Blanks |  |  | >1/2 MDL of analytes | [12 empty] | 20240206_HepClReRun_Ref_037 | CC13 | Analyte | 4.34 | 922854.563 | 17027436 | 39609.715 | 2329.869 | 48.75 | 70.26 | 44.1
| Technical Replicate Assessment |  |  | Replicates > 25% of Sample Mean | [12 empty] | 20240206_HepClReRun_Ref_042 | CC13 | Standard | 4.34 | 849091 | 16840478 | 43471.086 | 1953.232 | 48.75 | 58.82 | 20.6
| Beloq eLoQ | [15 empty] | 20240206_HepClReRun_Ref_047 | CC13 | Analyte | 4.34 | 1039245.25 | 19600076 | 38000.164 | 2734.844 | 48.75 | 82.56 | 69.4
[17 empty] | 20240206_HepClReRun_Ref_064 | CC13 | Standard | 4.34 | 946185.625 | 19434664 | 50045.719 | 1890.642 | 48.75 | 56.91 | 16.7
[17 empty] | 20240206_HepClReRun_Ref_080 | CC13 | Analyte | 4.33 | 1000756.313 | 19085440 | 46678.375 | 2143.94 | 48.75 | 64.61 | 32.5
[17 empty] | 20240206_HepClReRun_Ref_085 | CC13 | Analyte | 4.34 | 966794.625 | 18279832 | 39008.566 | 2478.416 | 48.75 | 74.77 | 53.4
[17 empty] | 20240206_HepClReRun_Ref_090 | CC13 | Analyte | 4.34 | 1063102.875 | 19946494 | 29047.742 | 3659.847 | 48.75 | 110.66 | 127
[17 empty] | 20240206_HepClReRun_Ref_022 | CC14 | Standard | 4.33 | 2365647 | 42121644 | 99196.797 | 2384.802 | 78.125 | 71.93 | -7.9
[17 empty] | 20240206_HepClReRun_Ref_065 | CC14 | Standard | 4.33 | 2664637.25 | 50079728 | 116990.578 | 2277.651 | 78.125 | 68.67 | -12.1
[17 empty] | 20240206_HepClReRun_Ref_023 | CC15 | Standard | 4.33 | 3551716.25 | 66769992 | 75629.461 | 4696.207 | 125 | 142.14 | 13.7
[17 empty] | 20240206_HepClReRun_Ref_066 | CC15 | Standard | 4.33 | 3385405.25 | 55499456 | 93066.141 | 3637.634 | 125 | 109.98 | -12
[17 empty] | 20240206_HepClReRun_Ref_024 | CC16 | Standard | 4.33 | 5274319 | 89664752 | 82667.789 | 6380.138 | 200 | 193.29 | -3.4
[17 empty] | 20240206_HepClReRun_Ref_067 | CC16 | Standard | 4.33 | 5772579 | 110625960 | 96650.32 | 5972.643 | 200 | 180.91 | -9.5
[17 empty] | 20240206_HepClReRun_Ref_025 | CC17 | Analyte | 4.33 | 8462794 | 160269728 | 103689.242 | 8161.69 | 325.625 | 247.41 | -24
[17 empty] | 20240206_HepClReRun_Ref_068 | CC17 | Analyte | 4.34 | 9179669 | 166081136 | 93195.711 | 9849.884 | 325.625 | 298.7 | -8.3
[17 empty] | 20240206_HepClReRun_Ref_028 | CCQC A | QC | 4.33 | 31258.584 | 606356 | 87326.75 | 35.795 | 0.625 | 0.57 | -9
[17 empty] | 20240206_HepClReRun_Ref_071 | CCQC A | QC | 4.34 | 32431.244 | 613701 | 95330.953 | 34.02 | 0.625 | 0.51 | -17.7
[17 empty] | 20240206_HepClReRun_Ref_029 | CCQC B | QC | 4.33 | 89345.578 | 1917429 | 81890.242 | 109.104 | 2.506 | 2.8 | 11.5
[17 empty] | 20240206_HepClReRun_Ref_072 | CCQC B | QC | 4.34 | 89657.414 | 1547955 | 79754.047 | 112.417 | 2.506 | 2.9 | 15.6
[17 empty] | 20240206_HepClReRun_Ref_030 | CCQC C | QC | 4.33 | 238971.328 | 4634862 | 103744.516 | 230.346 | 6.269 | 6.48 | 3.3
[17 empty] | 20240206_HepClReRun_Ref_073 | CCQC C | QC | 4.34 | 241690 | 4742805 | 116253.453 | 207.899 | 6.269 | 5.8 | -7.5
[17 empty] | 20240206_HepClReRun_Ref_031 | CCQC D | QC | 4.34 | 913307 | 16551990 | 72114.469 | 1266.468 | 25 | 37.95 | 51.8
[17 empty] | 20240206_HepClReRun_Ref_074 | CCQC D | QC | 4.34 | 1043704.25 | 20803620 | 62696.5 | 1664.693 | 25 | 50.05 | 100.2
[17 empty] | 20240206_HepClReRun_Ref_093 | Active T0 A | Analyte | 4.33 | 2294531.25 | 42666740 | 159468.484 | 1438.862 |  | 43.19
[17 empty] | 20240206_HepClReRun_Ref_094 | Active T0 B | Analyte | 4.33 | 2354053.25 | 45344624 | 137772.844 | 1708.648 |  | 51.39
[17 empty] | 20240206_HepClReRun_Ref_095 | Active T0 C | Analyte | 4.33 | 2053148 | 40953388 | 122639.313 | 1674.135 |  | 50.34
[17 empty] | 20240206_HepClReRun_Ref_096 | Active T15 A | Analyte | 4.33 | 2721176.25 | 52004224 | 191937.047 | 1417.744 |  | 42.55
[17 empty] | 20240206_HepClReRun_Ref_097 | Active T15 B | Analyte | 4.33 | 3183320 | 60800472 | 270446.25 | 1177.062 |  | 35.24
[17 empty] | 20240206_HepClReRun_Ref_098 | Active T15 C | Analyte | 4.32 | 2907235 | 55977992 | 228889.453 | 1270.148 |  | 38.07
[17 empty] | 20240206_HepClReRun_Ref_100 | Active T30 A | Analyte | 4.33 | 2042914 | 38590028 | 151167.922 | 1351.42 |  | 40.53
[17 empty] | 20240206_HepClReRun_Ref_101 | Active T30 B | Analyte | 4.33 | 1920954.625 | 33420922 | 137365.531 | 1398.426 |  | 41.96
[17 empty] | 20240206_HepClReRun_Ref_102 | Active T30 C | Analyte | 4.32 | 2511156.5 | 49497012 | 188103.422 | 1334.987 |  | 40.03
[17 empty] | 20240206_HepClReRun_Ref_103 | Active T60 A | Analyte | 4.32 | 1781806 | 33449774 | 140990.125 | 1263.781 |  | 37.87
[17 empty] | 20240206_HepClReRun_Ref_104 | Active T60 B | Analyte | 4.32 | 1678213.625 | 32093362 | 193763.313 | 866.115 |  | 25.79
[17 empty] | 20240206_HepClReRun_Ref_105 | Active T60 C | Analyte | 4.32 | 1736233.375 | 36583268 | 134827.359 | 1287.746 |  | 38.6
[17 empty] | 20240206_HepClReRun_Ref_109 | Active T120 A | Analyte | 4.33 | 532864.438 | 10242889 | 66314.43 | 803.542 |  | 23.89
[17 empty] | 20240206_HepClReRun_Ref_110 | Active T120 B | Analyte | 4.33 | 794152 | 14965115 | 70664.453 | 1123.835 |  | 33.62
[17 empty] | 20240206_HepClReRun_Ref_111 | Active T120 C | Analyte | 4.33 | 858142.438 | 17149398 | 90734.82 | 945.77 |  | 28.21
[17 empty] | 20240206_HepClReRun_Ref_112 | Active T240 A | Analyte | 4.33 | 524098.875 | 9971786 | 150100.688 | 349.165 |  | 10.09
[17 empty] | 20240206_HepClReRun_Ref_113 | Active T240 B | Analyte | 4.33 | 420662.5 | 8655304 | 121538.523 | 346.115 |  | 10
[17 empty] | 20240206_HepClReRun_Ref_114 | Active T240 C | Analyte | 4.33 | 499231.844 | 9808518 | 141360.891 | 353.161 |  | 10.21
[17 empty] | 20240206_HepClReRun_Ref_118 | Media T0 A | Analyte | 4.33 | 2019758.875 | 37338576 | 59452.156 | 3397.284 |  | 102.68
[17 empty] | 20240206_HepClReRun_Ref_119 | Media T0 B | Analyte | 4.33 | 1882942.125 | 36278996 | 73160.086 | 2573.729 |  | 77.66
[17 empty] | 20240206_HepClReRun_Ref_120 | Media T0 C | Analyte | 4.33 | 1664458.75 | 30950594 | 81241.563 | 2048.777 |  | 61.72
[17 empty] | 20240206_HepClReRun_Ref_121 | Media T15 A | Analyte | 4.32 | 3340735.25 | 58716056 | 203907.672 | 1638.357 |  | 49.25
[17 empty] | 20240206_HepClReRun_Ref_122 | Media T15 B | Analyte | 4.32 | 2807132.75 | 53510668 | 205969.297 | 1362.889 |  | 40.88
[17 empty] | 20240206_HepClReRun_Ref_123 | Media T15 C | Analyte | 4.33 | 2809799 | 51944668 | 123378.438 | 2277.383 |  | 68.66
[17 empty] | 20240206_HepClReRun_Ref_125 | Media T30 A | Analyte | 4.32 | 3186197.25 | 61430992 | 105990.336 | 3006.121 |  | 90.8
[17 empty] | 20240206_HepClReRun_Ref_126 | Media T30 B | Analyte | 4.33 | 2758034.25 | 51965436 | 92672.016 | 2976.124 |  | 89.89
[17 empty] | 20240206_HepClReRun_Ref_127 | Media T30 C | Analyte | 4.33 | 3248939 | 62116900 | 122775.945 | 2646.234 |  | 79.87
[17 empty] | 20240206_HepClReRun_Ref_128 | Media T60 A | Analyte | 4.32 | 2785065.25 | 52642948 | 109065.461 | 2553.572 |  | 77.05
[17 empty] | 20240206_HepClReRun_Ref_129 | Media T60 B | Analyte | 4.33 | 3777946.5 | 70176496 | 111220.977 | 3396.793 |  | 102.67
[17 empty] | 20240206_HepClReRun_Ref_130 | Media T60 C | Analyte | 4.33 | 2946401 | 55806236 | 88635.961 | 3324.16 |  | 100.46
[17 empty] | 20240206_HepClReRun_Ref_134 | Media T120 A | Analyte | 4.33 | 2475880.25 | 46387352 | 99175.742 | 2496.458 |  | 75.32
[17 empty] | 20240206_HepClReRun_Ref_135 | Media T120 B | Analyte | 4.33 | 2926991.5 | 53916236 | 153366.016 | 1908.501 |  | 57.46
[17 empty] | 20240206_HepClReRun_Ref_136 | Media T120 C | Analyte | 4.33 | 2552598.75 | 46398920 | 156841.313 | 1627.504 |  | 48.92
[17 empty] | 20240206_HepClReRun_Ref_137 | Media T240 A | Analyte | 4.32 | 2148663.25 | 40641980 | 126010.648 | 1705.144 |  | 51.28
[17 empty] | 20240206_HepClReRun_Ref_138 | Media T240 B | Analyte | 4.32 | 2479209 | 45458668 | 133898.75 | 1851.555 |  | 55.73
[17 empty] | 20240206_HepClReRun_Ref_139 | Media T240 C | Analyte | 4.32 | 2299225.5 | 45059124 | 103058.133 | 2230.999 |  | 67.25
[17 empty] | 20240206_HepClReRun_Ref_143 | Inactive T0 A | Analyte | 4.33 | 1680210.625 | 35287136 | 67915.203 | 2473.983 |  | 74.63
[17 empty] | 20240206_HepClReRun_Ref_144 | Inactive T0 B | Analyte | 4.33 | 1794048.625 | 34997684 | 82405.945 | 2177.086 |  | 65.62
[17 empty] | 20240206_HepClReRun_Ref_145 | Inactive T0 C | Analyte | 4.32 | 1585744.25 | 29815876 | 57695.191 | 2748.486 |  | 82.97
[17 empty] | 20240206_HepClReRun_Ref_146 | Inactive T240 A | Analyte | 4.33 | 2462339.25 | 45285364 | 110553.828 | 2227.276 |  | 67.14
[17 empty] | 20240206_HepClReRun_Ref_147 | Inactive T240 B | Analyte | 4.33 | 2377318.25 | 44228896 | 79907.016 | 2975.106 |  | 89.86
[17 empty] | 20240206_HepClReRun_Ref_148 | Inactive T240 C | Analyte | 4.32 | 2581464.75 | 45727672 | 103302.727 | 2498.932 |  | 75.39
[17 empty] | 20240206_HepClReRun_Ref_003 | Matrix Blank | Blank |  |  |  | 15.171
[17 empty] | 20240206_HepClReRun_Ref_004 | Crash Blank | Blank |  |  |  | 1.89
[17 empty] | 20240206_HepClReRun_Ref_001 | Blank | Blank | 4.34 | 29.845 | 1236
[17 empty] | 20240206_HepClReRun_Ref_002 | Blank | Blank |  |  |  | 11.192
[17 empty] | 20240206_HepClReRun_Ref_005 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_006 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_013 | Blank | Blank | 4.33 | 10.779 | 831 | 14.119 | 76.344 |  | 1.8
[17 empty] | 20240206_HepClReRun_Ref_020 | Blank | Blank | 4.34 | 1514.958 | 22611 | 11.739 | 12905.341 |  | 391.51
[17 empty] | 20240206_HepClReRun_Ref_026 | Blank | Blank | 4.33 | 17999.15 | 280841 | 7.415 | 242739.717 |  | 7373.29
[17 empty] | 20240206_HepClReRun_Ref_027 | Blank | Blank | 4.33 | 3748.776 | 50957 | 18.043 | 20776.9 |  | 630.63
[17 empty] | 20240206_HepClReRun_Ref_032 | Blank | Blank | 4.33 | 1958.397 | 36961 | 1.047 | 187048.424 |  | 5681.53
[17 empty] | 20240206_HepClReRun_Ref_033 | Blank | Blank | 4.34 | 286.214 | 5789 | 12.055 | 2374.235 |  | 71.6
[17 empty] | 20240206_HepClReRun_Ref_038 | Blank | Blank | 4.33 | 1089.875 | 24160
[17 empty] | 20240206_HepClReRun_Ref_043 | Blank | Blank | 4.33 | 1211.203 | 25076 | 9.778 | 12387.022 |  | 375.77
[17 empty] | 20240206_HepClReRun_Ref_048 | Blank | Blank | 4.33 | 1172.217 | 19600 | 16.982 | 6902.703 |  | 209.17
[17 empty] | 20240206_HepClReRun_Ref_049 | Blank | Blank | 4.33 | 231.257 | 5935
[17 empty] | 20240206_HepClReRun_Ref_056 | Blank | Blank
[17 empty] | 20240206_HepClReRun_Ref_063 | Blank | Blank | 4.33 | 1801.626 | 29332 | 1.61 | 111902.236 |  | 3398.78
[17 empty] | 20240206_HepClReRun_Ref_069 | Blank | Blank | 4.33 | 21903.471 | 355734
[17 empty] | 20240206_HepClReRun_Ref_070 | Blank | Blank | 4.33 | 4349.25 | 71317
[17 empty] | 20240206_HepClReRun_Ref_075 | Blank | Blank | 4.33 | 1827.312 | 31153 | 9.576 | 19082.206 |  | 579.15
[17 empty] | 20240206_HepClReRun_Ref_076 | Blank | Blank | 4.34 | 536.954 | 11782
[17 empty] | 20240206_HepClReRun_Ref_081 | Blank | Blank | 4.33 | 1757.605 | 35551 | 8.46 | 20775.473 |  | 630.59
[17 empty] | 20240206_HepClReRun_Ref_086 | Blank | Blank | 4.34 | 2099.177 | 30580 | 1.97 | 106557.208 |  | 3236.42
[17 empty] | 20240206_HepClReRun_Ref_091 | Blank | Blank | 4.32 | 1580.214 | 27385 | 9.14 | 17288.993 |  | 524.68
[17 empty] | 20240206_HepClReRun_Ref_092 | Blank | Blank | 4.35 | 354.868 | 7123 | 0.86 | 41263.721 |  | 1252.97
[17 empty] | 20240206_HepClReRun_Ref_099 | Blank | Blank | 4.32 | 5665.304 | 88952 | 10.602 | 53436.182 |  | 1622.74
[17 empty] | 20240206_HepClReRun_Ref_106 | Blank | Blank | 4.32 | 3944.261 | 65031
[17 empty] | 20240206_HepClReRun_Ref_108 | Blank | Blank | 4.32 | 2588.102 | 39532 | 3.136 | 82528.763 |  | 2506.49
[17 empty] | 20240206_HepClReRun_Ref_115 | Blank | Blank | 4.32 | 1192.125 | 20112 | 7.786 | 15311.135 |  | 464.59
[17 empty] | 20240206_HepClReRun_Ref_117 | Blank | Blank | 4.32 | 1896.822 | 26264
[17 empty] | 20240206_HepClReRun_Ref_124 | Blank | Blank | 4.32 | 5829.389 | 99283 | 2.987 | 195158.654 |  | 5927.9
[17 empty] | 20240206_HepClReRun_Ref_131 | Blank | Blank | 4.32 | 7125.924 | 108026
[17 empty] | 20240206_HepClReRun_Ref_133 | Blank | Blank | 4.32 | 2937.011 | 42697 | 7.218 | 40690.094 |  | 1235.54
[17 empty] | 20240206_HepClReRun_Ref_140 | Blank | Blank | 4.32 | 4341.378 | 63516 | 11.756 | 36929.04 |  | 1121.29
[17 empty] | 20240206_HepClReRun_Ref_142 | Blank | Blank | 4.33 | 2837.69 | 36534 | 2.465 | 115119.27 |  | 3496.51
[17 empty] | 20240206_HepClReRun_Ref_149 | Blank | Blank | 4.32 | 6456.986 | 100137 | 6.103 | 105800.197 |  | 3213.42
[17 empty] | 20240206_HepClReRun_Ref_150 | Blank | Blank | 4.32 | 1637.52 | 28260
[17 empty] | 20240206_HepClReRun_Ref_151 | Blank | Blank | 4.34 | 737.453 | 11797 | 11.433 | 6450.214 |  | 195.42
Compound 3: M8PFOA
[17 empty] | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev
[17 empty] | 20240206_HepClReRun_Ref_007 | CC01 | Analyte | 3.79 | 99147.227 | 3690400 |  | 99147.227 | 100 | 118.98 | 19
[17 empty] | 20240206_HepClReRun_Ref_050 | CC01 | Analyte | 3.78 | 91640.219 | 3479469 |  | 91640.219 | 100 | 109.97 | 10
[17 empty] | 20240206_HepClReRun_Ref_008 | CC02 | Analyte | 3.79 | 110029.922 | 4115721 |  | 110029.922 | 100 | 132.04 | 32
[17 empty] | 20240206_HepClReRun_Ref_051 | CC02 | Analyte | 3.78 | 118718.242 | 4508226 |  | 118718.242 | 100 | 142.47 | 42.5
[17 empty] | 20240206_HepClReRun_Ref_009 | CC03 | Analyte | 3.79 | 97847.359 | 3658686 |  | 97847.359 | 100 | 117.42 | 17.4
[17 empty] | 20240206_HepClReRun_Ref_052 | CC03 | Analyte | 3.79 | 104661.281 | 3932631 |  | 104661.281 | 100 | 125.6 | 25.6
[17 empty] | 20240206_HepClReRun_Ref_010 | CC04 | Analyte | 3.79 | 56906.484 | 2091590 |  | 56906.484 | 100 | 68.29 | -31.7
[17 empty] | 20240206_HepClReRun_Ref_034 | CC04 | Analyte | 3.79 | 50823.348 | 1909275 |  | 50823.348 | 100 | 60.99 | -39
[17 empty] | 20240206_HepClReRun_Ref_039 | CC04 | Analyte | 3.79 | 45127.563 | 1700804 |  | 45127.563 | 100 | 54.16 | -45.8
[17 empty] | 20240206_HepClReRun_Ref_044 | CC04 | Analyte | 3.79 | 50215.895 | 1900730 |  | 50215.895 | 100 | 60.26 | -39.7
[17 empty] | 20240206_HepClReRun_Ref_053 | CC04 | Analyte | 3.79 | 51603.223 | 1962689 |  | 51603.223 | 100 | 61.93 | -38.1
[17 empty] | 20240206_HepClReRun_Ref_077 | CC04 | Analyte | 3.79 | 55185.578 | 2096160 |  | 55185.578 | 100 | 66.23 | -33.8
[17 empty] | 20240206_HepClReRun_Ref_082 | CC04 | Analyte | 3.78 | 57444.809 | 2196922 |  | 57444.809 | 100 | 68.94 | -31.1
[17 empty] | 20240206_HepClReRun_Ref_087 | CC04 | Analyte | 3.79 | 46731.051 | 1784513 |  | 46731.051 | 100 | 56.08 | -43.9
[17 empty] | 20240206_HepClReRun_Ref_011 | CC05 | Standard | 3.79 | 62519.641 | 2550542 |  | 62519.641 | 100 | 75.03 | -25
[17 empty] | 20240206_HepClReRun_Ref_035 | CC05 | Standard | 3.79 | 81373.328 | 3072050 |  | 81373.328 | 100 | 97.65 | -2.3
[17 empty] | 20240206_HepClReRun_Ref_040 | CC05 | Standard | 3.78 | 66556.109 | 2338083 |  | 66556.109 | 100 | 79.87 | -20.1
[17 empty] | 20240206_HepClReRun_Ref_045 | CC05 | Standard | 3.79 | 75406.516 | 2832145 |  | 75406.516 | 100 | 90.49 | -9.5
[17 empty] | 20240206_HepClReRun_Ref_054 | CC05 | Standard | 3.79 | 76128.5 | 2872571 |  | 76128.5 | 100 | 91.36 | -8.6
[17 empty] | 20240206_HepClReRun_Ref_078 | CC05 | Standard | 3.79 | 80799.367 | 3053570 |  | 80799.367 | 100 | 96.96 | -3
[17 empty] | 20240206_HepClReRun_Ref_083 | CC05 | Standard | 3.78 | 73697.18 | 2806167 |  | 73697.18 | 100 | 88.44 | -11.6
[17 empty] | 20240206_HepClReRun_Ref_088 | CC05 | Standard | 3.78 | 71098.961 | 2720564 |  | 71098.961 | 100 | 85.32 | -14.7
[17 empty] | 20240206_HepClReRun_Ref_012 | CC06 | Standard | 3.79 | 88497.523 | 3262129 |  | 88497.523 | 100 | 106.2 | 6.2
[17 empty] | 20240206_HepClReRun_Ref_036 | CC06 | Standard | 3.79 | 96763.203 | 3627207 |  | 96763.203 | 100 | 116.12 | 16.1
[17 empty] | 20240206_HepClReRun_Ref_041 | CC06 | Standard | 3.79 | 86041.914 | 2949560 |  | 86041.914 | 100 | 103.25 | 3.3
[17 empty] | 20240206_HepClReRun_Ref_046 | CC06 | Standard | 3.79 | 89999.438 | 3457415 |  | 89999.438 | 100 | 108 | 8
[17 empty] | 20240206_HepClReRun_Ref_055 | CC06 | Standard | 3.78 | 86843.813 | 3293688 |  | 86843.813 | 100 | 104.22 | 4.2
[17 empty] | 20240206_HepClReRun_Ref_079 | CC06 | Standard | 3.78 | 93015.078 | 3374432 |  | 93015.078 | 100 | 111.62 | 11.6
[17 empty] | 20240206_HepClReRun_Ref_084 | CC06 | Standard | 3.78 | 102521.625 | 3753753 |  | 102521.625 | 100 | 123.03 | 23
[17 empty] | 20240206_HepClReRun_Ref_089 | CC06 | Standard | 3.79 | 93787.023 | 3541642 |  | 93787.023 | 100 | 112.55 | 12.5
[17 empty] | 20240206_HepClReRun_Ref_014 | CC07 | Standard | 3.79 | 76204.688 | 2781862 |  | 76204.688 | 100 | 91.45 | -8.6
[17 empty] | 20240206_HepClReRun_Ref_057 | CC07 | Standard | 3.78 | 83129.859 | 3025645 |  | 83129.859 | 100 | 99.76 | -0.2
[17 empty] | 20240206_HepClReRun_Ref_015 | CC08 | Standard | 3.79 | 76172.734 | 2958551 |  | 76172.734 | 100 | 91.41 | -8.6
[17 empty] | 20240206_HepClReRun_Ref_058 | CC08 | Standard | 3.79 | 95656.969 | 3566732 |  | 95656.969 | 100 | 114.79 | 14.8
[17 empty] | 20240206_HepClReRun_Ref_016 | CC09 | Standard | 3.79 | 79696.414 | 3734757 |  | 79696.414 | 100 | 95.64 | -4.4
[17 empty] | 20240206_HepClReRun_Ref_059 | CC09 | Standard | 3.79 | 92235.273 | 4208975 |  | 92235.273 | 100 | 110.69 | 10.7
[17 empty] | 20240206_HepClReRun_Ref_017 | CC10 | Standard | 3.79 | 81798.641 | 3844104 |  | 81798.641 | 100 | 98.16 | -1.8
[17 empty] | 20240206_HepClReRun_Ref_060 | CC10 | Standard | 3.79 | 97453.102 | 4459191 |  | 97453.102 | 100 | 116.95 | 16.9
[17 empty] | 20240206_HepClReRun_Ref_018 | CC11 | Standard | 3.79 | 72615.586 | 2704232 |  | 72615.586 | 100 | 87.14 | -12.9
[17 empty] | 20240206_HepClReRun_Ref_061 | CC11 | Standard | 3.79 | 86978.148 | 3261943 |  | 86978.148 | 100 | 104.38 | 4.4
[17 empty] | 20240206_HepClReRun_Ref_019 | CC12 | Standard | 3.79 | 76026 | 2870378 |  | 76026 | 100 | 91.23 | -8.8
[17 empty] | 20240206_HepClReRun_Ref_062 | CC12 | Standard | 3.78 | 85374.516 | 3281335 |  | 85374.516 | 100 | 102.45 | 2.5
[17 empty] | 20240206_HepClReRun_Ref_107 | CC12 | Standard | 3.78 | 105078.148 | 3696228 |  | 105078.148 | 100 | 126.1 | 26.1
[17 empty] | 20240206_HepClReRun_Ref_116 | CC12 | Standard | 3.78 | 92012.313 | 3411119 |  | 92012.313 | 100 | 110.42 | 10.4
[17 empty] | 20240206_HepClReRun_Ref_132 | CC12 | Standard | 3.78 | 99651.063 | 3638737 |  | 99651.063 | 100 | 119.59 | 19.6
[17 empty] | 20240206_HepClReRun_Ref_141 | CC12 | Standard | 3.78 | 87360.313 | 3242247 |  | 87360.313 | 100 | 104.84 | 4.8
[17 empty] | 20240206_HepClReRun_Ref_021 | CC13 | Standard | 3.79 | 46317.648 | 1555216 |  | 46317.648 | 100 | 55.58 | -44.4
[17 empty] | 20240206_HepClReRun_Ref_037 | CC13 | Analyte | 3.79 | 39609.715 | 1478978 |  | 39609.715 | 100 | 47.53 | -52.5
[17 empty] | 20240206_HepClReRun_Ref_042 | CC13 | Standard | 3.79 | 43471.086 | 1413638 |  | 43471.086 | 100 | 52.17 | -47.8
[17 empty] | 20240206_HepClReRun_Ref_047 | CC13 | Analyte | 3.79 | 38000.164 | 1436974 |  | 38000.164 | 100 | 45.6 | -54.4
[17 empty] | 20240206_HepClReRun_Ref_064 | CC13 | Standard | 3.79 | 50045.719 | 1651115 |  | 50045.719 | 100 | 60.06 | -39.9
[17 empty] | 20240206_HepClReRun_Ref_080 | CC13 | Analyte | 3.79 | 46678.375 | 1515470 |  | 46678.375 | 100 | 56.02 | -44
[17 empty] | 20240206_HepClReRun_Ref_085 | CC13 | Analyte | 3.78 | 39008.566 | 1492758 |  | 39008.566 | 100 | 46.81 | -53.2
[17 empty] | 20240206_HepClReRun_Ref_090 | CC13 | Analyte | 3.78 | 29047.742 | 1175002 |  | 29047.742 | 100 | 34.86 | -65.1
[17 empty] | 20240206_HepClReRun_Ref_022 | CC14 | Standard | 3.79 | 99196.797 | 4421671 |  | 99196.797 | 100 | 119.04 | 19
[17 empty] | 20240206_HepClReRun_Ref_065 | CC14 | Standard | 3.78 | 116990.578 | 5138411 |  | 116990.578 | 100 | 140.39 | 40.4
[17 empty] | 20240206_HepClReRun_Ref_023 | CC15 | Standard | 3.79 | 75629.461 | 2825389 |  | 75629.461 | 100 | 90.76 | -9.2
[17 empty] | 20240206_HepClReRun_Ref_066 | CC15 | Standard | 3.78 | 93066.141 | 3368788 |  | 93066.141 | 100 | 111.68 | 11.7
[17 empty] | 20240206_HepClReRun_Ref_024 | CC16 | Standard | 3.79 | 82667.789 | 3326529 |  | 82667.789 | 100 | 99.21 | -0.8
[17 empty] | 20240206_HepClReRun_Ref_067 | CC16 | Standard | 3.79 | 96650.32 | 3640127 |  | 96650.32 | 100 | 115.99 | 16
[17 empty] | 20240206_HepClReRun_Ref_025 | CC17 | Analyte | 3.79 | 103689.242 | 3858924 |  | 103689.242 | 100 | 124.43 | 24.4
[17 empty] | 20240206_HepClReRun_Ref_068 | CC17 | Analyte | 3.78 | 93195.711 | 3550052 |  | 93195.711 | 100 | 111.84 | 11.8
[17 empty] | 20240206_HepClReRun_Ref_028 | CCQC A | QC | 3.79 | 87326.75 | 3300575 |  | 87326.75 | 100 | 104.8 | 4.8
[17 empty] | 20240206_HepClReRun_Ref_071 | CCQC A | QC | 3.79 | 95330.953 | 3615751 |  | 95330.953 | 100 | 114.4 | 14.4
[17 empty] | 20240206_HepClReRun_Ref_029 | CCQC B | QC | 3.79 | 81890.242 | 3061537 |  | 81890.242 | 100 | 98.27 | -1.7
[17 empty] | 20240206_HepClReRun_Ref_072 | CCQC B | QC | 3.79 | 79754.047 | 3041261 |  | 79754.047 | 100 | 95.71 | -4.3
[17 empty] | 20240206_HepClReRun_Ref_030 | CCQC C | QC | 3.79 | 103744.516 | 3924788 |  | 103744.516 | 100 | 124.5 | 24.5
[17 empty] | 20240206_HepClReRun_Ref_073 | CCQC C | QC | 3.79 | 116253.453 | 4423421 |  | 116253.453 | 100 | 139.51 | 39.5
[17 empty] | 20240206_HepClReRun_Ref_031 | CCQC D | QC | 3.79 | 72114.469 | 2688628 |  | 72114.469 | 100 | 86.54 | -13.5
[17 empty] | 20240206_HepClReRun_Ref_074 | CCQC D | QC | 3.79 | 62696.5 | 2378496 |  | 62696.5 | 100 | 75.24 | -24.8
[17 empty] | 20240206_HepClReRun_Ref_093 | Active T0 A | Analyte | 3.78 | 159468.484 | 6234953 |  | 159468.484 | 100 | 191.37 | 91.4
[17 empty] | 20240206_HepClReRun_Ref_094 | Active T0 B | Analyte | 3.78 | 137772.844 | 5556907 |  | 137772.844 | 100 | 165.33 | 65.3
[17 empty] | 20240206_HepClReRun_Ref_095 | Active T0 C | Analyte | 3.78 | 122639.313 | 4941604 |  | 122639.313 | 100 | 147.17 | 47.2
[17 empty] | 20240206_HepClReRun_Ref_096 | Active T15 A | Analyte | 3.78 | 191937.047 | 7737548 |  | 191937.047 | 100 | 230.33 | 130.3
[17 empty] | 20240206_HepClReRun_Ref_097 | Active T15 B | Analyte | 3.78 | 270446.25 | 10975762 |  | 270446.25 | 100 | 324.55 | 224.5
[17 empty] | 20240206_HepClReRun_Ref_098 | Active T15 C | Analyte | 3.78 | 228889.453 | 9083325 |  | 228889.453 | 100 | 274.68 | 174.7
[17 empty] | 20240206_HepClReRun_Ref_100 | Active T30 A | Analyte | 3.78 | 151167.922 | 5595652 |  | 151167.922 | 100 | 181.41 | 81.4
[17 empty] | 20240206_HepClReRun_Ref_101 | Active T30 B | Analyte | 3.78 | 137365.531 | 4915725 |  | 137365.531 | 100 | 164.85 | 64.8
[17 empty] | 20240206_HepClReRun_Ref_102 | Active T30 C | Analyte | 3.78 | 188103.422 | 6732630 |  | 188103.422 | 100 | 225.73 | 125.7
[17 empty] | 20240206_HepClReRun_Ref_103 | Active T60 A | Analyte | 3.78 | 140990.125 | 4797147 |  | 140990.125 | 100 | 169.19 | 69.2
[17 empty] | 20240206_HepClReRun_Ref_104 | Active T60 B | Analyte | 3.78 | 193763.313 | 6576127 |  | 193763.313 | 100 | 232.53 | 132.5
[17 empty] | 20240206_HepClReRun_Ref_105 | Active T60 C | Analyte | 3.78 | 134827.359 | 4495299 |  | 134827.359 | 100 | 161.8 | 61.8
[17 empty] | 20240206_HepClReRun_Ref_109 | Active T120 A | Analyte | 3.78 | 66314.43 | 2659309 |  | 66314.43 | 100 | 79.58 | -20.4
[17 empty] | 20240206_HepClReRun_Ref_110 | Active T120 B | Analyte | 3.78 | 70664.453 | 2822867 |  | 70664.453 | 100 | 84.8 | -15.2
[17 empty] | 20240206_HepClReRun_Ref_111 | Active T120 C | Analyte | 3.79 | 90734.82 | 3446046 |  | 90734.82 | 100 | 108.89 | 8.9
[17 empty] | 20240206_HepClReRun_Ref_112 | Active T240 A | Analyte | 3.78 | 150100.688 | 5841353 |  | 150100.688 | 100 | 180.13 | 80.1
[17 empty] | 20240206_HepClReRun_Ref_113 | Active T240 B | Analyte | 3.78 | 121538.523 | 4873387 |  | 121538.523 | 100 | 145.85 | 45.9
[17 empty] | 20240206_HepClReRun_Ref_114 | Active T240 C | Analyte | 3.78 | 141360.891 | 5586628 |  | 141360.891 | 100 | 169.64 | 69.6
[17 empty] | 20240206_HepClReRun_Ref_118 | Media T0 A | Analyte | 3.78 | 59452.156 | 2364865 |  | 59452.156 | 100 | 71.35 | -28.7
[17 empty] | 20240206_HepClReRun_Ref_119 | Media T0 B | Analyte | 3.78 | 73160.086 | 2981668 |  | 73160.086 | 100 | 87.8 | -12.2
[17 empty] | 20240206_HepClReRun_Ref_120 | Media T0 C | Analyte | 3.78 | 81241.563 | 3158303 |  | 81241.563 | 100 | 97.49 | -2.5
[17 empty] | 20240206_HepClReRun_Ref_121 | Media T15 A | Analyte | 3.78 | 203907.672 | 8179395 |  | 203907.672 | 100 | 244.7 | 144.7
[17 empty] | 20240206_HepClReRun_Ref_122 | Media T15 B | Analyte | 3.78 | 205969.297 | 8093554 |  | 205969.297 | 100 | 247.17 | 147.2
[17 empty] | 20240206_HepClReRun_Ref_123 | Media T15 C | Analyte | 3.78 | 123378.438 | 5218319 |  | 123378.438 | 100 | 148.06 | 48.1
[17 empty] | 20240206_HepClReRun_Ref_125 | Media T30 A | Analyte | 3.78 | 105990.336 | 4481414 |  | 105990.336 | 100 | 127.19 | 27.2
[17 empty] | 20240206_HepClReRun_Ref_126 | Media T30 B | Analyte | 3.78 | 92672.016 | 3787446 |  | 92672.016 | 100 | 111.21 | 11.2
[17 empty] | 20240206_HepClReRun_Ref_127 | Media T30 C | Analyte | 3.78 | 122775.945 | 4730254 |  | 122775.945 | 100 | 147.34 | 47.3
[17 empty] | 20240206_HepClReRun_Ref_128 | Media T60 A | Analyte | 3.78 | 109065.461 | 4490368 |  | 109065.461 | 100 | 130.88 | 30.9
[17 empty] | 20240206_HepClReRun_Ref_129 | Media T60 B | Analyte | 3.78 | 111220.977 | 4359841 |  | 111220.977 | 100 | 133.47 | 33.5
[17 empty] | 20240206_HepClReRun_Ref_130 | Media T60 C | Analyte | 3.78 | 88635.961 | 3907847 |  | 88635.961 | 100 | 106.37 | 6.4
[17 empty] | 20240206_HepClReRun_Ref_134 | Media T120 A | Analyte | 3.78 | 99175.742 | 4210011 |  | 99175.742 | 100 | 119.02 | 19
[17 empty] | 20240206_HepClReRun_Ref_135 | Media T120 B | Analyte | 3.78 | 153366.016 | 6537087 |  | 153366.016 | 100 | 184.05 | 84
[17 empty] | 20240206_HepClReRun_Ref_136 | Media T120 C | Analyte | 3.78 | 156841.313 | 6584501 |  | 156841.313 | 100 | 188.22 | 88.2
[17 empty] | 20240206_HepClReRun_Ref_137 | Media T240 A | Analyte | 3.78 | 126010.648 | 5389455 |  | 126010.648 | 100 | 151.22 | 51.2
[17 empty] | 20240206_HepClReRun_Ref_138 | Media T240 B | Analyte | 3.78 | 133898.75 | 5696388 |  | 133898.75 | 100 | 160.68 | 60.7
[17 empty] | 20240206_HepClReRun_Ref_139 | Media T240 C | Analyte | 3.78 | 103058.133 | 4536677 |  | 103058.133 | 100 | 123.67 | 23.7
[17 empty] | 20240206_HepClReRun_Ref_143 | Inactive T0 A | Analyte | 3.78 | 67915.203 | 2974119 |  | 67915.203 | 100 | 81.5 | -18.5
[17 empty] | 20240206_HepClReRun_Ref_144 | Inactive T0 B | Analyte | 3.78 | 82405.945 | 3620762 |  | 82405.945 | 100 | 98.89 | -1.1
[17 empty] | 20240206_HepClReRun_Ref_145 | Inactive T0 C | Analyte | 3.78 | 57695.191 | 2476753 |  | 57695.191 | 100 | 69.24 | -30.8
[17 empty] | 20240206_HepClReRun_Ref_146 | Inactive T240 A | Analyte | 3.78 | 110553.828 | 4797423 |  | 110553.828 | 100 | 132.67 | 32.7
[17 empty] | 20240206_HepClReRun_Ref_147 | Inactive T240 B | Analyte | 3.78 | 79907.016 | 3404979 |  | 79907.016 | 100 | 95.89 | -4.1
[17 empty] | 20240206_HepClReRun_Ref_148 | Inactive T240 C | Analyte | 3.78 | 103302.727 | 4376327 |  | 103302.727 | 100 | 123.97 | 24
[17 empty] | 20240206_HepClReRun_Ref_003 | Matrix Blank | Blank | 3.82 | 15.171 | 367 |  | 15.171 | 100 | 0.02 | -100
[17 empty] | 20240206_HepClReRun_Ref_004 | Crash Blank | Blank | 3.78 | 1.89 | 77 |  | 1.89 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_001 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_002 | Blank | Blank | 3.84 | 11.192 | 512 |  | 11.192 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_005 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_006 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_013 | Blank | Blank | 3.79 | 14.119 | 559 |  | 14.119 | 100 | 0.02 | -100
[17 empty] | 20240206_HepClReRun_Ref_020 | Blank | Blank | 3.78 | 11.739 | 323 |  | 11.739 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_026 | Blank | Blank | 3.78 | 7.415 | 304 |  | 7.415 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_027 | Blank | Blank | 3.8 | 18.043 | 493 |  | 18.043 | 100 | 0.02 | -100
[17 empty] | 20240206_HepClReRun_Ref_032 | Blank | Blank | 3.79 | 1.047 | 123 |  | 1.047 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_033 | Blank | Blank | 3.84 | 12.055 | 251 |  | 12.055 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_038 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_043 | Blank | Blank | 3.79 | 9.778 | 359 |  | 9.778 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_048 | Blank | Blank | 3.81 | 16.982 | 459 |  | 16.982 | 100 | 0.02 | -100
[17 empty] | 20240206_HepClReRun_Ref_049 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_056 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_063 | Blank | Blank | 3.83 | 1.61 | 79 |  | 1.61 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_069 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_070 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_075 | Blank | Blank | 3.78 | 9.576 | 462 |  | 9.576 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_076 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_081 | Blank | Blank | 3.84 | 8.46 | 296 |  | 8.46 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_086 | Blank | Blank | 3.78 | 1.97 | 59 |  | 1.97 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_091 | Blank | Blank | 3.8 | 9.14 | 264 |  | 9.14 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_092 | Blank | Blank | 3.78 | 0.86 | 57 |  | 0.86 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_099 | Blank | Blank | 3.78 | 10.602 | 443 |  | 10.602 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_106 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_108 | Blank | Blank | 3.82 | 3.136 | 179 |  | 3.136 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_115 | Blank | Blank | 3.82 | 7.786 | 316 |  | 7.786 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_117 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_124 | Blank | Blank | 3.82 | 2.987 | 122 |  | 2.987 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_131 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_133 | Blank | Blank | 3.78 | 7.218 | 368 |  | 7.218 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_140 | Blank | Blank | 3.79 | 11.756 | 472 |  | 11.756 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_142 | Blank | Blank | 3.83 | 2.465 | 96 |  | 2.465 | 100 | 0 | -100
[17 empty] | 20240206_HepClReRun_Ref_149 | Blank | Blank | 3.84 | 6.103 | 264 |  | 6.103 | 100 | 0.01 | -100
[17 empty] | 20240206_HepClReRun_Ref_150 | Blank | Blank | [5 empty] | 100
[17 empty] | 20240206_HepClReRun_Ref_151 | Blank | Blank | 3.83 | 11.433 | 297 |  | 11.433 | 100 | 0.01 | -100

## Raw Data-NoRefCmpd
Quantify Compound Summary Report
Printed Thu Feb 08 11:11:15 2024
Compound 1: PFHxS
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20240206_HepCl_001 | Blank | Blank | 3.9 | 83.006 | 1481 | 81.312 | 102.083 |  | 7.31 |  | bb | 2024-02-06
2 | 2 | 20240206_HepCl_002 | Blank | Blank | 3.9 | 105.249 | 1243 | 121.729 | 86.462 |  | 6.19 |  | bb | 2024-02-06
3 | 3 | 20240206_HepCl_003 | Matrix Blank | Blank | 3.9 | 90.429 | 1328 | 160.252 | 56.429 |  | 4.03 |  | bb | 2024-02-06
4 | 4 | 20240206_HepCl_004 | Crash Blank | Blank | 3.9 | 85.059 | 1917 | 129.917 | 65.472 |  | 4.68 |  | bb | 2024-02-06
5 | 5 | 20240206_HepCl_005 | Blank | Blank | 3.89 | 81.687 | 1547 | 157.713 | 51.795 |  | 3.69 |  | bb | 2024-02-06
6 | 6 | 20240206_HepCl_006 | Blank | Blank | 3.89 | 81.964 | 1491 | 23.43 | 349.825 |  | 25.14 |  | bb | 2024-02-06
7 | 7 | 20240206_HepCl_007 | CC01 | Standard | 3.82 | 665.396 | 15303 | 13228.613 | 5.03 | 0.175 | 0.33 | 88.2 | bb | 2024-02-06
8 | 8 | 20240206_HepCl_008 | CC02 | Standard | 3.82 | 592.364 | 12612 | 13060.495 | 4.536 | 0.281 | 0.29 | 4.5 | bb | 2024-02-06
9 | 9 | 20240206_HepCl_009 | CC03 | Standard | 3.82 | 907.246 | 21508 | 12383.586 | 7.326 | 0.444 | 0.49 | 11.5 | bb | 2024-02-06
10 | 10 | 20240206_HepCl_010 | CC04 | Standard | 3.82 | 1198.078 | 30750 | 14313.511 | 8.37 | 0.713 | 0.57 | -20 | bb | 2024-02-06
11 | 11 | 20240206_HepCl_011 | CC05 | Standard | 3.82 | 2166.731 | 54951 | 10278.519 | 21.08 | 1.138 | 1.48 | 30.5 | bb | 2024-02-06
12 | 12 | 20240206_HepCl_012 | CC06 | Standard | 3.82 | 3070.819 | 77651 | 12826.483 | 23.941 | 1.825 | 1.69 | -7.4 | bb | 2024-02-06
13 | 13 | 20240206_HepCl_013 | Blank | Blank | 3.89 | 74.91 | 1535 | 119.286 | 62.799 |  | 4.49 |  | bb | 2024-02-06
14 | 14 | 20240206_HepCl_014 | CC07 | Standard | 3.82 | 4642.819 | 118697 | 13677.68 | 33.944 | 2.919 | 2.41 | -17.4 | bb | 2024-02-06
15 | 15 | 20240206_HepCl_015 | CC08 | Standard | 3.82 | 7219.885 | 193161 | 12671.938 | 56.975 | 4.669 | 4.07 | -12.9 | bb | 2024-02-06
16 | 16 | 20240206_HepCl_016 | CC09 | Standard | 3.82 | 9527.564 | 245663 | 9804.405 | 97.176 | 7.45 | 6.96 | -6.6 | bb | 2024-02-06
17 | 17 | 20240206_HepCl_017 | CC10 | Standard | 3.82 | 17529.984 | 443192 | 14090.989 | 124.406 | 11.875 | 8.92 | -24.9 | bb | 2024-02-06
18 | 18 | 20240206_HepCl_018 | CC11 | Standard | 3.82 | 30891.338 | 783737 | 12546.156 | 246.222 | 19.125 | 17.68 | -7.5 | bb | 2024-02-06
19 | 19 | 20240206_HepCl_019 | CC12 | Standard | 3.82 | 48357.563 | 1291990 | 12496.379 | 386.973 | 30.5 | 27.81 | -8.8 | bb | 2024-02-06
20 | 20 | 20240206_HepCl_020 | Blank | Blank | 3.89 | 90.015 | 1678 | 185.395 | 48.553 |  | 3.46 |  | bb | 2024-02-06
21 | 21 | 20240206_HepCl_021 | CC13 | Standard | 3.82 | 80208.766 | 2132800 | 12645.211 | 634.302 | 48.75 | 45.61 | -6.4 | bb | 2024-02-06
22 | 22 | 20240206_HepCl_022 | CC14 | Standard | 3.82 | 137819.438 | 3808929 | 13375.271 | 1030.405 | 78.125 | 74.11 | -5.1 | bb | 2024-02-06
23 | 23 | 20240206_HepCl_023 | CC15 | Standard | 3.82 | 221276.516 | 6151594 | 11024.747 | 2007.089 | 125 | 144.38 | 15.5 | bb | 2024-02-06
24 | 24 | 20240206_HepCl_024 | CC16 | Standard | 3.82 | 355811.5 | 9871077 | 15599.353 | 2280.938 | 200 | 164.09 | -18 | bb | 2024-02-06
25 | 25 | 20240206_HepCl_025 | CC17 | Standard | 3.82 | 572068.563 | 15426399 | 11725.054 | 4879.027 | 325.625 | 351.02 | 7.8 | bb | 2024-02-06
26 | 26 | 20240206_HepCl_026 | Blank | Blank | 3.89 | 67.843 | 1608 | 79.602 | 85.228 |  | 6.1 |  | bb | 2024-02-06
27 | 27 | 20240206_HepCl_027 | Blank | Blank | 3.89 | 88.088 | 1727 | 189.619 | 46.455 |  | 3.31 |  | bb | 2024-02-06
28 | 28 | 20240206_HepCl_028 | CCQC A | QC | 3.82 | 1670.837 | 42937 | 15079.548 | 11.08 | 0.625 | 0.76 | 22.3 | bb | 2024-02-06
29 | 29 | 20240206_HepCl_029 | CCQC B | QC | 3.82 | 6524.17 | 174808 | 17196.908 | 37.938 | 2.506 | 2.7 | 7.6 | bb | 2024-02-06
30 | 30 | 20240206_HepCl_030 | CCQC C | QC | 3.82 | 12457.592 | 324712 | 14326.982 | 86.952 | 6.269 | 6.22 | -0.7 | bb | 2024-02-06
31 | 31 | 20240206_HepCl_031 | CCQC D | QC | 3.82 | 52733.43 | 1380563 | 13883.457 | 379.829 | 25 | 27.3 | 9.2 | bb | 2024-02-06
32 | 32 | 20240206_HepCl_032 | Blank | Blank | 3.89 | 65.47 | 1652 | 0.622 | 10525.723 |  | 757.31 |  | bb | 2024-02-06
33 | 33 | 20240206_HepCl_033 | Blank | Blank | 3.89 | 69.892 | 1803 | 148.775 | 46.978 |  | 3.35 |  | bb | 2024-02-06
34 | 34 | 20240206_HepCl_034 | CC04 | Standard | 3.81 | 1166.829 | 28096 | 11906.045 | 9.8 | 0.713 | 0.67 | -5.6 | bb | 2024-02-06
35 | 35 | 20240206_HepCl_035 | CC05 | Standard | 3.82 | 2021.19 | 50930 | 12084.42 | 16.726 | 1.138 | 1.17 | 2.9 | bb | 2024-02-06
36 | 36 | 20240206_HepCl_036 | CC06 | Standard | 3.81 | 3175.081 | 83981 | 11320.455 | 28.047 | 1.825 | 1.99 | 8.8 | bb | 2024-02-06
37 | 37 | 20240206_HepCl_037 | CC13 | Standard | 3.82 | 73982.078 | 1876210 | 12121.979 | 610.314 | 48.75 | 43.88 | -10 | bb | 2024-02-06
38 | 38 | 20240206_HepCl_038 | Blank | Blank | 3.89 | 95.655 | 1716 | 193.399 | 49.46 |  | 3.53 |  | bb | 2024-02-06
39 | 39 | 20240206_HepCl_039 | CC04 | Standard | 3.82 | 1174.397 | 28083 | 12918.972 | 9.09 | 0.713 | 0.62 | -12.8 | bb | 2024-02-06
40 | 40 | 20240206_HepCl_040 | CC05 | Standard | 3.82 | 1709.357 | 41074 | 11582.531 | 14.758 | 1.138 | 1.03 | -9.5 | bb | 2024-02-06
41 | 41 | 20240206_HepCl_041 | CC06 | Standard | 3.82 | 3226.61 | 87559 | 13133.145 | 24.568 | 1.825 | 1.74 | -4.9 | bb | 2024-02-06
42 | 42 | 20240206_HepCl_042 | CC13 | Standard | 3.82 | 85995.836 | 2358451 | 12904.606 | 666.396 | 48.75 | 47.92 | -1.7 | bb | 2024-02-06
43 | 43 | 20240206_HepCl_043 | Blank | Blank | 3.88 | 95.703 | 1741 | 91.56 | 104.525 |  | 7.49 |  | bb | 2024-02-06
44 | 44 | 20240206_HepCl_044 | CC04 | Standard | 3.82 | 1285.76 | 31434 | 12008.481 | 10.707 | 0.713 | 0.74 | 3.6 | bb | 2024-02-06
45 | 45 | 20240206_HepCl_045 | CC05 | Standard | 3.82 | 1746.354 | 40763 | 11747.538 | 14.866 | 1.138 | 1.04 | -8.8 | bb | 2024-02-07
46 | 46 | 20240206_HepCl_046 | CC06 | Standard | 3.82 | 2751.479 | 67568 | 11973.123 | 22.98 | 1.825 | 1.62 | -11.2 | bb | 2024-02-07
47 | 47 | 20240206_HepCl_047 | CC13 | Standard | 3.82 | 85420.461 | 2232592 | 12636.182 | 675.999 | 48.75 | 48.61 | -0.3 | bb | 2024-02-07
48 | 48 | 20240206_HepCl_048 | Blank | Blank | 3.89 | 72.437 | 1606 | 45.974 | 157.561 |  | 11.3 |  | bb | 2024-02-07
49 | 49 | 20240206_HepCl_049 | Blank | Blank | 3.89 | 102.721 | 1862 | 30.281 | 339.226 |  | 24.38 |  | bb | 2024-02-07
50 | 50 | 20240206_HepCl_050 | CC01 | Standard | 3.82 | 644.485 | 14194 | 13672.755 | 4.714 | 0.175 | 0.31 | 75.2 | bb | 2024-02-07
51 | 51 | 20240206_HepCl_051 | CC02 | Standard | 3.82 | 651.616 | 14548 | 14006.428 | 4.652 | 0.281 | 0.3 | 7.4 | bb | 2024-02-07
52 | 52 | 20240206_HepCl_052 | CC03 | Standard | 3.82 | 811.079 | 17729 | 12324.795 | 6.581 | 0.444 | 0.44 | -0.6 | bb | 2024-02-07
53 | 53 | 20240206_HepCl_053 | CC04 | Standard | 3.82 | 1271.823 | 32126 | 12797.363 | 9.938 | 0.713 | 0.68 | -4.2 | bb | 2024-02-07
54 | 54 | 20240206_HepCl_054 | CC05 | Standard | 3.82 | 1961.893 | 48988 | 14087.408 | 13.927 | 1.138 | 0.97 | -14.8 | bb | 2024-02-07
55 | 55 | 20240206_HepCl_055 | CC06 | Standard | 3.82 | 3377.518 | 86556 | 12849.26 | 26.286 | 1.825 | 1.86 | 1.8 | bb | 2024-02-07
56 | 56 | 20240206_HepCl_056 | Blank | Blank | 3.88 | 89.901 | 1862 | 41.153 | 218.456 |  | 15.69 |  | bb | 2024-02-07
57 | 57 | 20240206_HepCl_057 | CC07 | Standard | 3.82 | 5172.834 | 138072 | 11294.601 | 45.799 | 2.919 | 3.26 | 11.8 | bb | 2024-02-07
58 | 58 | 20240206_HepCl_058 | CC08 | Standard | 3.82 | 8000.987 | 213848 | 13495.462 | 59.286 | 4.669 | 4.23 | -9.3 | bb | 2024-02-07
59 | 59 | 20240206_HepCl_059 | CC09 | Standard | 3.82 | 10462.401 | 267364 | 12311.349 | 84.982 | 7.45 | 6.08 | -18.4 | bb | 2024-02-07
60 | 60 | 20240206_HepCl_060 | CC10 | Standard | 3.82 | 20564.145 | 543171 | 12715.516 | 161.725 | 11.875 | 11.6 | -2.3 | bb | 2024-02-07
61 | 61 | 20240206_HepCl_061 | CC11 | Standard | 3.82 | 33561.285 | 903724 | 13536.77 | 247.927 | 19.125 | 17.81 | -6.9 | bb | 2024-02-07
62 | 62 | 20240206_HepCl_062 | CC12 | Standard | 3.82 | 59905.824 | 1629002 | 13279.041 | 451.131 | 30.5 | 32.43 | 6.3 | bb | 2024-02-07
63 | 63 | 20240206_HepCl_063 | Blank | Blank | 3.89 | 89.378 | 1697 | 128.157 | 69.741 |  | 4.99 |  | bb | 2024-02-07
64 | 64 | 20240206_HepCl_064 | CC13 | Standard | 3.82 | 88756.984 | 2438055 | 12511.845 | 709.384 | 48.75 | 51.01 | 4.6 | bb | 2024-02-07
65 | 65 | 20240206_HepCl_065 | CC14 | Standard | 3.82 | 135522 | 3683825 | 14586.896 | 929.067 | 78.125 | 66.82 | -14.5 | bb | 2024-02-07
66 | 66 | 20240206_HepCl_066 | CC15 | Standard | 3.82 | 234000.469 | 6380038 | 13164.907 | 1777.456 | 125 | 127.86 | 2.3 | bb | 2024-02-07
67 | 67 | 20240206_HepCl_067 | CC16 | Standard | 3.82 | 372184.063 | 10188026 | 15479.135 | 2404.424 | 200 | 172.97 | -13.5 | bb | 2024-02-07
68 | 68 | 20240206_HepCl_068 | CC17 | Standard | 3.82 | 638518.063 | 17553762 | 12682.382 | 5034.686 | 325.625 | 362.22 | 11.2 | bb | 2024-02-07
69 | 69 | 20240206_HepCl_069 | Blank | Blank | 3.9 | 63.475 | 1854 | 147.617 | 43 |  | 3.06 |  | bb | 2024-02-07
70 | 70 | 20240206_HepCl_070 | Blank | Blank | 3.89 | 78.396 | 1431 | 171.114 | 45.815 |  | 3.26 |  | bb | 2024-02-07
71 | 71 | 20240206_HepCl_071 | QC A | QC | 3.82 | 1721.165 | 40731 | 17512.383 | 9.828 | 0.625 | 0.67 | 7.9 | bb | 2024-02-07
72 | 72 | 20240206_HepCl_072 | QC B | QC | 3.82 | 6749.785 | 185797 | 17804.043 | 37.912 | 2.506 | 2.7 | 7.5 | bb | 2024-02-07
73 | 73 | 20240206_HepCl_073 | QC C | QC | 3.82 | 14612.47 | 390227 | 15293.799 | 95.545 | 6.269 | 6.84 | 9.1 | bb | 2024-02-07
74 | 74 | 20240206_HepCl_074 | QC D | QC | 3.82 | 62387.379 | 1673291 | 14440.207 | 432.039 | 25 | 31.05 | 24.2 | bb | 2024-02-07
75 | 75 | 20240206_HepCl_075 | Blank | Blank | 3.88 | 85.324 | 1906 | 212.657 | 40.123 |  | 2.85 |  | bb | 2024-02-07
76 | 76 | 20240206_HepCl_076 | Blank | Blank | 3.89 | 113.44 | 1819 | 110.334 | 102.815 |  | 7.37 |  | bb | 2024-02-07
77 | 77 | 20240206_HepCl_077 | CC04 | Standard | 3.82 | 1257.506 | 31257 | 10298.73 | 12.21 | 0.713 | 0.85 | 18.7 | bb | 2024-02-07
78 | 78 | 20240206_HepCl_078 | CC05 | Standard | 3.82 | 2021.016 | 50230 | 13693.741 | 14.759 | 1.138 | 1.03 | -9.5 | bb | 2024-02-07
79 | 79 | 20240206_HepCl_079 | CC06 | Standard | 3.82 | 3237.091 | 88521 | 13200.004 | 24.523 | 1.825 | 1.73 | -5.1 | bb | 2024-02-07
80 | 80 | 20240206_HepCl_080 | CC13 | Standard | 3.82 | 93043.977 | 2559129 | 12701.495 | 732.544 | 48.75 | 52.68 | 8.1 | bb | 2024-02-07
81 | 81 | 20240206_HepCl_081 | Blank | Blank | 3.89 | 114.742 | 1644 | 184.305 | 62.257 |  | 4.45 |  | bb | 2024-02-07
82 | 82 | 20240206_HepCl_082 | CC04 | Standard | 3.82 | 1128.267 | 27893 | 12015.393 | 9.39 | 0.713 | 0.64 | -9.7 | bb | 2024-02-07
83 | 83 | 20240206_HepCl_083 | CC05 | Standard | 3.82 | 2348.791 | 61487 | 15032.213 | 15.625 | 1.138 | 1.09 | -4 | bb | 2024-02-07
84 | 84 | 20240206_HepCl_084 | CC06 | Standard | 3.82 | 3132.317 | 79741 | 12810.994 | 24.45 | 1.825 | 1.73 | -5.4 | bb | 2024-02-07
85 | 85 | 20240206_HepCl_085 | CC13 | Standard | 3.82 | 80929.953 | 2105538 | 12612.149 | 641.682 | 48.75 | 46.14 | -5.4 | bb | 2024-02-07
86 | 86 | 20240206_HepCl_086 | Blank | Blank | 3.9 | 59.277 | 1597 | 48.337 | 122.633 |  | 8.79 |  | bb | 2024-02-07
87 | 87 | 20240206_HepCl_087 | CC04 | Standard | 3.82 | 1347.186 | 35163 | 15304.867 | 8.802 | 0.713 | 0.6 | -15.7 | bb | 2024-02-07
88 | 88 | 20240206_HepCl_088 | CC05 | Standard | 3.82 | 2182.837 | 54618 | 13440.987 | 16.24 | 1.138 | 1.14 | -0.1 | bb | 2024-02-07
89 | 89 | 20240206_HepCl_089 | CC06 | Standard | 3.82 | 3372.854 | 88752 | 13731.132 | 24.564 | 1.825 | 1.73 | -4.9 | bb | 2024-02-07
90 | 90 | 20240206_HepCl_090 | CC13 | Standard | 3.82 | 92521.656 | 2488279 | 11894.017 | 777.884 | 48.75 | 55.94 | 14.7 | bb | 2024-02-07
91 | 91 | 20240206_HepCl_091 | Blank | Blank | 3.89 | 66.831 | 1575 | 141.691 | 47.167 |  | 3.36 |  | bb | 2024-02-07
92 | 92 | 20240206_HepCl_092 | Blank | Blank | 3.89 | 140.583 | 1667 | 193.432 | 72.678 |  | 5.2 |  | bb | 2024-02-07
93 | 93 | 20240206_HepCl_093 | Active T0 A | Analyte | 3.82 | 86594.836 | 2325177 | 57196.434 | 151.399 |  | 10.86 |  | bb | 2024-02-07
94 | 94 | 20240206_HepCl_094 | Active T0 B | Analyte | 3.82 | 80860.594 | 1947570 | 52622.742 | 153.661 |  | 11.02 |  | bb | 2024-02-07
95 | 95 | 20240206_HepCl_095 | Active T0 C | Analyte | 3.82 | 84127 | 2268312 | 62977.305 | 133.583 |  | 9.58 |  | bb | 2024-02-07
96 | 96 | 20240206_HepCl_096 | Active T15 A | Analyte | 3.82 | 142816.328 | 3862124 | 54653.637 | 261.312 |  | 18.77 |  | bb | 2024-02-07
97 | 97 | 20240206_HepCl_097 | Active T15 B | Analyte | 3.82 | 97267.898 | 2538648 | 56251.129 | 172.917 |  | 12.41 |  | bb | 2024-02-07
98 | 98 | 20240206_HepCl_098 | Active T15 C | Analyte | 3.82 | 140795.516 | 3873275 | 46738.797 | 301.239 |  | 21.64 |  | bb | 2024-02-07
99 | 99 | 20240206_HepCl_099 | Blank | Blank | 3.89 | 66.924 | 1651 | 74.173 | 90.227 |  | 6.46 |  | bb | 2024-02-07
100 | 100 | 20240206_HepCl_100 | Active T30 A | Analyte | 3.82 | 100394.141 | 2828198 | 54554.977 | 184.024 |  | 13.21 |  | bb | 2024-02-07
101 | 101 | 20240206_HepCl_101 | Active T30 B | Analyte | 3.82 | 65320.938 | 1655156 | 57943.566 | 112.732 |  | 8.08 |  | bb | 2024-02-07
102 | 102 | 20240206_HepCl_102 | Active T30 C | Analyte | 3.82 | 69604.961 | 1912436 | 55111.348 | 126.299 |  | 9.05 |  | bb | 2024-02-07
103 | 103 | 20240206_HepCl_103 | Active T60 A | Analyte | 3.82 | 101476.094 | 2799021 | 59652.234 | 170.113 |  | 12.21 |  | bb | 2024-02-07
104 | 104 | 20240206_HepCl_104 | Active T60 B | Analyte | 3.82 | 77268.211 | 1955454 | 59296.051 | 130.309 |  | 9.34 |  | bb | 2024-02-07
105 | 105 | 20240206_HepCl_105 | Active T60 C | Analyte | 3.82 | 106937.984 | 2930513 | 65826.789 | 162.454 |  | 11.66 |  | bb | 2024-02-07
106 | 106 | 20240206_HepCl_106 | Blank | Blank | 3.9 | 80.859 | 1522 | 178.652 | 45.261 |  | 3.22 |  | bb | 2024-02-07
107 | 107 | 20240206_HepCl_107 | CC12 | Standard | 3.82 | 50414.809 | 1320994 | 12644.87 | 398.698 | 30.5 | 28.65 | -6.1 | bb | 2024-02-07
108 | 108 | 20240206_HepCl_108 | Blank | Blank | 3.89 | 99.444 | 1228 | 145.36 | 68.412 |  | 4.89 |  | bb | 2024-02-07
109 | 109 | 20240206_HepCl_109 | Active T120 A | Analyte | 3.82 | 93797.734 | 2549652 | 70924.125 | 132.251 |  | 9.48 |  | bb | 2024-02-07
110 | 110 | 20240206_HepCl_110 | Active T120 B | Analyte | 3.82 | 84825.398 | 2274040 | 72407.617 | 117.15 |  | 8.4 |  | bb | 2024-02-07
111 | 111 | 20240206_HepCl_111 | Active T120 C | Analyte | 3.82 | 83003.328 | 2307404 | 59619.336 | 139.222 |  | 9.98 |  | bb | 2024-02-07
112 | 112 | 20240206_HepCl_112 | Active T240 A | Analyte | 3.82 | 100694.75 | 2752351 | 65011.285 | 154.888 |  | 11.11 |  | bb | 2024-02-07
113 | 113 | 20240206_HepCl_113 | Active T240 B | Analyte | 3.82 | 113709.984 | 3235892 | 56173.52 | 202.426 |  | 14.53 |  | bb | 2024-02-07
114 | 114 | 20240206_HepCl_114 | Active T240 C | Analyte | 3.82 | 78901.656 | 2163128 | 60057.676 | 131.376 |  | 9.42 |  | bb | 2024-02-07
115 | 115 | 20240206_HepCl_115 | Blank | Blank | 3.89 | 138.643 | 3232 | 192.55 | 72.004 |  | 5.15 |  | bb | 2024-02-07
116 | 116 | 20240206_HepCl_116 | CC12 | Standard | 3.82 | 56819.621 | 1537002 | 13688.918 | 415.078 | 30.5 | 29.83 | -2.2 | bb | 2024-02-07
117 | 117 | 20240206_HepCl_117 | Blank | Blank | 3.89 | 81.174 | 1365 | 124.869 | 65.007 |  | 4.64 |  | bb | 2024-02-07
118 | 118 | 20240206_HepCl_118 | Media T0 A | Analyte | 3.82 | 100251.555 | 2701375 | 61092.672 | 164.098 |  | 11.77 |  | bb | 2024-02-07
119 | 119 | 20240206_HepCl_119 | Media T0 B | Analyte | 3.82 | 90407.914 | 2479259 | 57920.395 | 156.09 |  | 11.2 |  | bb | 2024-02-07
120 | 120 | 20240206_HepCl_120 | Media T0 C | Analyte | 3.82 | 78154.32 | 2074097 | 49501.477 | 157.883 |  | 11.33 |  | bb | 2024-02-07
121 | 121 | 20240206_HepCl_121 | Media T15 A | Analyte | 3.82 | 107893.641 | 2836428 | 50702.512 | 212.797 |  | 15.28 |  | bb | 2024-02-07
122 | 122 | 20240206_HepCl_122 | Media T15 B | Analyte | 3.82 | 73800.031 | 1945417 | 49429.43 | 149.304 |  | 10.71 |  | bb | 2024-02-07
123 | 123 | 20240206_HepCl_123 | Media T15 C | Analyte | 3.82 | 77809.773 | 2146174 | 58678.934 | 132.603 |  | 9.51 |  | bb | 2024-02-07
124 | 124 | 20240206_HepCl_124 | Blank | Blank | 3.9 | 75.368 | 1168 | 173.451 | 43.452 |  | 3.09 |  | bb | 2024-02-07
125 | 125 | 20240206_HepCl_125 | Media T30 A | Analyte | 3.82 | 72522.531 | 1973974 | 61647.859 | 117.64 |  | 8.43 |  | bb | 2024-02-07
126 | 126 | 20240206_HepCl_126 | Media T30 B | Analyte | 3.82 | 48810.473 | 1231894 | 61024.07 | 79.986 |  | 5.72 |  | bb | 2024-02-07
127 | 127 | 20240206_HepCl_127 | Media T30 C | Analyte | 3.82 | 67735.219 | 1841966 | 61411.234 | 110.298 |  | 7.9 |  | bb | 2024-02-07
128 | 128 | 20240206_HepCl_128 | Media T60 A | Analyte | 3.82 | 103834.203 | 2896273 | 61999.914 | 167.475 |  | 12.02 |  | bb | 2024-02-07
129 | 129 | 20240206_HepCl_129 | Media T60 B | Analyte | 3.82 | 88792.891 | 2441441 | 50020.82 | 177.512 |  | 12.74 |  | bb | 2024-02-07
130 | 130 | 20240206_HepCl_130 | Media T60 C | Analyte | 3.82 | 63867.574 | 1681176 | 62273.973 | 102.559 |  | 7.35 |  | bb | 2024-02-07
131 | 131 | 20240206_HepCl_131 | Blank | Blank | 3.9 | 91.398 | 1266 | 91.41 | 99.987 |  | 7.16 |  | bb | 2024-02-07
132 | 132 | 20240206_HepCl_132 | CC12 | Standard | 3.82 | 62562.719 | 1743878 | 13353.789 | 468.502 | 30.5 | 33.68 | 10.4 | bb | 2024-02-07
133 | 133 | 20240206_HepCl_133 | Blank | Blank | 3.74 | 10.297 | 251 | 139.391 | 7.387 |  | 0.5 |  | bb | 2024-02-07
134 | 134 | 20240206_HepCl_134 | Media T120 A | Analyte | 3.82 | 89899.516 | 2354959 | 67364.336 | 133.453 |  | 9.57 |  | bb | 2024-02-07
135 | 135 | 20240206_HepCl_135 | Media T120 B | Analyte | 3.82 | 69891.273 | 1889217 | 74916.945 | 93.292 |  | 6.68 |  | bb | 2024-02-07
136 | 136 | 20240206_HepCl_136 | Media T120 C | Analyte | 3.82 | 70295.148 | 1937328 | 65483.613 | 107.348 |  | 7.69 |  | bb | 2024-02-07
137 | 137 | 20240206_HepCl_137 | Media T240 A | Analyte | 3.82 | 86548.836 | 2110591 | 57828.598 | 149.664 |  | 10.74 |  | bb | 2024-02-07
138 | 138 | 20240206_HepCl_138 | Media T240 B | Analyte | 3.82 | 107041.172 | 2904692 | 60564.098 | 176.74 |  | 12.68 |  | bb | 2024-02-07
139 | 139 | 20240206_HepCl_139 | Media T240 C | Analyte | 3.82 | 86867.195 | 2313675 | 61393.453 | 141.493 |  | 10.15 |  | bb | 2024-02-07
140 | 140 | 20240206_HepCl_140 | Blank | Blank | 3.89 | 49.966 | 1137 | 54.732 | 91.292 |  | 6.54 |  | bb | 2024-02-07
141 | 141 | 20240206_HepCl_141 | CC12 | Standard | 3.82 | 63911.715 | 1723573 | 14119.17 | 452.659 | 30.5 | 32.54 | 6.7 | bb | 2024-02-07
142 | 142 | 20240206_HepCl_142 | Blank | Blank | 3.73 | 4.838 | 165 | 122.332 | 3.955 |  | 0.25 |  | bb | 2024-02-07
143 | 143 | 20240206_HepCl_143 | Inactive T0 A | Analyte | 3.82 | 69812.531 | 1812805 | 51183.547 | 136.396 |  | 9.78 |  | bb | 2024-02-07
144 | 144 | 20240206_HepCl_144 | Inactive T0 B | Analyte | 3.82 | 47744.313 | 1223507 | 60897.484 | 78.401 |  | 5.61 |  | bb | 2024-02-07
145 | 145 | 20240206_HepCl_145 | Inactive T0 C | Analyte | 3.82 | 54972.5 | 1454558 | 57956.52 | 94.851 |  | 6.79 |  | bb | 2024-02-07
146 | 146 | 20240206_HepCl_146 | Inactive T240 A | Analyte | 3.82 | 86538.484 | 2347388 | 54418.652 | 159.024 |  | 11.41 |  | bb | 2024-02-07
147 | 147 | 20240206_HepCl_147 | Inactive T240 B | Analyte | 3.82 | 64077.941 | 1700892 | 53906.863 | 118.868 |  | 8.52 |  | bb | 2024-02-07
148 | 148 | 20240206_HepCl_148 | Inactive T240 C | Analyte | 3.82 | 83679.945 | 2268970 | 59785.922 | 139.966 |  | 10.04 |  | bb | 2024-02-07
149 | 149 | 20240206_HepCl_149 | Blank | Blank | 3.9 | 68.14 | 1366 | 228.384 | 29.836 |  | 2.11 |  | bb | 2024-02-07
150 | 150 | 20240206_HepCl_150 | Blank | Blank | 3.89 | 61.63 | 1495 | 26.952 | 228.666 |  | 16.42 |  | bb | 2024-02-07
151 | 151 | 20240206_HepCl_151 | Blank | Blank | 3.89 | 61.384 | 1342 | 176.011 | 34.875 |  | 2.48 |  | bb | 2024-02-07
Compound 2: M3PFHxS
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20240206_HepCl_001 | Blank | Blank | 3.78 | 81.312 | 812 |  | 81.312 | 100 | 0.63 | -99.4 | bb | 2024-02-06
2 | 2 | 20240206_HepCl_002 | Blank | Blank | 3.8 | 121.729 | 673 |  | 121.729 | 100 | 0.95 | -99.1 | bb | 2024-02-06
3 | 3 | 20240206_HepCl_003 | Matrix Blank | Blank | 3.77 | 160.252 | 829 |  | 160.252 | 100 | 1.25 | -98.8 | bb | 2024-02-06
4 | 4 | 20240206_HepCl_004 | Crash Blank | Blank | 3.75 | 129.917 | 978 |  | 129.917 | 100 | 1.01 | -99 | bb | 2024-02-06
5 | 5 | 20240206_HepCl_005 | Blank | Blank | 3.81 | 157.713 | 840 |  | 157.713 | 100 | 1.23 | -98.8 | bb | 2024-02-06
6 | 6 | 20240206_HepCl_006 | Blank | Blank | 3.78 | 23.43 | 417 |  | 23.43 | 100 | 0.18 | -99.8 | bb | 2024-02-06
7 | 7 | 20240206_HepCl_007 | CC01 | Standard | 3.82 | 13228.613 | 413852 |  | 13228.613 | 100 | 102.85 | 2.9 | bb | 2024-02-06
8 | 8 | 20240206_HepCl_008 | CC02 | Standard | 3.82 | 13060.495 | 418054 |  | 13060.495 | 100 | 101.55 | 1.5 | bb | 2024-02-06
9 | 9 | 20240206_HepCl_009 | CC03 | Standard | 3.82 | 12383.586 | 377565 |  | 12383.586 | 100 | 96.28 | -3.7 | bb | 2024-02-06
10 | 10 | 20240206_HepCl_010 | CC04 | Standard | 3.82 | 14313.511 | 451057 |  | 14313.511 | 100 | 111.29 | 11.3 | bb | 2024-02-06
11 | 11 | 20240206_HepCl_011 | CC05 | Standard | 3.82 | 10278.519 | 300601 |  | 10278.519 | 100 | 79.92 | -20.1 | bb | 2024-02-06
12 | 12 | 20240206_HepCl_012 | CC06 | Standard | 3.82 | 12826.483 | 409590 |  | 12826.483 | 100 | 99.73 | -0.3 | bb | 2024-02-06
13 | 13 | 20240206_HepCl_013 | Blank | Blank | 3.76 | 119.286 | 826 |  | 119.286 | 100 | 0.93 | -99.1 | bb | 2024-02-06
14 | 14 | 20240206_HepCl_014 | CC07 | Standard | 3.82 | 13677.68 | 437952 |  | 13677.68 | 100 | 106.34 | 6.3 | bb | 2024-02-06
15 | 15 | 20240206_HepCl_015 | CC08 | Standard | 3.82 | 12671.938 | 390296 |  | 12671.938 | 100 | 98.52 | -1.5 | bb | 2024-02-06
16 | 16 | 20240206_HepCl_016 | CC09 | Standard | 3.82 | 9804.405 | 296585 |  | 9804.405 | 100 | 76.23 | -23.8 | bb | 2024-02-06
17 | 17 | 20240206_HepCl_017 | CC10 | Standard | 3.82 | 14090.989 | 442205 |  | 14090.989 | 100 | 109.56 | 9.6 | bb | 2024-02-06
18 | 18 | 20240206_HepCl_018 | CC11 | Standard | 3.82 | 12546.156 | 389787 |  | 12546.156 | 100 | 97.55 | -2.5 | bb | 2024-02-06
19 | 19 | 20240206_HepCl_019 | CC12 | Standard | 3.82 | 12496.379 | 384143 |  | 12496.379 | 100 | 97.16 | -2.8 | bb | 2024-02-06
20 | 20 | 20240206_HepCl_020 | Blank | Blank | 3.77 | 185.395 | 1192 |  | 185.395 | 100 | 1.44 | -98.6 | bb | 2024-02-06
21 | 21 | 20240206_HepCl_021 | CC13 | Standard | 3.82 | 12645.211 | 393316 |  | 12645.211 | 100 | 98.32 | -1.7 | bb | 2024-02-06
22 | 22 | 20240206_HepCl_022 | CC14 | Standard | 3.82 | 13375.271 | 417579 |  | 13375.271 | 100 | 103.99 | 4 | bb | 2024-02-06
23 | 23 | 20240206_HepCl_023 | CC15 | Standard | 3.82 | 11024.747 | 338200 |  | 11024.747 | 100 | 85.72 | -14.3 | bb | 2024-02-06
24 | 24 | 20240206_HepCl_024 | CC16 | Standard | 3.82 | 15599.353 | 485161 |  | 15599.353 | 100 | 121.29 | 21.3 | bb | 2024-02-06
25 | 25 | 20240206_HepCl_025 | CC17 | Standard | 3.82 | 11725.054 | 355157 |  | 11725.054 | 100 | 91.16 | -8.8 | bb | 2024-02-06
26 | 26 | 20240206_HepCl_026 | Blank | Blank | 3.78 | 79.602 | 700 |  | 79.602 | 100 | 0.62 | -99.4 | bb | 2024-02-06
27 | 27 | 20240206_HepCl_027 | Blank | Blank | 3.77 | 189.619 | 1152 |  | 189.619 | 100 | 1.47 | -98.5 | bb | 2024-02-06
28 | 28 | 20240206_HepCl_028 | CCQC A | QC | 3.82 | 15079.548 | 473390 |  | 15079.548 | 100 | 117.24 | 17.2 | bb | 2024-02-06
29 | 29 | 20240206_HepCl_029 | CCQC B | QC | 3.82 | 17196.908 | 543029 |  | 17196.908 | 100 | 133.71 | 33.7 | bb | 2024-02-06
30 | 30 | 20240206_HepCl_030 | CCQC C | QC | 3.82 | 14326.982 | 457026 |  | 14326.982 | 100 | 111.39 | 11.4 | bb | 2024-02-06
31 | 31 | 20240206_HepCl_031 | CCQC D | QC | 3.82 | 13883.457 | 426754 |  | 13883.457 | 100 | 107.94 | 7.9 | bb | 2024-02-06
32 | 32 | 20240206_HepCl_032 | Blank | Blank | 3.73 | 0.622 | 35 |  | 0.622 | 100 | 0 | -100 | bb | 2024-02-06
33 | 33 | 20240206_HepCl_033 | Blank | Blank | 3.77 | 148.775 | 910 |  | 148.775 | 100 | 1.16 | -98.8 | bb | 2024-02-06
34 | 34 | 20240206_HepCl_034 | CC04 | Standard | 3.81 | 11906.045 | 371992 |  | 11906.045 | 100 | 92.57 | -7.4 | bb | 2024-02-06
35 | 35 | 20240206_HepCl_035 | CC05 | Standard | 3.81 | 12084.42 | 372086 |  | 12084.42 | 100 | 93.96 | -6 | bb | 2024-02-06
36 | 36 | 20240206_HepCl_036 | CC06 | Standard | 3.81 | 11320.455 | 343551 |  | 11320.455 | 100 | 88.02 | -12 | bb | 2024-02-06
37 | 37 | 20240206_HepCl_037 | CC13 | Standard | 3.82 | 12121.979 | 372326 |  | 12121.979 | 100 | 94.25 | -5.8 | bb | 2024-02-06
38 | 38 | 20240206_HepCl_038 | Blank | Blank | 3.78 | 193.399 | 1174 |  | 193.399 | 100 | 1.5 | -98.5 | bb | 2024-02-06
39 | 39 | 20240206_HepCl_039 | CC04 | Standard | 3.82 | 12918.972 | 393904 |  | 12918.972 | 100 | 100.45 | 0.4 | bb | 2024-02-06
40 | 40 | 20240206_HepCl_040 | CC05 | Standard | 3.82 | 11582.531 | 353675 |  | 11582.531 | 100 | 90.05 | -9.9 | bb | 2024-02-06
41 | 41 | 20240206_HepCl_041 | CC06 | Standard | 3.82 | 13133.145 | 404806 |  | 13133.145 | 100 | 102.11 | 2.1 | bb | 2024-02-06
42 | 42 | 20240206_HepCl_042 | CC13 | Standard | 3.82 | 12904.606 | 398315 |  | 12904.606 | 100 | 100.33 | 0.3 | bb | 2024-02-06
43 | 43 | 20240206_HepCl_043 | Blank | Blank | 3.77 | 91.56 | 786 |  | 91.56 | 100 | 0.71 | -99.3 | bb | 2024-02-06
44 | 44 | 20240206_HepCl_044 | CC04 | Standard | 3.81 | 12008.481 | 363469 |  | 12008.481 | 100 | 93.37 | -6.6 | bb | 2024-02-06
45 | 45 | 20240206_HepCl_045 | CC05 | Standard | 3.82 | 11747.538 | 362078 |  | 11747.538 | 100 | 91.34 | -8.7 | bb | 2024-02-07
46 | 46 | 20240206_HepCl_046 | CC06 | Standard | 3.81 | 11973.123 | 361518 |  | 11973.123 | 100 | 93.09 | -6.9 | bb | 2024-02-07
47 | 47 | 20240206_HepCl_047 | CC13 | Standard | 3.82 | 12636.182 | 381013 |  | 12636.182 | 100 | 98.25 | -1.8 | bb | 2024-02-07
48 | 48 | 20240206_HepCl_048 | Blank | Blank | 3.77 | 45.974 | 527 |  | 45.974 | 100 | 0.36 | -99.6 | bb | 2024-02-07
49 | 49 | 20240206_HepCl_049 | Blank | Blank | 3.79 | 30.281 | 527 |  | 30.281 | 100 | 0.24 | -99.8 | bb | 2024-02-07
50 | 50 | 20240206_HepCl_050 | CC01 | Standard | 3.82 | 13672.755 | 422542 |  | 13672.755 | 100 | 106.31 | 6.3 | bb | 2024-02-07
51 | 51 | 20240206_HepCl_051 | CC02 | Standard | 3.82 | 14006.428 | 444985 |  | 14006.428 | 100 | 108.9 | 8.9 | bb | 2024-02-07
52 | 52 | 20240206_HepCl_052 | CC03 | Standard | 3.82 | 12324.795 | 377716 |  | 12324.795 | 100 | 95.83 | -4.2 | bb | 2024-02-07
53 | 53 | 20240206_HepCl_053 | CC04 | Standard | 3.82 | 12797.363 | 393599 |  | 12797.363 | 100 | 99.5 | -0.5 | bb | 2024-02-07
54 | 54 | 20240206_HepCl_054 | CC05 | Standard | 3.82 | 14087.408 | 443564 |  | 14087.408 | 100 | 109.53 | 9.5 | bb | 2024-02-07
55 | 55 | 20240206_HepCl_055 | CC06 | Standard | 3.82 | 12849.26 | 390870 |  | 12849.26 | 100 | 99.9 | -0.1 | bb | 2024-02-07
56 | 56 | 20240206_HepCl_056 | Blank | Blank | 3.78 | 41.153 | 652 |  | 41.153 | 100 | 0.32 | -99.7 | bb | 2024-02-07
57 | 57 | 20240206_HepCl_057 | CC07 | Standard | 3.82 | 11294.601 | 335217 |  | 11294.601 | 100 | 87.82 | -12.2 | bb | 2024-02-07
58 | 58 | 20240206_HepCl_058 | CC08 | Standard | 3.82 | 13495.462 | 416955 |  | 13495.462 | 100 | 104.93 | 4.9 | bb | 2024-02-07
59 | 59 | 20240206_HepCl_059 | CC09 | Standard | 3.82 | 12311.349 | 379701 |  | 12311.349 | 100 | 95.72 | -4.3 | bb | 2024-02-07
60 | 60 | 20240206_HepCl_060 | CC10 | Standard | 3.82 | 12715.516 | 385560 |  | 12715.516 | 100 | 98.86 | -1.1 | bb | 2024-02-07
61 | 61 | 20240206_HepCl_061 | CC11 | Standard | 3.82 | 13536.77 | 411755 |  | 13536.77 | 100 | 105.25 | 5.2 | bb | 2024-02-07
62 | 62 | 20240206_HepCl_062 | CC12 | Standard | 3.82 | 13279.041 | 410264 |  | 13279.041 | 100 | 103.24 | 3.2 | bb | 2024-02-07
63 | 63 | 20240206_HepCl_063 | Blank | Blank | 3.8 | 128.157 | 907 |  | 128.157 | 100 | 1 | -99 | bb | 2024-02-07
64 | 64 | 20240206_HepCl_064 | CC13 | Standard | 3.82 | 12511.845 | 372955 |  | 12511.845 | 100 | 97.28 | -2.7 | bb | 2024-02-07
65 | 65 | 20240206_HepCl_065 | CC14 | Standard | 3.82 | 14586.896 | 459507 |  | 14586.896 | 100 | 113.41 | 13.4 | bb | 2024-02-07
66 | 66 | 20240206_HepCl_066 | CC15 | Standard | 3.82 | 13164.907 | 405608 |  | 13164.907 | 100 | 102.36 | 2.4 | bb | 2024-02-07
67 | 67 | 20240206_HepCl_067 | CC16 | Standard | 3.82 | 15479.135 | 480047 |  | 15479.135 | 100 | 120.35 | 20.4 | bb | 2024-02-07
68 | 68 | 20240206_HepCl_068 | CC17 | Standard | 3.82 | 12682.382 | 373498 |  | 12682.382 | 100 | 98.61 | -1.4 | bb | 2024-02-07
69 | 69 | 20240206_HepCl_069 | Blank | Blank | 3.79 | 147.617 | 971 |  | 147.617 | 100 | 1.15 | -98.9 | bb | 2024-02-07
70 | 70 | 20240206_HepCl_070 | Blank | Blank | 3.78 | 171.114 | 972 |  | 171.114 | 100 | 1.33 | -98.7 | bb | 2024-02-07
71 | 71 | 20240206_HepCl_071 | QC A | QC | 3.82 | 17512.383 | 543966 |  | 17512.383 | 100 | 136.16 | 36.2 | bb | 2024-02-07
72 | 72 | 20240206_HepCl_072 | QC B | QC | 3.82 | 17804.043 | 566058 |  | 17804.043 | 100 | 138.43 | 38.4 | bb | 2024-02-07
73 | 73 | 20240206_HepCl_073 | QC C | QC | 3.82 | 15293.799 | 475924 |  | 15293.799 | 100 | 118.91 | 18.9 | bb | 2024-02-07
74 | 74 | 20240206_HepCl_074 | QC D | QC | 3.82 | 14440.207 | 442584 |  | 14440.207 | 100 | 112.27 | 12.3 | bb | 2024-02-07
75 | 75 | 20240206_HepCl_075 | Blank | Blank | 3.78 | 212.657 | 1293 |  | 212.657 | 100 | 1.65 | -98.3 | bb | 2024-02-07
76 | 76 | 20240206_HepCl_076 | Blank | Blank | 3.77 | 110.334 | 1025 |  | 110.334 | 100 | 0.86 | -99.1 | bb | 2024-02-07
77 | 77 | 20240206_HepCl_077 | CC04 | Standard | 3.82 | 10298.73 | 301611 |  | 10298.73 | 100 | 80.07 | -19.9 | bb | 2024-02-07
78 | 78 | 20240206_HepCl_078 | CC05 | Standard | 3.82 | 13693.741 | 424112 |  | 13693.741 | 100 | 106.47 | 6.5 | bb | 2024-02-07
79 | 79 | 20240206_HepCl_079 | CC06 | Standard | 3.82 | 13200.004 | 412523 |  | 13200.004 | 100 | 102.63 | 2.6 | bb | 2024-02-07
80 | 80 | 20240206_HepCl_080 | CC13 | Standard | 3.82 | 12701.495 | 388987 |  | 12701.495 | 100 | 98.75 | -1.2 | bb | 2024-02-07
81 | 81 | 20240206_HepCl_081 | Blank | Blank | 3.8 | 184.305 | 1099 |  | 184.305 | 100 | 1.43 | -98.6 | bb | 2024-02-07
82 | 82 | 20240206_HepCl_082 | CC04 | Standard | 3.82 | 12015.393 | 359388 |  | 12015.393 | 100 | 93.42 | -6.6 | bb | 2024-02-07
83 | 83 | 20240206_HepCl_083 | CC05 | Standard | 3.82 | 15032.213 | 466578 |  | 15032.213 | 100 | 116.88 | 16.9 | bb | 2024-02-07
84 | 84 | 20240206_HepCl_084 | CC06 | Standard | 3.82 | 12810.994 | 387593 |  | 12810.994 | 100 | 99.61 | -0.4 | bb | 2024-02-07
85 | 85 | 20240206_HepCl_085 | CC13 | Standard | 3.82 | 12612.149 | 388733 |  | 12612.149 | 100 | 98.06 | -1.9 | bb | 2024-02-07
86 | 86 | 20240206_HepCl_086 | Blank | Blank | 3.8 | 48.337 | 697 |  | 48.337 | 100 | 0.38 | -99.6 | bb | 2024-02-07
87 | 87 | 20240206_HepCl_087 | CC04 | Standard | 3.82 | 15304.867 | 492807 |  | 15304.867 | 100 | 119 | 19 | bb | 2024-02-07
88 | 88 | 20240206_HepCl_088 | CC05 | Standard | 3.82 | 13440.987 | 419437 |  | 13440.987 | 100 | 104.5 | 4.5 | bb | 2024-02-07
89 | 89 | 20240206_HepCl_089 | CC06 | Standard | 3.82 | 13731.132 | 432971 |  | 13731.132 | 100 | 106.76 | 6.8 | bb | 2024-02-07
90 | 90 | 20240206_HepCl_090 | CC13 | Standard | 3.82 | 11894.017 | 349842 |  | 11894.017 | 100 | 92.48 | -7.5 | bb | 2024-02-07
91 | 91 | 20240206_HepCl_091 | Blank | Blank | 3.79 | 141.691 | 640 |  | 141.691 | 100 | 1.1 | -98.9 | bb | 2024-02-07
92 | 92 | 20240206_HepCl_092 | Blank | Blank | 3.78 | 193.432 | 872 |  | 193.432 | 100 | 1.5 | -98.5 | bb | 2024-02-07
93 | 93 | 20240206_HepCl_093 | Active T0 A | Analyte | 3.82 | 57196.434 | 1753631 |  | 57196.434 | 100 | 444.7 | 344.7 | bb | 2024-02-07
94 | 94 | 20240206_HepCl_094 | Active T0 B | Analyte | 3.82 | 52622.742 | 1567755 |  | 52622.742 | 100 | 409.14 | 309.1 | bb | 2024-02-07
95 | 95 | 20240206_HepCl_095 | Active T0 C | Analyte | 3.82 | 62977.305 | 2001805 |  | 62977.305 | 100 | 489.65 | 389.7 | bb | 2024-02-07
96 | 96 | 20240206_HepCl_096 | Active T15 A | Analyte | 3.82 | 54653.637 | 1667995 |  | 54653.637 | 100 | 424.93 | 324.9 | bb | 2024-02-07
97 | 97 | 20240206_HepCl_097 | Active T15 B | Analyte | 3.82 | 56251.129 | 1713435 |  | 56251.129 | 100 | 437.35 | 337.4 | bb | 2024-02-07
98 | 98 | 20240206_HepCl_098 | Active T15 C | Analyte | 3.82 | 46738.797 | 1393550 |  | 46738.797 | 100 | 363.4 | 263.4 | bb | 2024-02-07
99 | 99 | 20240206_HepCl_099 | Blank | Blank | 3.78 | 74.173 | 512 |  | 74.173 | 100 | 0.58 | -99.4 | bb | 2024-02-07
100 | 100 | 20240206_HepCl_100 | Active T30 A | Analyte | 3.82 | 54554.977 | 1695409 |  | 54554.977 | 100 | 424.17 | 324.2 | bb | 2024-02-07
101 | 101 | 20240206_HepCl_101 | Active T30 B | Analyte | 3.82 | 57943.566 | 1843299 |  | 57943.566 | 100 | 450.51 | 350.5 | bb | 2024-02-07
102 | 102 | 20240206_HepCl_102 | Active T30 C | Analyte | 3.82 | 55111.348 | 1760062 |  | 55111.348 | 100 | 428.49 | 328.5 | bb | 2024-02-07
103 | 103 | 20240206_HepCl_103 | Active T60 A | Analyte | 3.82 | 59652.234 | 1889510 |  | 59652.234 | 100 | 463.8 | 363.8 | bb | 2024-02-07
104 | 104 | 20240206_HepCl_104 | Active T60 B | Analyte | 3.82 | 59296.051 | 1854851 |  | 59296.051 | 100 | 461.03 | 361 | bb | 2024-02-07
105 | 105 | 20240206_HepCl_105 | Active T60 C | Analyte | 3.82 | 65826.789 | 2056766 |  | 65826.789 | 100 | 511.81 | 411.8 | bb | 2024-02-07
106 | 106 | 20240206_HepCl_106 | Blank | Blank | 3.77 | 178.652 | 1024 |  | 178.652 | 100 | 1.39 | -98.6 | bb | 2024-02-07
107 | 107 | 20240206_HepCl_107 | CC12 | Standard | 3.82 | 12644.87 | 393543 |  | 12644.87 | 100 | 98.31 | -1.7 | bb | 2024-02-07
108 | 108 | 20240206_HepCl_108 | Blank | Blank | 3.66 | 145.36 | 734 |  | 145.36 | 100 | 1.13 | -98.9 | bb | 2024-02-07
109 | 109 | 20240206_HepCl_109 | Active T120 A | Analyte | 3.82 | 70924.125 | 2216592 |  | 70924.125 | 100 | 551.44 | 451.4 | bb | 2024-02-07
110 | 110 | 20240206_HepCl_110 | Active T120 B | Analyte | 3.82 | 72407.617 | 2285308 |  | 72407.617 | 100 | 562.97 | 463 | bb | 2024-02-07
111 | 111 | 20240206_HepCl_111 | Active T120 C | Analyte | 3.82 | 59619.336 | 1828233 |  | 59619.336 | 100 | 463.54 | 363.5 | bb | 2024-02-07
112 | 112 | 20240206_HepCl_112 | Active T240 A | Analyte | 3.82 | 65011.285 | 2034666 |  | 65011.285 | 100 | 505.46 | 405.5 | bb | 2024-02-07
113 | 113 | 20240206_HepCl_113 | Active T240 B | Analyte | 3.82 | 56173.52 | 1712859 |  | 56173.52 | 100 | 436.75 | 336.8 | bb | 2024-02-07
114 | 114 | 20240206_HepCl_114 | Active T240 C | Analyte | 3.82 | 60057.676 | 1861754 |  | 60057.676 | 100 | 466.95 | 367 | bb | 2024-02-07
115 | 115 | 20240206_HepCl_115 | Blank | Blank | 3.81 | 192.55 | 724 |  | 192.55 | 100 | 1.5 | -98.5 | bb | 2024-02-07
116 | 116 | 20240206_HepCl_116 | CC12 | Standard | 3.82 | 13688.918 | 427866 |  | 13688.918 | 100 | 106.43 | 6.4 | bb | 2024-02-07
117 | 117 | 20240206_HepCl_117 | Blank | Blank | 3.66 | 124.869 | 828 |  | 124.869 | 100 | 0.97 | -99 | bb | 2024-02-07
118 | 118 | 20240206_HepCl_118 | Media T0 A | Analyte | 3.82 | 61092.672 | 1890146 |  | 61092.672 | 100 | 475 | 375 | bb | 2024-02-07
119 | 119 | 20240206_HepCl_119 | Media T0 B | Analyte | 3.82 | 57920.395 | 1824850 |  | 57920.395 | 100 | 450.33 | 350.3 | bb | 2024-02-07
120 | 120 | 20240206_HepCl_120 | Media T0 C | Analyte | 3.82 | 49501.477 | 1486998 |  | 49501.477 | 100 | 384.88 | 284.9 | bb | 2024-02-07
121 | 121 | 20240206_HepCl_121 | Media T15 A | Analyte | 3.82 | 50702.512 | 1553175 |  | 50702.512 | 100 | 394.21 | 294.2 | bb | 2024-02-07
122 | 122 | 20240206_HepCl_122 | Media T15 B | Analyte | 3.82 | 49429.43 | 1506484 |  | 49429.43 | 100 | 384.32 | 284.3 | bb | 2024-02-07
123 | 123 | 20240206_HepCl_123 | Media T15 C | Analyte | 3.82 | 58678.934 | 1838720 |  | 58678.934 | 100 | 456.23 | 356.2 | bb | 2024-02-07
124 | 124 | 20240206_HepCl_124 | Blank | Blank | 3.77 | 173.451 | 797 |  | 173.451 | 100 | 1.35 | -98.7 | bb | 2024-02-07
125 | 125 | 20240206_HepCl_125 | Media T30 A | Analyte | 3.82 | 61647.859 | 1945396 |  | 61647.859 | 100 | 479.31 | 379.3 | bb | 2024-02-07
126 | 126 | 20240206_HepCl_126 | Media T30 B | Analyte | 3.82 | 61024.07 | 1875551 |  | 61024.07 | 100 | 474.46 | 374.5 | bb | 2024-02-07
127 | 127 | 20240206_HepCl_127 | Media T30 C | Analyte | 3.82 | 61411.234 | 1934272 |  | 61411.234 | 100 | 477.47 | 377.5 | bb | 2024-02-07
128 | 128 | 20240206_HepCl_128 | Media T60 A | Analyte | 3.82 | 61999.914 | 1957041 |  | 61999.914 | 100 | 482.05 | 382.1 | bb | 2024-02-07
129 | 129 | 20240206_HepCl_129 | Media T60 B | Analyte | 3.82 | 50020.82 | 1529021 |  | 50020.82 | 100 | 388.91 | 288.9 | bb | 2024-02-07
130 | 130 | 20240206_HepCl_130 | Media T60 C | Analyte | 3.82 | 62273.973 | 1903030 |  | 62273.973 | 100 | 484.18 | 384.2 | bb | 2024-02-07
131 | 131 | 20240206_HepCl_131 | Blank | Blank | 3.75 | 91.41 | 723 |  | 91.41 | 100 | 0.71 | -99.3 | bb | 2024-02-07
132 | 132 | 20240206_HepCl_132 | CC12 | Standard | 3.82 | 13353.789 | 412544 |  | 13353.789 | 100 | 103.83 | 3.8 | bb | 2024-02-07
133 | 133 | 20240206_HepCl_133 | Blank | Blank | 3.78 | 139.391 | 641 |  | 139.391 | 100 | 1.08 | -98.9 | bb | 2024-02-07
134 | 134 | 20240206_HepCl_134 | Media T120 A | Analyte | 3.82 | 67364.336 | 2137155 |  | 67364.336 | 100 | 523.76 | 423.8 | bb | 2024-02-07
135 | 135 | 20240206_HepCl_135 | Media T120 B | Analyte | 3.82 | 74916.945 | 2339204 |  | 74916.945 | 100 | 582.48 | 482.5 | bb | 2024-02-07
136 | 136 | 20240206_HepCl_136 | Media T120 C | Analyte | 3.82 | 65483.613 | 2024615 |  | 65483.613 | 100 | 509.14 | 409.1 | bb | 2024-02-07
137 | 137 | 20240206_HepCl_137 | Media T240 A | Analyte | 3.82 | 57828.598 | 1748022 |  | 57828.598 | 100 | 449.62 | 349.6 | bb | 2024-02-07
138 | 138 | 20240206_HepCl_138 | Media T240 B | Analyte | 3.82 | 60564.098 | 1891748 |  | 60564.098 | 100 | 470.89 | 370.9 | bb | 2024-02-07
139 | 139 | 20240206_HepCl_139 | Media T240 C | Analyte | 3.82 | 61393.453 | 1952005 |  | 61393.453 | 100 | 477.34 | 377.3 | bb | 2024-02-07
140 | 140 | 20240206_HepCl_140 | Blank | Blank | 3.78 | 54.732 | 682 |  | 54.732 | 100 | 0.43 | -99.6 | bb | 2024-02-07
141 | 141 | 20240206_HepCl_141 | CC12 | Standard | 3.82 | 14119.17 | 438989 |  | 14119.17 | 100 | 109.78 | 9.8 | bb | 2024-02-07
142 | 142 | 20240206_HepCl_142 | Blank | Blank | 3.8 | 122.332 | 558 |  | 122.332 | 100 | 0.95 | -99 | bb | 2024-02-07
143 | 143 | 20240206_HepCl_143 | Inactive T0 A | Analyte | 3.82 | 51183.547 | 1550672 |  | 51183.547 | 100 | 397.95 | 298 | bb | 2024-02-07
144 | 144 | 20240206_HepCl_144 | Inactive T0 B | Analyte | 3.82 | 60897.484 | 1924150 |  | 60897.484 | 100 | 473.48 | 373.5 | bb | 2024-02-07
145 | 145 | 20240206_HepCl_145 | Inactive T0 C | Analyte | 3.82 | 57956.52 | 1805876 |  | 57956.52 | 100 | 450.61 | 350.6 | bb | 2024-02-07
146 | 146 | 20240206_HepCl_146 | Inactive T240 A | Analyte | 3.82 | 54418.652 | 1682014 |  | 54418.652 | 100 | 423.11 | 323.1 | bb | 2024-02-07
147 | 147 | 20240206_HepCl_147 | Inactive T240 B | Analyte | 3.82 | 53906.863 | 1653033 |  | 53906.863 | 100 | 419.13 | 319.1 | bb | 2024-02-07
148 | 148 | 20240206_HepCl_148 | Inactive T240 C | Analyte | 3.82 | 59785.922 | 1831872 |  | 59785.922 | 100 | 464.84 | 364.8 | bb | 2024-02-07
149 | 149 | 20240206_HepCl_149 | Blank | Blank | 3.77 | 228.384 | 845 |  | 228.384 | 100 | 1.78 | -98.2 | bb | 2024-02-07
150 | 150 | 20240206_HepCl_150 | Blank | Blank | 3.79 | 26.952 | 415 |  | 26.952 | 100 | 0.21 | -99.8 | bb | 2024-02-07
151 | 151 | 20240206_HepCl_151 | Blank | Blank | 3.79 | 176.011 | 908 |  | 176.011 | 100 | 1.37 | -98.6 | bb | 2024-02-07
Compound 3: PFOA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20240206_HepCl_001 | Blank | Blank | 3.87 | 4602.562 | 58641 | 51.652 | 8910.714 |  | 699.36 |  | bb | 2024-02-06
2 | 2 | 20240206_HepCl_002 | Blank | Blank | 3.86 | 5433.033 | 70617 | 2.539 | 213983.182 |  | 16806.45 |  | bb | 2024-02-06
3 | 3 | 20240206_HepCl_003 | Matrix Blank | Blank | 3.85 | 5091.157 | 66742 | 77.175 | 6596.899 |  | 517.62 |  | bb | 2024-02-06
4 | 4 | 20240206_HepCl_004 | Crash Blank | Blank | 3.84 | 4666.37 | 63665 | 7.834 | 59565.611 |  | 4677.96 |  | bb | 2024-02-06
5 | 5 | 20240206_HepCl_005 | Blank | Blank | 3.84 | 5930.396 | 71987 | 12.389 | 47868.238 |  | 3759.21 |  | bb | 2024-02-06
6 | 6 | 20240206_HepCl_006 | Blank | Blank | 3.85 | 5161.199 | 67760 | 7.742 | 66664.932 |  | 5235.57 |  | bb | 2024-02-06
7 | 7 | 20240206_HepCl_007 | CC01 | Standard | 3.79 | 14689.672 | 349695 | 134058.391 | 10.958 | 0.175 | 0.34 | 93.2 | bb | 2024-02-06
8 | 8 | 20240206_HepCl_008 | CC02 | Standard | 3.79 | 14954.547 | 335498 | 118457.406 | 12.624 | 0.281 | 0.47 | 66.7 | bb | 2024-02-06
9 | 9 | 20240206_HepCl_009 | CC03 | Standard | 3.79 | 16010.505 | 399784 | 146184.828 | 10.952 | 0.444 | 0.34 | -23.9 | bb | 2024-02-06
10 | 10 | 20240206_HepCl_010 | CC04 | Standard | 3.79 | 18814.031 | 485555 | 134612.875 | 13.976 | 0.713 | 0.58 | -19.3 | bb | 2024-02-06
11 | 11 | 20240206_HepCl_011 | CC05 | Standard | 3.79 | 30746.496 | 862894 | 134493.047 | 22.861 | 1.138 | 1.27 | 11.9 | bb | 2024-02-06
12 | 12 | 20240206_HepCl_012 | CC06 | Standard | 3.79 | 33982.055 | 948781 | 114462.625 | 29.688 | 1.825 | 1.81 | -0.9 | bb | 2024-02-06
13 | 13 | 20240206_HepCl_013 | Blank | Blank | 3.85 | 4084.238 | 59924 | 27.52 | 14840.981 |  | 1165.14 |  | bb | 2024-02-06
14 | 14 | 20240206_HepCl_014 | CC07 | Standard | 3.79 | 54403.395 | 1644471 | 124790.297 | 43.596 | 2.919 | 2.9 | -0.6 | bb | 2024-02-06
15 | 15 | 20240206_HepCl_015 | CC08 | Standard | 3.79 | 83895.859 | 2620377 | 129505.023 | 64.782 | 4.669 | 4.57 | -2.2 | bb | 2024-02-06
16 | 16 | 20240206_HepCl_016 | CC09 | Standard | 3.79 | 116868.016 | 3671228 | 117331.039 | 99.605 | 7.45 | 7.3 | -2 | bb | 2024-02-06
17 | 17 | 20240206_HepCl_017 | CC10 | Standard | 3.79 | 193112.781 | 6144241 | 125046.508 | 154.433 | 11.875 | 11.61 | -2.3 | bb | 2024-02-06
18 | 18 | 20240206_HepCl_018 | CC11 | Standard | 3.79 | 378762.781 | 12584447 | 119904.609 | 315.887 | 19.125 | 24.29 | 27 | bb | 2024-02-06
19 | 19 | 20240206_HepCl_019 | CC12 | Standard | 3.79 | 543942.688 | 17636362 | 114519.438 | 474.978 | 30.5 | 36.78 | 20.6 | bb | 2024-02-06
20 | 20 | 20240206_HepCl_020 | Blank | Blank | 3.85 | 4801.735 | 63427 | [5 empty] | bb | 2024-02-06
21 | 21 | 20240206_HepCl_021 | CC13 | Standard | 3.79 | 655501.875 | 20798108 | 119743.57 | 547.421 | 48.75 | 42.47 | -12.9 | bb | 2024-02-06
22 | 22 | 20240206_HepCl_022 | CC14 | Standard | 3.79 | 1384856.5 | 45921412 | 111885.367 | 1237.746 | 78.125 | 96.69 | 23.8 | bb | 2024-02-06
23 | 23 | 20240206_HepCl_023 | CC15 | Standard | 3.79 | 1852453.625 | 59662784 | 89403.023 | 2072.026 | 125 | 162.22 | 29.8 | bb | 2024-02-06
24 | 24 | 20240206_HepCl_024 | CC16 | Standard | 3.79 | 2635860 | 83402032 | 133968.5 | 1967.522 | 200 | 154.01 | -23 | bb | 2024-02-06
25 | 25 | 20240206_HepCl_025 | CC17 | Standard | 3.79 | 3855763.5 | 121183752 | 90058.109 | 4281.417 | 325.625 | 335.75 | 3.1 | bb | 2024-02-06
26 | 26 | 20240206_HepCl_026 | Blank | Blank | 3.84 | 5251.62 | 64577 | 14.134 | 37155.936 |  | 2917.83 |  | bb | 2024-02-06
27 | 27 | 20240206_HepCl_027 | Blank | Blank | 3.86 | 5246.035 | 71244 | 12.725 | 41226.208 |  | 3237.53 |  | bb | 2024-02-06
28 | 28 | 20240206_HepCl_028 | CCQC A | QC | 3.79 | 21164.85 | 567620 | 144180.359 | 14.679 | 0.625 | 0.63 | 0.9 | bb | 2024-02-06
29 | 29 | 20240206_HepCl_029 | CCQC B | QC | 3.79 | 77265.156 | 2400383 | 177202.609 | 43.603 | 2.506 | 2.9 | 15.8 | bb | 2024-02-06
30 | 30 | 20240206_HepCl_030 | CCQC C | QC | 3.79 | 150378.359 | 4879617 | 138244.625 | 108.777 | 6.269 | 8.02 | 28 | bb | 2024-02-06
31 | 31 | 20240206_HepCl_031 | CCQC D | QC | 3.79 | 558387.188 | 18370440 | 142977.109 | 390.543 | 25 | 30.15 | 20.6 | bb | 2024-02-06
32 | 32 | 20240206_HepCl_032 | Blank | Blank | 3.84 | 5866.67 | 66857 | 66.006 | 8888.086 |  | 697.58 |  | bb | 2024-02-06
33 | 33 | 20240206_HepCl_033 | Blank | Blank | 3.84 | 4312.259 | 64444 | 0.973 | 443192.086 |  | 34809.31 |  | bb | 2024-02-06
34 | 34 | 20240206_HepCl_034 | CC04 | Standard | 3.78 | 18807.299 | 455470 | 136284.063 | 13.8 | 0.713 | 0.56 | -21.2 | bb | 2024-02-06
35 | 35 | 20240206_HepCl_035 | CC05 | Standard | 3.78 | 29587.498 | 837094 | 135748.172 | 21.796 | 1.138 | 1.19 | 4.6 | bb | 2024-02-06
36 | 36 | 20240206_HepCl_036 | CC06 | Standard | 3.79 | 35359.563 | 976728 | 130111.891 | 27.176 | 1.825 | 1.61 | -11.7 | bb | 2024-02-06
37 | 37 | 20240206_HepCl_037 | CC13 | Standard | 3.79 | 745636.563 | 24235634 | 135464.547 | 550.429 | 48.75 | 42.71 | -12.4 | bb | 2024-02-06
38 | 38 | 20240206_HepCl_038 | Blank | Blank | 3.85 | 5647.38 | 69157 | 25.41 | 22225.03 |  | 1745.11 |  | bb | 2024-02-06
39 | 39 | 20240206_HepCl_039 | CC04 | Standard | 3.79 | 20343.488 | 512426 | 149225.266 | 13.633 | 0.713 | 0.55 | -23.1 | bb | 2024-02-06
40 | 40 | 20240206_HepCl_040 | CC05 | Standard | 3.79 | 25224.449 | 708081 | 136307.844 | 18.506 | 1.138 | 0.93 | -18.2 | bb | 2024-02-06
41 | 41 | 20240206_HepCl_041 | CC06 | Standard | 3.79 | 39211.633 | 1135086 | 122278.68 | 32.067 | 1.825 | 2 | 9.4 | bb | 2024-02-06
42 | 42 | 20240206_HepCl_042 | CC13 | Standard | 3.79 | 699896.938 | 22405544 | 116004.852 | 603.334 | 48.75 | 46.87 | -3.9 | bb | 2024-02-06
43 | 43 | 20240206_HepCl_043 | Blank | Blank | 3.86 | 5520.508 | 70644 | 12.584 | 43869.263 |  | 3445.12 |  | bb | 2024-02-06
44 | 44 | 20240206_HepCl_044 | CC04 | Standard | 3.79 | 17677.65 | 434824 | 146873.625 | 12.036 | 0.713 | 0.42 | -40.7 | bb | 2024-02-06
45 | 45 | 20240206_HepCl_045 | CC05 | Standard | 3.78 | 29065.047 | 801278 | 143390.891 | 20.27 | 1.138 | 1.07 | -6 | bb | 2024-02-07
46 | 46 | 20240206_HepCl_046 | CC06 | Standard | 3.78 | 39930.25 | 1149416 | 135189.078 | 29.537 | 1.825 | 1.8 | -1.5 | bb | 2024-02-07
47 | 47 | 20240206_HepCl_047 | CC13 | Standard | 3.79 | 904464 | 29545996 | 140587.422 | 643.346 | 48.75 | 50.01 | 2.6 | bb | 2024-02-07
48 | 48 | 20240206_HepCl_048 | Blank | Blank | 3.85 | 5790.475 | 66684 | 13.116 | 44148.178 |  | 3467.03 |  | bb | 2024-02-07
49 | 49 | 20240206_HepCl_049 | Blank | Blank | 3.85 | 4424.665 | 60633 | 8.044 | 55005.781 |  | 4319.82 |  | bb | 2024-02-07
50 | 50 | 20240206_HepCl_050 | CC01 | Standard | 3.79 | 15336.969 | 336924 | 146270.547 | 10.485 | 0.175 | 0.3 | 72 | bb | 2024-02-07
51 | 51 | 20240206_HepCl_051 | CC02 | Standard | 3.79 | 16400.055 | 407683 | 153555.141 | 10.68 | 0.281 | 0.32 | 12.5 | bb | 2024-02-07
52 | 52 | 20240206_HepCl_052 | CC03 | Standard | 3.79 | 14273.191 | 344687 | 133897.094 | 10.66 | 0.444 | 0.31 | -29.1 | bb | 2024-02-07
53 | 53 | 20240206_HepCl_053 | CC04 | Standard | 3.79 | 16063.406 | 364068 | 112335.555 | 14.299 | 0.713 | 0.6 | -15.7 | bb | 2024-02-07
54 | 54 | 20240206_HepCl_054 | CC05 | Standard | 3.79 | 28396.881 | 758421 | 117635.977 | 24.14 | 1.138 | 1.37 | 20.7 | bb | 2024-02-07
55 | 55 | 20240206_HepCl_055 | CC06 | Standard | 3.79 | 40665.832 | 1205194 | 142075.469 | 28.623 | 1.825 | 1.73 | -5.4 | bb | 2024-02-07
56 | 56 | 20240206_HepCl_056 | Blank | Blank | 3.85 | 5508.639 | 79349 | 61.414 | 8969.68 |  | 703.99 |  | bb | 2024-02-07
57 | 57 | 20240206_HepCl_057 | CC07 | Standard | 3.79 | 52478.242 | 1539972 | 117855.555 | 44.528 | 2.919 | 2.97 | 1.9 | bb | 2024-02-07
58 | 58 | 20240206_HepCl_058 | CC08 | Standard | 3.79 | 87295.852 | 2678251 | 113011.789 | 77.245 | 4.669 | 5.54 | 18.8 | bb | 2024-02-07
59 | 59 | 20240206_HepCl_059 | CC09 | Standard | 3.79 | 110376.977 | 3406501 | 118829.922 | 92.887 | 7.45 | 6.77 | -9.1 | bb | 2024-02-07
60 | 60 | 20240206_HepCl_060 | CC10 | Standard | 3.79 | 226155.516 | 7283410 | 152427.219 | 148.37 | 11.875 | 11.13 | -6.3 | bb | 2024-02-07
61 | 61 | 20240206_HepCl_061 | CC11 | Standard | 3.79 | 378595.563 | 12403840 | 134633.641 | 281.204 | 19.125 | 21.56 | 12.8 | bb | 2024-02-07
62 | 62 | 20240206_HepCl_062 | CC12 | Standard | 3.79 | 542615.563 | 17402342 | 130157.563 | 416.891 | 30.5 | 32.22 | 5.6 | bb | 2024-02-07
63 | 63 | 20240206_HepCl_063 | Blank | Blank | 3.85 | 4456.979 | 67963 | [5 empty] | bb | 2024-02-07
64 | 64 | 20240206_HepCl_064 | CC13 | Standard | 3.79 | 780961.375 | 25277208 | 130902.141 | 596.599 | 48.75 | 46.34 | -5 | bb | 2024-02-07
65 | 65 | 20240206_HepCl_065 | CC14 | Standard | 3.79 | 1138025.875 | 36036232 | 122410.234 | 929.682 | 78.125 | 72.5 | -7.2 | bb | 2024-02-07
66 | 66 | 20240206_HepCl_066 | CC15 | Standard | 3.79 | 1911083.125 | 61873040 | 116367.969 | 1642.276 | 125 | 128.47 | 2.8 | bb | 2024-02-07
67 | 67 | 20240206_HepCl_067 | CC16 | Standard | 3.79 | 2784664 | 89580352 | 132073.734 | 2108.416 | 200 | 165.08 | -17.5 | bb | 2024-02-07
68 | 68 | 20240206_HepCl_068 | CC17 | Standard | 3.79 | 4283628.5 | 135021648 | 98631.766 | 4343.052 | 325.625 | 340.6 | 4.6 | bb | 2024-02-07
69 | 69 | 20240206_HepCl_069 | Blank | Blank | 3.85 | 4351.339 | 61376 | 70.645 | 6159.444 |  | 483.26 |  | bb | 2024-02-07
70 | 70 | 20240206_HepCl_070 | Blank | Blank | 3.85 | 5999.731 | 69681 | [5 empty] | bb | 2024-02-07
71 | 71 | 20240206_HepCl_071 | QC A | QC | 3.79 | 27920.525 | 739345 | 167490.781 | 16.67 | 0.625 | 0.79 | 25.9 | bb | 2024-02-07
72 | 72 | 20240206_HepCl_072 | QC B | QC | 3.79 | 63021.105 | 1860668 | 132315.688 | 47.629 | 2.506 | 3.22 | 28.4 | bb | 2024-02-07
73 | 73 | 20240206_HepCl_073 | QC C | QC | 3.79 | 162964.266 | 5223161 | 156897.484 | 103.867 | 6.269 | 7.64 | 21.8 | bb | 2024-02-07
74 | 74 | 20240206_HepCl_074 | QC D | QC | 3.79 | 623981.313 | 20126016 | 166386.391 | 375.019 | 25 | 28.93 | 15.7 | bb | 2024-02-07
75 | 75 | 20240206_HepCl_075 | Blank | Blank | 3.85 | 5932.333 | 74277 | 79.723 | 7441.181 |  | 583.93 |  | bb | 2024-02-07
76 | 76 | 20240206_HepCl_076 | Blank | Blank | 3.85 | 4732 | 69217 | 99.62 | 4750.05 |  | 372.56 |  | bb | 2024-02-07
77 | 77 | 20240206_HepCl_077 | CC04 | Standard | 3.79 | 19587.879 | 493252 | 150515.109 | 13.014 | 0.713 | 0.5 | -29.9 | bb | 2024-02-07
78 | 78 | 20240206_HepCl_078 | CC05 | Standard | 3.79 | 25835.053 | 668365 | 137131.031 | 18.84 | 1.138 | 0.96 | -15.9 | bb | 2024-02-07
79 | 79 | 20240206_HepCl_079 | CC06 | Standard | 3.79 | 37115.809 | 1071561 | 129202.703 | 28.727 | 1.825 | 1.73 | -5 | bb | 2024-02-07
80 | 80 | 20240206_HepCl_080 | CC13 | Standard | 3.78 | 922489.438 | 29977964 | 138754.016 | 664.838 | 48.75 | 51.7 | 6 | bb | 2024-02-07
81 | 81 | 20240206_HepCl_081 | Blank | Blank | 3.85 | 4536.292 | 61774 | 14.27 | 31789.012 |  | 2496.3 |  | bb | 2024-02-07
82 | 82 | 20240206_HepCl_082 | CC04 | Standard | 3.79 | 15517.276 | 358209 | 129994.75 | 11.937 | 0.713 | 0.41 | -41.8 | bb | 2024-02-07
83 | 83 | 20240206_HepCl_083 | CC05 | Standard | 3.79 | 32774.672 | 913665 | 128168.617 | 25.572 | 1.138 | 1.49 | 30.6 | bb | 2024-02-07
84 | 84 | 20240206_HepCl_084 | CC06 | Standard | 3.79 | 38892.891 | 1098905 | 135884.641 | 28.622 | 1.825 | 1.73 | -5.5 | bb | 2024-02-07
85 | 85 | 20240206_HepCl_085 | CC13 | Standard | 3.79 | 820857.625 | 26492756 | 142142.094 | 577.491 | 48.75 | 44.84 | -8 | bb | 2024-02-07
86 | 86 | 20240206_HepCl_086 | Blank | Blank | 3.85 | 4412.375 | 65235 | 10.153 | 43458.83 |  | 3412.88 |  | bb | 2024-02-07
87 | 87 | 20240206_HepCl_087 | CC04 | Standard | 3.79 | 17412.711 | 454139 | 157955.078 | 11.024 | 0.713 | 0.34 | -51.8 | bb | 2024-02-07
88 | 88 | 20240206_HepCl_088 | CC05 | Standard | 3.79 | 27265.17 | 740062 | 141630.219 | 19.251 | 1.138 | 0.99 | -13 | bb | 2024-02-07
89 | 89 | 20240206_HepCl_089 | CC06 | Standard | 3.79 | 38642.188 | 1151668 | 131731.141 | 29.334 | 1.825 | 1.78 | -2.4 | bb | 2024-02-07
90 | 90 | 20240206_HepCl_090 | CC13 | Standard | 3.79 | 858363.625 | 27619504 | 125548.438 | 683.691 | 48.75 | 53.18 | 9.1 | bb | 2024-02-07
91 | 91 | 20240206_HepCl_091 | Blank | Blank | 3.85 | 4792.176 | 72944 | 8.037 | 59626.428 |  | 4682.74 |  | bb | 2024-02-07
92 | 92 | 20240206_HepCl_092 | Blank | Blank | 3.85 | 5023.125 | 73481 | 40.381 | 12439.328 |  | 976.5 |  | bb | 2024-02-07
93 | 93 | 20240206_HepCl_093 | Active T0 A | Analyte | 3.79 | 757865.313 | 23980454 | 483745.031 | 156.666 |  | 11.78 |  | bb | 2024-02-07
94 | 94 | 20240206_HepCl_094 | Active T0 B | Analyte | 3.79 | 685363.375 | 21845232 | 459075.188 | 149.292 |  | 11.2 |  | bb | 2024-02-07
95 | 95 | 20240206_HepCl_095 | Active T0 C | Analyte | 3.79 | 748576.5 | 24609110 | 528794.375 | 141.563 |  | 10.6 |  | bb | 2024-02-07
96 | 96 | 20240206_HepCl_096 | Active T15 A | Analyte | 3.79 | 1012456.563 | 33141090 | 551724.25 | 183.508 |  | 13.89 |  | bb | 2024-02-07
97 | 97 | 20240206_HepCl_097 | Active T15 B | Analyte | 3.79 | 968402.5 | 32242352 | 583719.563 | 165.902 |  | 12.51 |  | bb | 2024-02-07
98 | 98 | 20240206_HepCl_098 | Active T15 C | Analyte | 3.79 | 786365.438 | 25254146 | 514480.375 | 152.847 |  | 11.48 |  | bb | 2024-02-07
99 | 99 | 20240206_HepCl_099 | Blank | Blank | 3.85 | 4249.628 | 62655 | 5.725 | 74229.31 |  | 5829.7 |  | bb | 2024-02-07
100 | 100 | 20240206_HepCl_100 | Active T30 A | Analyte | 3.79 | 726149.313 | 23132068 | 487161.469 | 149.057 |  | 11.18 |  | bb | 2024-02-07
101 | 101 | 20240206_HepCl_101 | Active T30 B | Analyte | 3.78 | 778169.25 | 25243274 | 572075.313 | 136.026 |  | 10.16 |  | bb | 2024-02-07
102 | 102 | 20240206_HepCl_102 | Active T30 C | Analyte | 3.79 | 675882.688 | 21901790 | 481638.969 | 140.33 |  | 10.5 |  | bb | 2024-02-07
103 | 103 | 20240206_HepCl_103 | Active T60 A | Analyte | 3.79 | 1018539.625 | 33397900 | 530218.813 | 192.098 |  | 14.57 |  | bb | 2024-02-07
104 | 104 | 20240206_HepCl_104 | Active T60 B | Analyte | 3.79 | 787809.75 | 25440130 | 598683.313 | 131.59 |  | 9.81 |  | bb | 2024-02-07
105 | 105 | 20240206_HepCl_105 | Active T60 C | Analyte | 3.79 | 989299.5 | 32508514 | 600279.688 | 164.806 |  | 12.42 |  | bb | 2024-02-07
106 | 106 | 20240206_HepCl_106 | Blank | Blank | 3.85 | 5541.751 | 75448 | [5 empty] | bb | 2024-02-07
107 | 107 | 20240206_HepCl_107 | CC12 | Standard | 3.78 | 468996.781 | 14921748 | 116307.844 | 403.237 | 30.5 | 31.15 | 2.1 | bb | 2024-02-07
108 | 108 | 20240206_HepCl_108 | Blank | Blank | 3.85 | 4669.948 | 65573 | 12.753 | 36618.427 |  | 2875.61 |  | bb | 2024-02-07
109 | 109 | 20240206_HepCl_109 | Active T120 A | Analyte | 3.79 | 843010.313 | 27373468 | 748670.625 | 112.601 |  | 8.32 |  | bb | 2024-02-07
110 | 110 | 20240206_HepCl_110 | Active T120 B | Analyte | 3.79 | 697223.625 | 22892238 | 611171.313 | 114.08 |  | 8.44 |  | bb | 2024-02-07
111 | 111 | 20240206_HepCl_111 | Active T120 C | Analyte | 3.79 | 590744.375 | 18706572 | 630458.313 | 93.701 |  | 6.84 |  | bb | 2024-02-07
112 | 112 | 20240206_HepCl_112 | Active T240 A | Analyte | 3.79 | 589344.875 | 18359838 | 530024 | 111.192 |  | 8.21 |  | bb | 2024-02-07
113 | 113 | 20240206_HepCl_113 | Active T240 B | Analyte | 3.79 | 637305.75 | 20351406 | 553846.313 | 115.069 |  | 8.52 |  | bb | 2024-02-07
114 | 114 | 20240206_HepCl_114 | Active T240 C | Analyte | 3.79 | 594142.438 | 19279252 | 517922.813 | 114.716 |  | 8.49 |  | bb | 2024-02-07
115 | 115 | 20240206_HepCl_115 | Blank | Blank | 3.84 | 3708.198 | 57257 | 0.544 | 681654.044 |  | 53538.93 |  | bb | 2024-02-07
116 | 116 | 20240206_HepCl_116 | CC12 | Standard | 3.78 | 508351.656 | 16565745 | 136432.594 | 372.603 | 30.5 | 28.74 | -5.8 | bb | 2024-02-07
117 | 117 | 20240206_HepCl_117 | Blank | Blank | 3.85 | 4831.553 | 65807 | 1.556 | 310511.118 |  | 24388.09 |  | bb | 2024-02-07
118 | 118 | 20240206_HepCl_118 | Media T0 A | Analyte | 3.79 | 706007.5 | 22494514 | 483696.719 | 145.961 |  | 10.94 |  | bb | 2024-02-07
119 | 119 | 20240206_HepCl_119 | Media T0 B | Analyte | 3.78 | 715865.375 | 23240018 | 554689.25 | 129.057 |  | 9.61 |  | bb | 2024-02-07
120 | 120 | 20240206_HepCl_120 | Media T0 C | Analyte | 3.79 | 741524.938 | 24445986 | 605034.563 | 122.559 |  | 9.1 |  | bb | 2024-02-07
121 | 121 | 20240206_HepCl_121 | Media T15 A | Analyte | 3.79 | 804611.063 | 25870144 | 436108.25 | 184.498 |  | 13.97 |  | bb | 2024-02-07
122 | 122 | 20240206_HepCl_122 | Media T15 B | Analyte | 3.79 | 898459.813 | 29312212 | 497355.656 | 180.647 |  | 13.67 |  | bb | 2024-02-07
123 | 123 | 20240206_HepCl_123 | Media T15 C | Analyte | 3.79 | 632550.5 | 20060792 | 458635.219 | 137.92 |  | 10.31 |  | bb | 2024-02-07
124 | 124 | 20240206_HepCl_124 | Blank | Blank | 3.84 | 5092.387 | 64352 | 121.379 | 4195.443 |  | 329 |  | bb | 2024-02-07
125 | 125 | 20240206_HepCl_125 | Media T30 A | Analyte | 3.79 | 692136.75 | 22767466 | 549553.5 | 125.945 |  | 9.37 |  | bb | 2024-02-07
126 | 126 | 20240206_HepCl_126 | Media T30 B | Analyte | 3.79 | 528796.875 | 16542680 | 542602.625 | 97.456 |  | 7.13 |  | bb | 2024-02-07
127 | 127 | 20240206_HepCl_127 | Media T30 C | Analyte | 3.79 | 671984.125 | 22347616 | 572704.688 | 117.335 |  | 8.69 |  | bb | 2024-02-07
128 | 128 | 20240206_HepCl_128 | Media T60 A | Analyte | 3.79 | 831415.125 | 26780574 | 517056.781 | 160.798 |  | 12.11 |  | bb | 2024-02-07
129 | 129 | 20240206_HepCl_129 | Media T60 B | Analyte | 3.79 | 825454.438 | 26125740 | 545332.438 | 151.367 |  | 11.37 |  | bb | 2024-02-07
130 | 130 | 20240206_HepCl_130 | Media T60 C | Analyte | 3.79 | 899184.625 | 28970670 | 564425.188 | 159.31 |  | 11.99 |  | bb | 2024-02-07
131 | 131 | 20240206_HepCl_131 | Blank | Blank | 3.85 | 5546.071 | 72105 | 9.231 | 60080.934 |  | 4718.44 |  | bb | 2024-02-07
132 | 132 | 20240206_HepCl_132 | CC12 | Standard | 3.79 | 557822.938 | 18217662 | 125093.547 | 445.925 | 30.5 | 34.5 | 13.1 | bb | 2024-02-07
133 | 133 | 20240206_HepCl_133 | Blank | Blank | 3.85 | 4060.826 | 54872 | 101.073 | 4017.716 |  | 315.04 |  | bb | 2024-02-07
134 | 134 | 20240206_HepCl_134 | Media T120 A | Analyte | 3.79 | 488209.906 | 15396140 | 646892.625 | 75.47 |  | 5.41 |  | bb | 2024-02-07
135 | 135 | 20240206_HepCl_135 | Media T120 B | Analyte | 3.79 | 454990.031 | 14362970 | 527762.875 | 86.211 |  | 6.25 |  | bb | 2024-02-07
136 | 136 | 20240206_HepCl_136 | Media T120 C | Analyte | 3.79 | 586462.188 | 19158654 | 616629.188 | 95.108 |  | 6.95 |  | bb | 2024-02-07
137 | 137 | 20240206_HepCl_137 | Media T240 A | Analyte | 3.79 | 529178.5 | 16877646 | 554051.313 | 95.511 |  | 6.98 |  | bb | 2024-02-07
138 | 138 | 20240206_HepCl_138 | Media T240 B | Analyte | 3.79 | 547773.438 | 17375820 | 507204.156 | 107.999 |  | 7.96 |  | bb | 2024-02-07
139 | 139 | 20240206_HepCl_139 | Media T240 C | Analyte | 3.79 | 736954 | 24175192 | 631669 | 116.668 |  | 8.64 |  | bb | 2024-02-07
140 | 140 | 20240206_HepCl_140 | Blank | Blank | 3.85 | 4071.576 | 58685 | 16.476 | 24712.163 |  | 1940.46 |  | bb | 2024-02-07
141 | 141 | 20240206_HepCl_141 | CC12 | Standard | 3.79 | 532644.188 | 17493900 | 142639.484 | 373.42 | 30.5 | 28.81 | -5.6 | bb | 2024-02-07
142 | 142 | 20240206_HepCl_142 | Blank | Blank | 3.85 | 4597.399 | 64995 | 99.922 | 4600.988 |  | 360.85 |  | bb | 2024-02-07
143 | 143 | 20240206_HepCl_143 | Inactive T0 A | Analyte | 3.79 | 582762.813 | 19011156 | 525006.813 | 111.001 |  | 8.2 |  | bb | 2024-02-07
144 | 144 | 20240206_HepCl_144 | Inactive T0 B | Analyte | 3.79 | 544444.063 | 17852800 | 551036.875 | 98.804 |  | 7.24 |  | bb | 2024-02-07
145 | 145 | 20240206_HepCl_145 | Inactive T0 C | Analyte | 3.79 | 446171.031 | 14175868 | 475163.781 | 93.898 |  | 6.85 |  | bb | 2024-02-07
146 | 146 | 20240206_HepCl_146 | Inactive T240 A | Analyte | 3.79 | 550984.25 | 17694306 | 448450.656 | 122.864 |  | 9.13 |  | bb | 2024-02-07
147 | 147 | 20240206_HepCl_147 | Inactive T240 B | Analyte | 3.79 | 445698.344 | 14486281 | 490642.688 | 90.84 |  | 6.61 |  | bb | 2024-02-07
148 | 148 | 20240206_HepCl_148 | Inactive T240 C | Analyte | 3.79 | 624689.375 | 20648322 | 618428.688 | 101.012 |  | 7.41 |  | bb | 2024-02-07
149 | 149 | 20240206_HepCl_149 | Blank | Blank | 3.85 | 5672.912 | 66329 | 104.156 | 5446.553 |  | 427.27 |  | bb | 2024-02-07
150 | 150 | 20240206_HepCl_150 | Blank | Blank | 3.85 | 4435.503 | 61997 | 107.154 | 4139.372 |  | 324.6 |  | bb | 2024-02-07
151 | 151 | 20240206_HepCl_151 | Blank | Blank | 3.85 | 4759.074 | 64106 | 108.747 | 4376.281 |  | 343.21 |  | bb | 2024-02-07
Compound 4: M8PFOA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20240206_HepCl_001 | Blank | Blank | 3.77 | 51.652 | 463 |  | 51.652 | 100 | 0.04 | -100 | bb | 2024-02-06
2 | 2 | 20240206_HepCl_002 | Blank | Blank | 3.86 | 2.539 | 144 |  | 2.539 | 100 | 0 | -100 | bb | 2024-02-06
3 | 3 | 20240206_HepCl_003 | Matrix Blank | Blank | 3.88 | 77.175 | 526 |  | 77.175 | 100 | 0.06 | -99.9 | bb | 2024-02-06
4 | 4 | 20240206_HepCl_004 | Crash Blank | Blank | 3.9 | 7.834 | 292 |  | 7.834 | 100 | 0.01 | -100 | bb | 2024-02-06
5 | 5 | 20240206_HepCl_005 | Blank | Blank | 3.86 | 12.389 | 310 |  | 12.389 | 100 | 0.01 | -100 | bb | 2024-02-06
6 | 6 | 20240206_HepCl_006 | Blank | Blank | 3.8 | 7.742 | 287 |  | 7.742 | 100 | 0.01 | -100 | bb | 2024-02-06
7 | 7 | 20240206_HepCl_007 | CC01 | Standard | 3.79 | 134058.391 | 4381584 |  | 134058.391 | 100 | 103.53 | 3.5 | bb | 2024-02-06
8 | 8 | 20240206_HepCl_008 | CC02 | Standard | 3.79 | 118457.406 | 3813211 |  | 118457.406 | 100 | 91.48 | -8.5 | bb | 2024-02-06
9 | 9 | 20240206_HepCl_009 | CC03 | Standard | 3.79 | 146184.828 | 4807071 |  | 146184.828 | 100 | 112.9 | 12.9 | bb | 2024-02-06
10 | 10 | 20240206_HepCl_010 | CC04 | Standard | 3.79 | 134612.875 | 4458353 |  | 134612.875 | 100 | 103.96 | 4 | bb | 2024-02-06
11 | 11 | 20240206_HepCl_011 | CC05 | Standard | 3.79 | 134493.047 | 4452551 |  | 134493.047 | 100 | 103.87 | 3.9 | bb | 2024-02-06
12 | 12 | 20240206_HepCl_012 | CC06 | Standard | 3.79 | 114462.625 | 3698309 |  | 114462.625 | 100 | 88.4 | -11.6 | bb | 2024-02-06
13 | 13 | 20240206_HepCl_013 | Blank | Blank | 3.64 | 27.52 | 389 |  | 27.52 | 100 | 0.02 | -100 | bb | 2024-02-06
14 | 14 | 20240206_HepCl_014 | CC07 | Standard | 3.79 | 124790.297 | 4079174 |  | 124790.297 | 100 | 96.37 | -3.6 | bb | 2024-02-06
15 | 15 | 20240206_HepCl_015 | CC08 | Standard | 3.79 | 129505.023 | 4280320 |  | 129505.023 | 100 | 100.02 | 0 | bb | 2024-02-06
16 | 16 | 20240206_HepCl_016 | CC09 | Standard | 3.79 | 117331.039 | 3835934 |  | 117331.039 | 100 | 90.61 | -9.4 | bb | 2024-02-06
17 | 17 | 20240206_HepCl_017 | CC10 | Standard | 3.79 | 125046.508 | 4131017 |  | 125046.508 | 100 | 96.57 | -3.4 | bb | 2024-02-06
18 | 18 | 20240206_HepCl_018 | CC11 | Standard | 3.79 | 119904.609 | 3949575 |  | 119904.609 | 100 | 92.6 | -7.4 | bb | 2024-02-06
19 | 19 | 20240206_HepCl_019 | CC12 | Standard | 3.79 | 114519.438 | 3692099 |  | 114519.438 | 100 | 88.44 | -11.6 | bb | 2024-02-06
20 | 20 | 20240206_HepCl_020 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-06
21 | 21 | 20240206_HepCl_021 | CC13 | Standard | 3.79 | 119743.57 | 3841606 |  | 119743.57 | 100 | 92.48 | -7.5 | bb | 2024-02-06
22 | 22 | 20240206_HepCl_022 | CC14 | Standard | 3.79 | 111885.367 | 3616760 |  | 111885.367 | 100 | 86.41 | -13.6 | bb | 2024-02-06
23 | 23 | 20240206_HepCl_023 | CC15 | Standard | 3.79 | 89403.023 | 2825222 |  | 89403.023 | 100 | 69.04 | -31 | bb | 2024-02-06
24 | 24 | 20240206_HepCl_024 | CC16 | Standard | 3.78 | 133968.5 | 4275505 |  | 133968.5 | 100 | 103.46 | 3.5 | bb | 2024-02-06
25 | 25 | 20240206_HepCl_025 | CC17 | Standard | 3.79 | 90058.109 | 2840044 |  | 90058.109 | 100 | 69.55 | -30.4 | bb | 2024-02-06
26 | 26 | 20240206_HepCl_026 | Blank | Blank | 3.77 | 14.134 | 366 |  | 14.134 | 100 | 0.01 | -100 | bb | 2024-02-06
27 | 27 | 20240206_HepCl_027 | Blank | Blank | 3.68 | 12.725 | 262 |  | 12.725 | 100 | 0.01 | -100 | bb | 2024-02-06
28 | 28 | 20240206_HepCl_028 | CCQC A | QC | 3.79 | 144180.359 | 4730226 |  | 144180.359 | 100 | 111.35 | 11.3 | bb | 2024-02-06
29 | 29 | 20240206_HepCl_029 | CCQC B | QC | 3.79 | 177202.609 | 5804371 |  | 177202.609 | 100 | 136.85 | 36.9 | bb | 2024-02-06
30 | 30 | 20240206_HepCl_030 | CCQC C | QC | 3.78 | 138244.625 | 4529511 |  | 138244.625 | 100 | 106.76 | 6.8 | bb | 2024-02-06
31 | 31 | 20240206_HepCl_031 | CCQC D | QC | 3.79 | 142977.109 | 4675255 |  | 142977.109 | 100 | 110.42 | 10.4 | bb | 2024-02-06
32 | 32 | 20240206_HepCl_032 | Blank | Blank | 3.8 | 66.006 | 444 |  | 66.006 | 100 | 0.05 | -99.9 | bb | 2024-02-06
33 | 33 | 20240206_HepCl_033 | Blank | Blank | 3.69 | 0.973 | 51 |  | 0.973 | 100 | 0 | -100 | bb | 2024-02-06
34 | 34 | 20240206_HepCl_034 | CC04 | Standard | 3.78 | 136284.063 | 4536227 |  | 136284.063 | 100 | 105.25 | 5.3 | bb | 2024-02-06
35 | 35 | 20240206_HepCl_035 | CC05 | Standard | 3.78 | 135748.172 | 4522542 |  | 135748.172 | 100 | 104.84 | 4.8 | bb | 2024-02-06
36 | 36 | 20240206_HepCl_036 | CC06 | Standard | 3.78 | 130111.891 | 4285330 |  | 130111.891 | 100 | 100.48 | 0.5 | bb | 2024-02-06
37 | 37 | 20240206_HepCl_037 | CC13 | Standard | 3.79 | 135464.547 | 4472643 |  | 135464.547 | 100 | 104.62 | 4.6 | bb | 2024-02-06
38 | 38 | 20240206_HepCl_038 | Blank | Blank | 3.84 | 25.41 | 427 |  | 25.41 | 100 | 0.02 | -100 | bb | 2024-02-06
39 | 39 | 20240206_HepCl_039 | CC04 | Standard | 3.79 | 149225.266 | 4943138 |  | 149225.266 | 100 | 115.24 | 15.2 | bb | 2024-02-06
40 | 40 | 20240206_HepCl_040 | CC05 | Standard | 3.79 | 136307.844 | 4519393 |  | 136307.844 | 100 | 105.27 | 5.3 | bb | 2024-02-06
41 | 41 | 20240206_HepCl_041 | CC06 | Standard | 3.79 | 122278.68 | 3964777 |  | 122278.68 | 100 | 94.43 | -5.6 | bb | 2024-02-06
42 | 42 | 20240206_HepCl_042 | CC13 | Standard | 3.78 | 116004.852 | 3779762 |  | 116004.852 | 100 | 89.59 | -10.4 | bb | 2024-02-06
43 | 43 | 20240206_HepCl_043 | Blank | Blank | 3.96 | 12.584 | 294 |  | 12.584 | 100 | 0.01 | -100 | bb | 2024-02-06
44 | 44 | 20240206_HepCl_044 | CC04 | Standard | 3.78 | 146873.625 | 4790017 |  | 146873.625 | 100 | 113.43 | 13.4 | bb | 2024-02-06
45 | 45 | 20240206_HepCl_045 | CC05 | Standard | 3.78 | 143390.891 | 4802327 |  | 143390.891 | 100 | 110.74 | 10.7 | bb | 2024-02-07
46 | 46 | 20240206_HepCl_046 | CC06 | Standard | 3.78 | 135189.078 | 4472953 |  | 135189.078 | 100 | 104.4 | 4.4 | bb | 2024-02-07
47 | 47 | 20240206_HepCl_047 | CC13 | Standard | 3.79 | 140587.422 | 4566574 |  | 140587.422 | 100 | 108.57 | 8.6 | bb | 2024-02-07
48 | 48 | 20240206_HepCl_048 | Blank | Blank | 3.83 | 13.116 | 293 |  | 13.116 | 100 | 0.01 | -100 | bb | 2024-02-07
49 | 49 | 20240206_HepCl_049 | Blank | Blank | 3.72 | 8.044 | 271 |  | 8.044 | 100 | 0.01 | -100 | bb | 2024-02-07
50 | 50 | 20240206_HepCl_050 | CC01 | Standard | 3.79 | 146270.547 | 4814792 |  | 146270.547 | 100 | 112.96 | 13 | bb | 2024-02-07
51 | 51 | 20240206_HepCl_051 | CC02 | Standard | 3.79 | 153555.141 | 5056891 |  | 153555.141 | 100 | 118.59 | 18.6 | bb | 2024-02-07
52 | 52 | 20240206_HepCl_052 | CC03 | Standard | 3.78 | 133897.094 | 4409536 |  | 133897.094 | 100 | 103.41 | 3.4 | bb | 2024-02-07
53 | 53 | 20240206_HepCl_053 | CC04 | Standard | 3.78 | 112335.555 | 3608396 |  | 112335.555 | 100 | 86.76 | -13.2 | bb | 2024-02-07
54 | 54 | 20240206_HepCl_054 | CC05 | Standard | 3.79 | 117635.977 | 3856072 |  | 117635.977 | 100 | 90.85 | -9.2 | bb | 2024-02-07
55 | 55 | 20240206_HepCl_055 | CC06 | Standard | 3.79 | 142075.469 | 4662613 |  | 142075.469 | 100 | 109.72 | 9.7 | bb | 2024-02-07
56 | 56 | 20240206_HepCl_056 | Blank | Blank | 3.86 | 61.414 | 389 |  | 61.414 | 100 | 0.05 | -100 | bb | 2024-02-07
57 | 57 | 20240206_HepCl_057 | CC07 | Standard | 3.78 | 117855.555 | 3814840 |  | 117855.555 | 100 | 91.02 | -9 | bb | 2024-02-07
58 | 58 | 20240206_HepCl_058 | CC08 | Standard | 3.78 | 113011.789 | 3641974 |  | 113011.789 | 100 | 87.28 | -12.7 | bb | 2024-02-07
59 | 59 | 20240206_HepCl_059 | CC09 | Standard | 3.79 | 118829.922 | 3856934 |  | 118829.922 | 100 | 91.77 | -8.2 | bb | 2024-02-07
60 | 60 | 20240206_HepCl_060 | CC10 | Standard | 3.78 | 152427.219 | 5080944 |  | 152427.219 | 100 | 117.72 | 17.7 | bb | 2024-02-07
61 | 61 | 20240206_HepCl_061 | CC11 | Standard | 3.79 | 134633.641 | 4370997 |  | 134633.641 | 100 | 103.98 | 4 | bb | 2024-02-07
62 | 62 | 20240206_HepCl_062 | CC12 | Standard | 3.79 | 130157.563 | 4261513 |  | 130157.563 | 100 | 100.52 | 0.5 | bb | 2024-02-07
63 | 63 | 20240206_HepCl_063 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-07
64 | 64 | 20240206_HepCl_064 | CC13 | Standard | 3.79 | 130902.141 | 4297562 |  | 130902.141 | 100 | 101.09 | 1.1 | bb | 2024-02-07
65 | 65 | 20240206_HepCl_065 | CC14 | Standard | 3.79 | 122410.234 | 3944801 |  | 122410.234 | 100 | 94.54 | -5.5 | bb | 2024-02-07
66 | 66 | 20240206_HepCl_066 | CC15 | Standard | 3.79 | 116367.969 | 3750585 |  | 116367.969 | 100 | 89.87 | -10.1 | bb | 2024-02-07
67 | 67 | 20240206_HepCl_067 | CC16 | Standard | 3.79 | 132073.734 | 4259799 |  | 132073.734 | 100 | 102 | 2 | bb | 2024-02-07
68 | 68 | 20240206_HepCl_068 | CC17 | Standard | 3.78 | 98631.766 | 3073554 |  | 98631.766 | 100 | 76.17 | -23.8 | bb | 2024-02-07
69 | 69 | 20240206_HepCl_069 | Blank | Blank | 3.89 | 70.645 | 511 |  | 70.645 | 100 | 0.05 | -99.9 | bb | 2024-02-07
70 | 70 | 20240206_HepCl_070 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-07
71 | 71 | 20240206_HepCl_071 | QC A | QC | 3.79 | 167490.781 | 5517542 |  | 167490.781 | 100 | 129.35 | 29.4 | bb | 2024-02-07
72 | 72 | 20240206_HepCl_072 | QC B | QC | 3.79 | 132315.688 | 4281620 |  | 132315.688 | 100 | 102.19 | 2.2 | bb | 2024-02-07
73 | 73 | 20240206_HepCl_073 | QC C | QC | 3.79 | 156897.484 | 5229336 |  | 156897.484 | 100 | 121.17 | 21.2 | bb | 2024-02-07
74 | 74 | 20240206_HepCl_074 | QC D | QC | 3.78 | 166386.391 | 5484595 |  | 166386.391 | 100 | 128.5 | 28.5 | bb | 2024-02-07
75 | 75 | 20240206_HepCl_075 | Blank | Blank | 3.79 | 79.723 | 480 |  | 79.723 | 100 | 0.06 | -99.9 | bb | 2024-02-07
76 | 76 | 20240206_HepCl_076 | Blank | Blank | 3.83 | 99.62 | 430 |  | 99.62 | 100 | 0.08 | -99.9 | bb | 2024-02-07
77 | 77 | 20240206_HepCl_077 | CC04 | Standard | 3.78 | 150515.109 | 5014824 |  | 150515.109 | 100 | 116.24 | 16.2 | bb | 2024-02-07
78 | 78 | 20240206_HepCl_078 | CC05 | Standard | 3.78 | 137131.031 | 4567322 |  | 137131.031 | 100 | 105.9 | 5.9 | bb | 2024-02-07
79 | 79 | 20240206_HepCl_079 | CC06 | Standard | 3.78 | 129202.703 | 4249507 |  | 129202.703 | 100 | 99.78 | -0.2 | bb | 2024-02-07
80 | 80 | 20240206_HepCl_080 | CC13 | Standard | 3.78 | 138754.016 | 4544742 |  | 138754.016 | 100 | 107.16 | 7.2 | bb | 2024-02-07
81 | 81 | 20240206_HepCl_081 | Blank | Blank | 3.8 | 14.27 | 304 |  | 14.27 | 100 | 0.01 | -100 | bb | 2024-02-07
82 | 82 | 20240206_HepCl_082 | CC04 | Standard | 3.78 | 129994.75 | 4272615 |  | 129994.75 | 100 | 100.39 | 0.4 | bb | 2024-02-07
83 | 83 | 20240206_HepCl_083 | CC05 | Standard | 3.78 | 128168.617 | 4134631 |  | 128168.617 | 100 | 98.98 | -1 | bb | 2024-02-07
84 | 84 | 20240206_HepCl_084 | CC06 | Standard | 3.78 | 135884.641 | 4425236 |  | 135884.641 | 100 | 104.94 | 4.9 | bb | 2024-02-07
85 | 85 | 20240206_HepCl_085 | CC13 | Standard | 3.78 | 142142.094 | 4720642 |  | 142142.094 | 100 | 109.77 | 9.8 | bb | 2024-02-07
86 | 86 | 20240206_HepCl_086 | Blank | Blank | 3.75 | 10.153 | 372 |  | 10.153 | 100 | 0.01 | -100 | bb | 2024-02-07
87 | 87 | 20240206_HepCl_087 | CC04 | Standard | 3.79 | 157955.078 | 5231489 |  | 157955.078 | 100 | 121.99 | 22 | bb | 2024-02-07
88 | 88 | 20240206_HepCl_088 | CC05 | Standard | 3.79 | 141630.219 | 4617237 |  | 141630.219 | 100 | 109.38 | 9.4 | bb | 2024-02-07
89 | 89 | 20240206_HepCl_089 | CC06 | Standard | 3.79 | 131731.141 | 4373728 |  | 131731.141 | 100 | 101.73 | 1.7 | bb | 2024-02-07
90 | 90 | 20240206_HepCl_090 | CC13 | Standard | 3.78 | 125548.438 | 4030725 |  | 125548.438 | 100 | 96.96 | -3 | bb | 2024-02-07
91 | 91 | 20240206_HepCl_091 | Blank | Blank | 3.78 | 8.037 | 275 |  | 8.037 | 100 | 0.01 | -100 | bb | 2024-02-07
92 | 92 | 20240206_HepCl_092 | Blank | Blank | 3.83 | 40.381 | 290 |  | 40.381 | 100 | 0.03 | -100 | bb | 2024-02-07
93 | 93 | 20240206_HepCl_093 | Active T0 A | Analyte | 3.79 | 483745.031 | 15100567 |  | 483745.031 | 100 | 373.59 | 273.6 | bb | 2024-02-07
94 | 94 | 20240206_HepCl_094 | Active T0 B | Analyte | 3.78 | 459075.188 | 14633644 |  | 459075.188 | 100 | 354.54 | 254.5 | bb | 2024-02-07
95 | 95 | 20240206_HepCl_095 | Active T0 C | Analyte | 3.79 | 528794.375 | 17161112 |  | 528794.375 | 100 | 408.38 | 308.4 | bb | 2024-02-07
96 | 96 | 20240206_HepCl_096 | Active T15 A | Analyte | 3.79 | 551724.25 | 17942246 |  | 551724.25 | 100 | 426.09 | 326.1 | bb | 2024-02-07
97 | 97 | 20240206_HepCl_097 | Active T15 B | Analyte | 3.79 | 583719.563 | 19246358 |  | 583719.563 | 100 | 450.8 | 350.8 | bb | 2024-02-07
98 | 98 | 20240206_HepCl_098 | Active T15 C | Analyte | 3.78 | 514480.375 | 16539207 |  | 514480.375 | 100 | 397.33 | 297.3 | bb | 2024-02-07
99 | 99 | 20240206_HepCl_099 | Blank | Blank | 3.66 | 5.725 | 214 |  | 5.725 | 100 | 0 | -100 | bb | 2024-02-07
100 | 100 | 20240206_HepCl_100 | Active T30 A | Analyte | 3.78 | 487161.469 | 16129338 |  | 487161.469 | 100 | 376.23 | 276.2 | bb | 2024-02-07
101 | 101 | 20240206_HepCl_101 | Active T30 B | Analyte | 3.78 | 572075.313 | 19031012 |  | 572075.313 | 100 | 441.81 | 341.8 | bb | 2024-02-07
102 | 102 | 20240206_HepCl_102 | Active T30 C | Analyte | 3.79 | 481638.969 | 15452899 |  | 481638.969 | 100 | 371.96 | 272 | bb | 2024-02-07
103 | 103 | 20240206_HepCl_103 | Active T60 A | Analyte | 3.79 | 530218.813 | 17092980 |  | 530218.813 | 100 | 409.48 | 309.5 | bb | 2024-02-07
104 | 104 | 20240206_HepCl_104 | Active T60 B | Analyte | 3.78 | 598683.313 | 19767818 |  | 598683.313 | 100 | 462.36 | 362.4 | bb | 2024-02-07
105 | 105 | 20240206_HepCl_105 | Active T60 C | Analyte | 3.79 | 600279.688 | 19534292 |  | 600279.688 | 100 | 463.59 | 363.6 | bb | 2024-02-07
106 | 106 | 20240206_HepCl_106 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-07
107 | 107 | 20240206_HepCl_107 | CC12 | Standard | 3.78 | 116307.844 | 3760237 |  | 116307.844 | 100 | 89.82 | -10.2 | bb | 2024-02-07
108 | 108 | 20240206_HepCl_108 | Blank | Blank | 3.84 | 12.753 | 406 |  | 12.753 | 100 | 0.01 | -100 | bb | 2024-02-07
109 | 109 | 20240206_HepCl_109 | Active T120 A | Analyte | 3.78 | 748670.625 | 24375130 |  | 748670.625 | 100 | 578.19 | 478.2 | bb | 2024-02-07
110 | 110 | 20240206_HepCl_110 | Active T120 B | Analyte | 3.78 | 611171.313 | 19885652 |  | 611171.313 | 100 | 472 | 372 | bb | 2024-02-07
111 | 111 | 20240206_HepCl_111 | Active T120 C | Analyte | 3.78 | 630458.313 | 20525438 |  | 630458.313 | 100 | 486.9 | 386.9 | bb | 2024-02-07
112 | 112 | 20240206_HepCl_112 | Active T240 A | Analyte | 3.78 | 530024 | 16970868 |  | 530024 | 100 | 409.33 | 309.3 | bb | 2024-02-07
113 | 113 | 20240206_HepCl_113 | Active T240 B | Analyte | 3.78 | 553846.313 | 17926820 |  | 553846.313 | 100 | 427.73 | 327.7 | bb | 2024-02-07
114 | 114 | 20240206_HepCl_114 | Active T240 C | Analyte | 3.78 | 517922.813 | 16951756 |  | 517922.813 | 100 | 399.99 | 300 | bb | 2024-02-07
115 | 115 | 20240206_HepCl_115 | Blank | Blank | 3.86 | 0.544 | 62 |  | 0.544 | 100 | 0 | -100 | bb | 2024-02-07
116 | 116 | 20240206_HepCl_116 | CC12 | Standard | 3.78 | 136432.594 | 4517728 |  | 136432.594 | 100 | 105.37 | 5.4 | bb | 2024-02-07
117 | 117 | 20240206_HepCl_117 | Blank | Blank | 3.78 | 1.556 | 170 |  | 1.556 | 100 | 0 | -100 | bb | 2024-02-07
118 | 118 | 20240206_HepCl_118 | Media T0 A | Analyte | 3.78 | 483696.719 | 15596621 |  | 483696.719 | 100 | 373.55 | 273.6 | bb | 2024-02-07
119 | 119 | 20240206_HepCl_119 | Media T0 B | Analyte | 3.78 | 554689.25 | 18180086 |  | 554689.25 | 100 | 428.38 | 328.4 | bb | 2024-02-07
120 | 120 | 20240206_HepCl_120 | Media T0 C | Analyte | 3.78 | 605034.563 | 20058984 |  | 605034.563 | 100 | 467.26 | 367.3 | bb | 2024-02-07
121 | 121 | 20240206_HepCl_121 | Media T15 A | Analyte | 3.79 | 436108.25 | 14171177 |  | 436108.25 | 100 | 336.8 | 236.8 | bb | 2024-02-07
122 | 122 | 20240206_HepCl_122 | Media T15 B | Analyte | 3.78 | 497355.656 | 16241408 |  | 497355.656 | 100 | 384.1 | 284.1 | bb | 2024-02-07
123 | 123 | 20240206_HepCl_123 | Media T15 C | Analyte | 3.79 | 458635.219 | 14618891 |  | 458635.219 | 100 | 354.2 | 254.2 | bb | 2024-02-07
124 | 124 | 20240206_HepCl_124 | Blank | Blank | 3.83 | 121.379 | 853 |  | 121.379 | 100 | 0.09 | -99.9 | bb | 2024-02-07
125 | 125 | 20240206_HepCl_125 | Media T30 A | Analyte | 3.79 | 549553.5 | 18123876 |  | 549553.5 | 100 | 424.41 | 324.4 | bb | 2024-02-07
126 | 126 | 20240206_HepCl_126 | Media T30 B | Analyte | 3.79 | 542602.625 | 17445314 |  | 542602.625 | 100 | 419.05 | 319 | bb | 2024-02-07
127 | 127 | 20240206_HepCl_127 | Media T30 C | Analyte | 3.79 | 572704.688 | 18743662 |  | 572704.688 | 100 | 442.29 | 342.3 | bb | 2024-02-07
128 | 128 | 20240206_HepCl_128 | Media T60 A | Analyte | 3.78 | 517056.781 | 16824996 |  | 517056.781 | 100 | 399.32 | 299.3 | bb | 2024-02-07
129 | 129 | 20240206_HepCl_129 | Media T60 B | Analyte | 3.78 | 545332.438 | 17639016 |  | 545332.438 | 100 | 421.15 | 321.2 | bb | 2024-02-07
130 | 130 | 20240206_HepCl_130 | Media T60 C | Analyte | 3.79 | 564425.188 | 17941832 |  | 564425.188 | 100 | 435.9 | 335.9 | bb | 2024-02-07
131 | 131 | 20240206_HepCl_131 | Blank | Blank | 3.78 | 9.231 | 356 |  | 9.231 | 100 | 0.01 | -100 | bb | 2024-02-07
132 | 132 | 20240206_HepCl_132 | CC12 | Standard | 3.78 | 125093.547 | 4092294 |  | 125093.547 | 100 | 96.61 | -3.4 | bb | 2024-02-07
133 | 133 | 20240206_HepCl_133 | Blank | Blank | 3.86 | 101.073 | 831 |  | 101.073 | 100 | 0.08 | -99.9 | bb | 2024-02-07
134 | 134 | 20240206_HepCl_134 | Media T120 A | Analyte | 3.79 | 646892.625 | 21053214 |  | 646892.625 | 100 | 499.59 | 399.6 | bb | 2024-02-07
135 | 135 | 20240206_HepCl_135 | Media T120 B | Analyte | 3.78 | 527762.875 | 16439441 |  | 527762.875 | 100 | 407.58 | 307.6 | bb | 2024-02-07
136 | 136 | 20240206_HepCl_136 | Media T120 C | Analyte | 3.79 | 616629.188 | 19846670 |  | 616629.188 | 100 | 476.22 | 376.2 | bb | 2024-02-07
137 | 137 | 20240206_HepCl_137 | Media T240 A | Analyte | 3.79 | 554051.313 | 17937966 |  | 554051.313 | 100 | 427.89 | 327.9 | bb | 2024-02-07
138 | 138 | 20240206_HepCl_138 | Media T240 B | Analyte | 3.78 | 507204.156 | 16314303 |  | 507204.156 | 100 | 391.71 | 291.7 | bb | 2024-02-07
139 | 139 | 20240206_HepCl_139 | Media T240 C | Analyte | 3.78 | 631669 | 20682628 |  | 631669 | 100 | 487.83 | 387.8 | bb | 2024-02-07
140 | 140 | 20240206_HepCl_140 | Blank | Blank | 3.86 | 16.476 | 597 |  | 16.476 | 100 | 0.01 | -100 | bb | 2024-02-07
141 | 141 | 20240206_HepCl_141 | CC12 | Standard | 3.79 | 142639.484 | 4738849 |  | 142639.484 | 100 | 110.16 | 10.2 | bb | 2024-02-07
142 | 142 | 20240206_HepCl_142 | Blank | Blank | 3.81 | 99.922 | 601 |  | 99.922 | 100 | 0.08 | -99.9 | bb | 2024-02-07
143 | 143 | 20240206_HepCl_143 | Inactive T0 A | Analyte | 3.78 | 525006.813 | 16898870 |  | 525006.813 | 100 | 405.46 | 305.5 | bb | 2024-02-07
144 | 144 | 20240206_HepCl_144 | Inactive T0 B | Analyte | 3.79 | 551036.875 | 18172490 |  | 551036.875 | 100 | 425.56 | 325.6 | bb | 2024-02-07
145 | 145 | 20240206_HepCl_145 | Inactive T0 C | Analyte | 3.79 | 475163.781 | 15171547 |  | 475163.781 | 100 | 366.96 | 267 | bb | 2024-02-07
146 | 146 | 20240206_HepCl_146 | Inactive T240 A | Analyte | 3.79 | 448450.656 | 14187672 |  | 448450.656 | 100 | 346.33 | 246.3 | bb | 2024-02-07
147 | 147 | 20240206_HepCl_147 | Inactive T240 B | Analyte | 3.79 | 490642.688 | 15946417 |  | 490642.688 | 100 | 378.92 | 278.9 | bb | 2024-02-07
148 | 148 | 20240206_HepCl_148 | Inactive T240 C | Analyte | 3.79 | 618428.688 | 20617388 |  | 618428.688 | 100 | 477.61 | 377.6 | bb | 2024-02-07
149 | 149 | 20240206_HepCl_149 | Blank | Blank | 3.84 | 104.156 | 560 |  | 104.156 | 100 | 0.08 | -99.9 | bb | 2024-02-07
150 | 150 | 20240206_HepCl_150 | Blank | Blank | 3.78 | 107.154 | 666 |  | 107.154 | 100 | 0.08 | -99.9 | bb | 2024-02-07
151 | 151 | 20240206_HepCl_151 | Blank | Blank | 3.87 | 108.747 | 553 |  | 108.747 | 100 | 0.08 | -99.9 | bb | 2024-02-07
Compound 5: PFOSA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20240206_HepCl_001 | Blank | Blank | 5.17 | 109.911 | 1041 | 42.896 | 256.227 |  | 24.98 |  | bb | 2024-02-06
2 | 2 | 20240206_HepCl_002 | Blank | Blank | 5.16 | 24.497 | 556 | 17.555 | 139.544 |  | 13.54 |  | bb | 2024-02-06
3 | 3 | 20240206_HepCl_003 | Matrix Blank | Blank |  |  |  | 16.108 |  |  |  |  | bb | 2024-02-06
4 | 4 | 20240206_HepCl_004 | Crash Blank | Blank | 5.17 | 64.09 | 314 | 25.148 | 254.851 |  | 24.84 |  | bb | 2024-02-06
5 | 5 | 20240206_HepCl_005 | Blank | Blank | 5.34 | 20.382 | 174 | 5.246 | 388.525 |  | 37.95 |  | bb | 2024-02-06
6 | 6 | 20240206_HepCl_006 | Blank | Blank | 5.05 | 3.732 | 137 | 4.082 | 91.426 |  | 8.82 |  | bb | 2024-02-06
7 | 7 | 20240206_HepCl_007 | CC01 | Standard | 5.16 | 1261.853 | 31157 | 28015.162 | 4.504 | 0.175 | 0.3 | 72.1 | bb | 2024-02-06
8 | 8 | 20240206_HepCl_008 | CC02 | Standard | 5.16 | 1216.324 | 29260 | 29313.949 | 4.149 | 0.281 | 0.27 | -5.3 | bb | 2024-02-06
9 | 9 | 20240206_HepCl_009 | CC03 | Standard | 5.17 | 1417.638 | 35882 | 29950.836 | 4.733 | 0.444 | 0.32 | -27.1 | bb | 2024-02-06
10 | 10 | 20240206_HepCl_010 | CC04 | Standard | 5.17 | 2318.271 | 56932 | 28339.805 | 8.18 | 0.713 | 0.66 | -7.1 | bb | 2024-02-06
11 | 11 | 20240206_HepCl_011 | CC05 | Standard | 5.17 | 3038.386 | 74959 | 25832.488 | 11.762 | 1.138 | 1.01 | -11 | bb | 2024-02-06
12 | 12 | 20240206_HepCl_012 | CC06 | Standard | 5.17 | 5194.874 | 130056 | 26238.695 | 19.799 | 1.825 | 1.8 | -1.3 | bb | 2024-02-06
13 | 13 | 20240206_HepCl_013 | Blank | Blank | 5.19 | 15.857 | 175 | 17.19 | 92.245 |  | 8.9 |  | bb | 2024-02-06
14 | 14 | 20240206_HepCl_014 | CC07 | Standard | 5.17 | 7565.646 | 184579 | 27822.641 | 27.192 | 2.919 | 2.53 | -13.5 | bb | 2024-02-06
15 | 15 | 20240206_HepCl_015 | CC08 | Standard | 5.17 | 12126.696 | 291778 | 27143.525 | 44.676 | 4.669 | 4.24 | -9.2 | bb | 2024-02-06
16 | 16 | 20240206_HepCl_016 | CC09 | Standard | 5.17 | 13285.316 | 318537 | 17637.496 | 75.324 | 7.45 | 7.24 | -2.8 | bb | 2024-02-06
17 | 17 | 20240206_HepCl_017 | CC10 | Standard | 5.17 | 32425.984 | 802510 | 26663.059 | 121.614 | 11.875 | 11.78 | -0.8 | bb | 2024-02-06
18 | 18 | 20240206_HepCl_018 | CC11 | Standard | 5.17 | 61522.891 | 1515608 | 27305.922 | 225.31 | 19.125 | 21.95 | 14.8 | bb | 2024-02-06
19 | 19 | 20240206_HepCl_019 | CC12 | Standard | 5.17 | 87626.531 | 2146827 | 27131.459 | 322.97 | 30.5 | 31.52 | 3.4 | bb | 2024-02-06
20 | 20 | 20240206_HepCl_020 | Blank | Blank | 5.17 | 23.177 | 373 | 7.525 | 308 |  | 30.05 |  | bb | 2024-02-06
21 | 21 | 20240206_HepCl_021 | CC13 | Standard | 5.16 | 141763.813 | 3414605 | 27395.945 | 517.463 | 48.75 | 50.59 | 3.8 | bb | 2024-02-06
22 | 22 | 20240206_HepCl_022 | CC14 | Standard | 5.17 | 205663.422 | 4891867 | 26067.559 | 788.963 | 78.125 | 77.21 | -1.2 | bb | 2024-02-06
23 | 23 | 20240206_HepCl_023 | CC15 | Standard | 5.16 | 309681.5 | 7288266 | 23888.445 | 1296.365 | 125 | 126.95 | 1.6 | bb | 2024-02-06
24 | 24 | 20240206_HepCl_024 | CC16 | Standard | 5.16 | 749450.375 | 18072108 | 45147.977 | 1659.987 | 200 | 162.6 | -18.7 | bb | 2024-02-06
25 | 25 | 20240206_HepCl_025 | CC17 | Standard | 5.17 | 1270558.75 | 29727522 | 35427.227 | 3586.391 | 325.625 | 351.46 | 7.9 | bb | 2024-02-06
26 | 26 | 20240206_HepCl_026 | Blank | Blank | 5.16 | 246.718 | 3720 | 6.313 | 3908.094 |  | 383 |  | bb | 2024-02-06
27 | 27 | 20240206_HepCl_027 | Blank | Blank | 5.16 | 106.462 | 944 | 36.603 | 290.856 |  | 28.37 |  | bb | 2024-02-06
28 | 28 | 20240206_HepCl_028 | CCQC A | QC | 5.17 | 3432.701 | 82059 | 37207.656 | 9.226 | 0.625 | 0.76 | 22.3 | bb | 2024-02-06
29 | 29 | 20240206_HepCl_029 | CCQC B | QC | 5.16 | 11178.361 | 261624 | 34276.102 | 32.613 | 2.506 | 3.06 | 22 | bb | 2024-02-06
30 | 30 | 20240206_HepCl_030 | CCQC C | QC | 5.16 | 21085.359 | 512127 | 27290.953 | 77.261 | 6.269 | 7.43 | 18.6 | bb | 2024-02-06
31 | 31 | 20240206_HepCl_031 | CCQC D | QC | 5.17 | 130586.773 | 3169813 | 38428.234 | 339.82 | 25 | 33.17 | 32.7 | bb | 2024-02-06
32 | 32 | 20240206_HepCl_032 | Blank | Blank | 5.16 | 21.903 | 545 | 10.67 | 205.276 |  | 19.98 |  | bb | 2024-02-06
33 | 33 | 20240206_HepCl_033 | Blank | Blank | 5.16 | 66.152 | 425 | 14.221 | 465.171 |  | 45.46 |  | bb | 2024-02-06
34 | 34 | 20240206_HepCl_034 | CC04 | Standard | 5.16 | 2133.638 | 50852 | 24535.77 | 8.696 | 0.713 | 0.71 | 0 | bb | 2024-02-06
35 | 35 | 20240206_HepCl_035 | CC05 | Standard | 5.16 | 2891.436 | 68256 | 23397 | 12.358 | 1.138 | 1.07 | -5.8 | bb | 2024-02-06
36 | 36 | 20240206_HepCl_036 | CC06 | Standard | 5.16 | 4643.119 | 112331 | 23039.303 | 20.153 | 1.825 | 1.84 | 0.6 | bb | 2024-02-06
37 | 37 | 20240206_HepCl_037 | CC13 | Standard | 5.16 | 133907.734 | 3295336 | 26302.416 | 509.108 | 48.75 | 49.77 | 2.1 | bb | 2024-02-06
38 | 38 | 20240206_HepCl_038 | Blank | Blank | 5.16 | 56.02 | 487 | 14.408 | 388.812 |  | 37.98 |  | bb | 2024-02-06
39 | 39 | 20240206_HepCl_039 | CC04 | Standard | 5.16 | 2245.164 | 54945 | 26556.42 | 8.454 | 0.713 | 0.69 | -3.4 | bb | 2024-02-06
40 | 40 | 20240206_HepCl_040 | CC05 | Standard | 5.16 | 2739.347 | 66732 | 23445.711 | 11.684 | 1.138 | 1.01 | -11.6 | bb | 2024-02-06
41 | 41 | 20240206_HepCl_041 | CC06 | Standard | 5.16 | 4788.455 | 119334 | 24264.488 | 19.734 | 1.825 | 1.79 | -1.7 | bb | 2024-02-06
42 | 42 | 20240206_HepCl_042 | CC13 | Standard | 5.16 | 134816.469 | 3379367 | 25713.629 | 524.3 | 48.75 | 51.26 | 5.1 | bb | 2024-02-06
43 | 43 | 20240206_HepCl_043 | Blank | Blank | 5.16 | 28.789 | 384 | 11.28 | 255.222 |  | 24.88 |  | bb | 2024-02-06
44 | 44 | 20240206_HepCl_044 | CC04 | Standard | 5.16 | 2043.82 | 51768 | 25185.447 | 8.115 | 0.713 | 0.66 | -8 | bb | 2024-02-06
45 | 45 | 20240206_HepCl_045 | CC05 | Standard | 5.16 | 2716.965 | 65917 | 22468.701 | 12.092 | 1.138 | 1.05 | -8.1 | bb | 2024-02-07
46 | 46 | 20240206_HepCl_046 | CC06 | Standard | 5.16 | 4706.088 | 116252 | 22817.016 | 20.625 | 1.825 | 1.88 | 3.1 | bb | 2024-02-07
47 | 47 | 20240206_HepCl_047 | CC13 | Standard | 5.17 | 126875.133 | 3033352 | 24279.836 | 522.554 | 48.75 | 51.09 | 4.8 | bb | 2024-02-07
48 | 48 | 20240206_HepCl_048 | Blank | Blank | 5.16 | 14.72 | 369 | 12.853 | 114.526 |  | 11.09 |  | bb | 2024-02-07
49 | 49 | 20240206_HepCl_049 | Blank | Blank | 5.09 | 5.489 | 174 | 2.151 | 255.184 |  | 24.88 |  | bb | 2024-02-07
50 | 50 | 20240206_HepCl_050 | CC01 | Standard | 5.17 | 1132.14 | 28317 | 24507.627 | 4.62 | 0.175 | 0.31 | 78.6 | bb | 2024-02-07
51 | 51 | 20240206_HepCl_051 | CC02 | Standard | 5.17 | 989.433 | 24612 | 23632.779 | 4.187 | 0.281 | 0.27 | -4 | bb | 2024-02-07
52 | 52 | 20240206_HepCl_052 | CC03 | Standard | 5.17 | 1100.143 | 26431 | 23528.35 | 4.676 | 0.444 | 0.32 | -28.3 | bb | 2024-02-07
53 | 53 | 20240206_HepCl_053 | CC04 | Standard | 5.17 | 1953.565 | 45869 | 23608.605 | 8.275 | 0.713 | 0.67 | -5.8 | bb | 2024-02-07
54 | 54 | 20240206_HepCl_054 | CC05 | Standard | 5.17 | 2599.149 | 62042 | 22096.361 | 11.763 | 1.138 | 1.01 | -11 | bb | 2024-02-07
55 | 55 | 20240206_HepCl_055 | CC06 | Standard | 5.17 | 4287.854 | 104165 | 21123.363 | 20.299 | 1.825 | 1.85 | 1.4 | bb | 2024-02-07
56 | 56 | 20240206_HepCl_056 | Blank | Blank | 5.18 | 9.435 | 176 | 22.688 | 41.586 |  | 3.94 |  | bb | 2024-02-07
57 | 57 | 20240206_HepCl_057 | CC07 | Standard | 5.17 | 5888.268 | 141624 | 22068.023 | 26.682 | 2.919 | 2.48 | -15.2 | bb | 2024-02-07
58 | 58 | 20240206_HepCl_058 | CC08 | Standard | 5.17 | 9765.783 | 237113 | 22590.502 | 43.23 | 4.669 | 4.1 | -12.2 | bb | 2024-02-07
59 | 59 | 20240206_HepCl_059 | CC09 | Standard | 5.17 | 10607.894 | 255942 | 14474.795 | 73.285 | 7.45 | 7.04 | -5.4 | bb | 2024-02-07
60 | 60 | 20240206_HepCl_060 | CC10 | Standard | 5.17 | 27365.682 | 683191 | 21487.777 | 127.355 | 11.875 | 12.35 | 4 | bb | 2024-02-07
61 | 61 | 20240206_HepCl_061 | CC11 | Standard | 5.17 | 52900.586 | 1319640 | 22796.053 | 232.06 | 19.125 | 22.61 | 18.2 | bb | 2024-02-07
62 | 62 | 20240206_HepCl_062 | CC12 | Standard | 5.18 | 67234.352 | 1651028 | 20643.158 | 325.698 | 30.5 | 31.79 | 4.2 | bb | 2024-02-07
63 | 63 | 20240206_HepCl_063 | Blank | Blank | 5.18 | 16.734 | 442 | 0.246 | 6802.439 |  | 666.75 |  | bb | 2024-02-07
64 | 64 | 20240206_HepCl_064 | CC13 | Standard | 5.18 | 116863.461 | 2871577 | 21099.6 | 553.866 | 48.75 | 54.16 | 11.1 | bb | 2024-02-07
65 | 65 | 20240206_HepCl_065 | CC14 | Standard | 5.18 | 160091.453 | 3896586 | 19912.988 | 803.955 | 78.125 | 78.68 | 0.7 | bb | 2024-02-07
66 | 66 | 20240206_HepCl_066 | CC15 | Standard | 5.18 | 236571.453 | 5677405 | 18836.156 | 1255.943 | 125 | 122.99 | -1.6 | bb | 2024-02-07
67 | 67 | 20240206_HepCl_067 | CC16 | Standard | 5.18 | 634856.063 | 14937564 | 37734.23 | 1682.441 | 200 | 164.8 | -17.6 | bb | 2024-02-07
68 | 68 | 20240206_HepCl_068 | CC17 | Standard | 5.18 | 1065313.625 | 25544994 | 31416.877 | 3390.896 | 325.625 | 332.29 | 2 | bb | 2024-02-07
69 | 69 | 20240206_HepCl_069 | Blank | Blank | 5.18 | 180.023 | 3059 | 41.587 | 432.883 |  | 42.3 |  | bb | 2024-02-07
70 | 70 | 20240206_HepCl_070 | Blank | Blank | 5.18 | 118.174 | 1030 | 18.371 | 643.264 |  | 62.92 |  | bb | 2024-02-07
71 | 71 | 20240206_HepCl_071 | QC A | QC | 5.18 | 2971.267 | 73145 | 32357.793 | 9.183 | 0.625 | 0.76 | 21.6 | bb | 2024-02-07
72 | 72 | 20240206_HepCl_072 | QC B | QC | 5.18 | 9997.105 | 241535 | 30769.189 | 32.491 | 2.506 | 3.04 | 21.5 | bb | 2024-02-07
73 | 73 | 20240206_HepCl_073 | QC C | QC | 5.18 | 15434.472 | 363637 | 18361.693 | 84.058 | 6.269 | 8.1 | 29.2 | bb | 2024-02-07
74 | 74 | 20240206_HepCl_074 | QC D | QC | 5.18 | 113470.266 | 2831497 | 33615.02 | 337.558 | 25 | 32.95 | 31.8 | bb | 2024-02-07
75 | 75 | 20240206_HepCl_075 | Blank | Blank | 5.19 | 9.891 | 317 | 11.861 | 83.391 |  | 8.03 |  | bb | 2024-02-07
76 | 76 | 20240206_HepCl_076 | Blank | Blank | 5.17 | 32.667 | 298 | 12.105 | 269.864 |  | 26.32 |  | bb | 2024-02-07
77 | 77 | 20240206_HepCl_077 | CC04 | Standard | 5.18 | 1691.695 | 39901 | 20045.84 | 8.439 | 0.713 | 0.69 | -3.6 | bb | 2024-02-07
78 | 78 | 20240206_HepCl_078 | CC05 | Standard | 5.18 | 2125.962 | 50142 | 17964.746 | 11.834 | 1.138 | 1.02 | -10.3 | bb | 2024-02-07
79 | 79 | 20240206_HepCl_079 | CC06 | Standard | 5.19 | 3641.675 | 90102 | 18868.227 | 19.301 | 1.825 | 1.75 | -4 | bb | 2024-02-07
80 | 80 | 20240206_HepCl_080 | CC13 | Standard | 5.19 | 104348.242 | 2556895 | 17471.68 | 597.242 | 48.75 | 58.41 | 19.8 | bb | 2024-02-07
81 | 81 | 20240206_HepCl_081 | Blank | Blank | 5.19 | 16.401 | 424 | 23.058 | 71.129 |  | 6.83 |  | bb | 2024-02-07
82 | 82 | 20240206_HepCl_082 | CC04 | Standard | 5.19 | 1678.866 | 41565 | 20375.537 | 8.24 | 0.713 | 0.67 | -6.3 | bb | 2024-02-07
83 | 83 | 20240206_HepCl_083 | CC05 | Standard | 5.19 | 1908.929 | 45048 | 16582.211 | 11.512 | 1.138 | 0.99 | -13.1 | bb | 2024-02-07
84 | 84 | 20240206_HepCl_084 | CC06 | Standard | 5.19 | 3367.372 | 79964 | 16807.939 | 20.034 | 1.825 | 1.82 | -0.1 | bb | 2024-02-07
85 | 85 | 20240206_HepCl_085 | CC13 | Standard | 5.19 | 97814.875 | 2321435 | 19143.832 | 510.947 | 48.75 | 49.95 | 2.5 | bb | 2024-02-07
86 | 86 | 20240206_HepCl_086 | Blank | Blank | 5.19 | 56.057 | 655 | 6.612 | 847.807 |  | 82.98 |  | bb | 2024-02-07
87 | 87 | 20240206_HepCl_087 | CC04 | Standard | 5.19 | 1601.046 | 38244 | 19531.541 | 8.197 | 0.713 | 0.66 | -6.9 | bb | 2024-02-07
88 | 88 | 20240206_HepCl_088 | CC05 | Standard | 5.2 | 2017.531 | 48224 | 17208.271 | 11.724 | 1.138 | 1.01 | -11.3 | bb | 2024-02-07
89 | 89 | 20240206_HepCl_089 | CC06 | Standard | 5.2 | 3380.709 | 82340 | 16733.023 | 20.204 | 1.825 | 1.84 | 0.8 | bb | 2024-02-07
90 | 90 | 20240206_HepCl_090 | CC13 | Standard | 5.2 | 95417.172 | 2251454 | 17418.449 | 547.794 | 48.75 | 53.56 | 9.9 | bb | 2024-02-07
91 | 91 | 20240206_HepCl_091 | Blank | Blank | 5.2 | 60.385 | 584 | 33.975 | 177.734 |  | 17.28 |  | bb | 2024-02-07
92 | 92 | 20240206_HepCl_092 | Blank | Blank | 5.2 | 4.567 | 152 | 27.48 | 16.619 |  | 1.49 |  | bb | 2024-02-07
93 | 93 | 20240206_HepCl_093 | Active T0 A | Analyte | 5.2 | 422083.375 | 10185595 | 76879.266 | 549.021 |  | 53.68 |  | bb | 2024-02-07
94 | 94 | 20240206_HepCl_094 | Active T0 B | Analyte | 5.2 | 440704.281 | 10793917 | 74541.352 | 591.221 |  | 57.82 |  | bb | 2024-02-07
95 | 95 | 20240206_HepCl_095 | Active T0 C | Analyte | 5.2 | 380902.75 | 9238476 | 73446.242 | 518.614 |  | 50.7 |  | bb | 2024-02-07
96 | 96 | 20240206_HepCl_096 | Active T15 A | Analyte | 5.2 | 341947.656 | 8135609 | 71795.5 | 476.28 |  | 46.55 |  | bb | 2024-02-07
97 | 97 | 20240206_HepCl_097 | Active T15 B | Analyte | 5.2 | 323529.906 | 7916220 | 67235.516 | 481.189 |  | 47.03 |  | bb | 2024-02-07
98 | 98 | 20240206_HepCl_098 | Active T15 C | Analyte | 5.2 | 307835.531 | 7611555 | 59377.094 | 518.442 |  | 50.69 |  | bb | 2024-02-07
99 | 99 | 20240206_HepCl_099 | Blank | Blank | 5.2 | 64.883 | 768 | 11.866 | 546.798 |  | 53.47 |  | bb | 2024-02-07
100 | 100 | 20240206_HepCl_100 | Active T30 A | Analyte | 5.2 | 296008.063 | 7315709 | 59786.223 | 495.111 |  | 48.4 |  | bb | 2024-02-07
101 | 101 | 20240206_HepCl_101 | Active T30 B | Analyte | 5.2 | 257113.063 | 6303558 | 60600.289 | 424.277 |  | 41.45 |  | bb | 2024-02-07
102 | 102 | 20240206_HepCl_102 | Active T30 C | Analyte | 5.2 | 250668 | 6286585 | 64703.051 | 387.413 |  | 37.84 |  | bb | 2024-02-07
103 | 103 | 20240206_HepCl_103 | Active T60 A | Analyte | 5.2 | 270694.688 | 6658858 | 71666.102 | 377.716 |  | 36.89 |  | bb | 2024-02-07
104 | 104 | 20240206_HepCl_104 | Active T60 B | Analyte | 5.2 | 272790.063 | 6748657 | 68193.617 | 400.023 |  | 39.08 |  | bb | 2024-02-07
105 | 105 | 20240206_HepCl_105 | Active T60 C | Analyte | 5.2 | 268667.219 | 6495771 | 66932.391 | 401.401 |  | 39.21 |  | bb | 2024-02-07
106 | 106 | 20240206_HepCl_106 | Blank | Blank | 5.2 | 63.298 | 885 | 14.916 | 424.363 |  | 41.46 |  | bb | 2024-02-07
107 | 107 | 20240206_HepCl_107 | CC12 | Standard | 5.2 | 45782.332 | 1091899 | 14109.446 | 324.48 | 30.5 | 31.67 | 3.8 | bb | 2024-02-07
108 | 108 | 20240206_HepCl_108 | Blank | Blank | 5.19 | 34.506 | 428 | 24.083 | 143.279 |  | 13.91 |  | bb | 2024-02-07
109 | 109 | 20240206_HepCl_109 | Active T120 A | Analyte | 5.2 | 211335.281 | 7100464 | 67789.539 | 311.752 |  | 30.42 |  | bb | 2024-02-07
110 | 110 | 20240206_HepCl_110 | Active T120 B | Analyte | 5.2 | 213620.984 | 6756371 | 70369.023 | 303.572 |  | 29.62 |  | bb | 2024-02-07
111 | 111 | 20240206_HepCl_111 | Active T120 C | Analyte | 5.2 | 221547.641 | 6922652 | 72949.461 | 303.7 |  | 29.63 |  | bb | 2024-02-07
112 | 112 | 20240206_HepCl_112 | Active T240 A | Analyte | 5.2 | 124771.945 | 5112219 | 69046.656 | 180.707 |  | 17.58 |  | bb | 2024-02-07
113 | 113 | 20240206_HepCl_113 | Active T240 B | Analyte | 5.2 | 116021.547 | 4662540 | 69858.797 | 166.08 |  | 16.14 |  | bb | 2024-02-07
114 | 114 | 20240206_HepCl_114 | Active T240 C | Analyte | 5.2 | 146123.844 | 5632798 | 72495.445 | 201.563 |  | 19.62 |  | bb | 2024-02-07
115 | 115 | 20240206_HepCl_115 | Blank | Blank | 5.2 | 54.096 | 305 | 8.215 | 658.503 |  | 64.42 |  | bb | 2024-02-07
116 | 116 | 20240206_HepCl_116 | CC12 | Standard | 5.2 | 40330.492 | 982018 | 12328.883 | 327.122 | 30.5 | 31.93 | 4.7 | bb | 2024-02-07
117 | 117 | 20240206_HepCl_117 | Blank | Blank | 5.2 | 35.661 | 504 | 6.233 | 572.132 |  | 55.95 |  | bb | 2024-02-07
118 | 118 | 20240206_HepCl_118 | Media T0 A | Analyte | 5.2 | 338138.125 | 8250365 | 62798.77 | 538.447 |  | 52.65 |  | bb | 2024-02-07
119 | 119 | 20240206_HepCl_119 | Media T0 B | Analyte | 5.2 | 322311.656 | 7793794 | 60950.219 | 528.811 |  | 51.7 |  | bb | 2024-02-07
120 | 120 | 20240206_HepCl_120 | Media T0 C | Analyte | 5.2 | 336100.875 | 8085665 | 61018.313 | 550.82 |  | 53.86 |  | bb | 2024-02-07
121 | 121 | 20240206_HepCl_121 | Media T15 A | Analyte | 5.21 | 357696.781 | 8728060 | 53439.043 | 669.355 |  | 65.48 |  | bb | 2024-02-07
122 | 122 | 20240206_HepCl_122 | Media T15 B | Analyte | 5.21 | 331523.156 | 8058599 | 55735.617 | 594.814 |  | 58.17 |  | bb | 2024-02-07
123 | 123 | 20240206_HepCl_123 | Media T15 C | Analyte | 5.21 | 245906.297 | 5946874 | 58680.23 | 419.062 |  | 40.94 |  | bb | 2024-02-07
124 | 124 | 20240206_HepCl_124 | Blank | Blank | 5.21 | 96.891 | 1086 | 39.24 | 246.919 |  | 24.07 |  | bb | 2024-02-07
125 | 125 | 20240206_HepCl_125 | Media T30 A | Analyte | 5.21 | 193000.906 | 4462446 | 48207.898 | 400.351 |  | 39.11 |  | bb | 2024-02-07
126 | 126 | 20240206_HepCl_126 | Media T30 B | Analyte | 5.21 | 188581.766 | 4653999 | 54628.371 | 345.208 |  | 33.7 |  | bb | 2024-02-07
127 | 127 | 20240206_HepCl_127 | Media T30 C | Analyte | 5.21 | 186013.891 | 4614897 | 57590.84 | 322.992 |  | 31.52 |  | bb | 2024-02-07
128 | 128 | 20240206_HepCl_128 | Media T60 A | Analyte | 5.21 | 342255.781 | 8209658 | 62483.848 | 547.751 |  | 53.56 |  | bb | 2024-02-07
129 | 129 | 20240206_HepCl_129 | Media T60 B | Analyte | 5.21 | 303342.594 | 7345818 | 54148.367 | 560.206 |  | 54.78 |  | bb | 2024-02-07
130 | 130 | 20240206_HepCl_130 | Media T60 C | Analyte | 5.21 | 281791.906 | 6948520 | 55963.391 | 503.529 |  | 49.22 |  | bb | 2024-02-07
131 | 131 | 20240206_HepCl_131 | Blank | Blank | 5.22 | 92.428 | 994 | 30.685 | 301.216 |  | 29.39 |  | bb | 2024-02-07
132 | 132 | 20240206_HepCl_132 | CC12 | Standard | 5.21 | 38098.008 | 909254 | 11270.036 | 338.047 | 30.5 | 33 | 8.2 | bb | 2024-02-07
133 | 133 | 20240206_HepCl_133 | Blank | Blank | 5.21 | 16.904 | 414 | 4.249 | 397.835 |  | 38.86 |  | bb | 2024-02-07
134 | 134 | 20240206_HepCl_134 | Media T120 A | Analyte | 5.21 | 324197.281 | 8006442 | 69880.008 | 463.934 |  | 45.34 |  | bb | 2024-02-07
135 | 135 | 20240206_HepCl_135 | Media T120 B | Analyte | 5.21 | 308013.719 | 7359169 | 67209.852 | 458.287 |  | 44.79 |  | bb | 2024-02-07
136 | 136 | 20240206_HepCl_136 | Media T120 C | Analyte | 5.21 | 296929.656 | 7207137 | 69773.344 | 425.563 |  | 41.58 |  | bb | 2024-02-07
137 | 137 | 20240206_HepCl_137 | Media T240 A | Analyte | 5.22 | 322356.75 | 7625099 | 66457.195 | 485.059 |  | 47.41 |  | bb | 2024-02-07
138 | 138 | 20240206_HepCl_138 | Media T240 B | Analyte | 5.21 | 305873.188 | 7348542 | 61993.867 | 493.393 |  | 48.23 |  | bb | 2024-02-07
139 | 139 | 20240206_HepCl_139 | Media T240 C | Analyte | 5.21 | 310387.938 | 7491211 | 67426.344 | 460.336 |  | 44.99 |  | bb | 2024-02-07
140 | 140 | 20240206_HepCl_140 | Blank | Blank | 5.22 | 52.428 | 1022 | 31.17 | 168.2 |  | 16.35 |  | bb | 2024-02-07
141 | 141 | 20240206_HepCl_141 | CC12 | Standard | 5.22 | 33885.363 | 798328 | 10386.876 | 326.232 | 30.5 | 31.84 | 4.4 | bb | 2024-02-07
142 | 142 | 20240206_HepCl_142 | Blank | Blank | 5.2 | 21.74 | 410 | 4.231 | 513.827 |  | 50.23 |  | bb | 2024-02-07
143 | 143 | 20240206_HepCl_143 | Inactive T0 A | Analyte | 5.21 | 342590.844 | 8160378 | 76440.156 | 448.182 |  | 43.8 |  | bb | 2024-02-07
144 | 144 | 20240206_HepCl_144 | Inactive T0 B | Analyte | 5.22 | 312290.625 | 7718516 | 78526.922 | 397.686 |  | 38.85 |  | bb | 2024-02-07
145 | 145 | 20240206_HepCl_145 | Inactive T0 C | Analyte | 5.22 | 284597.656 | 7016211 | 74402.867 | 382.509 |  | 37.36 |  | bb | 2024-02-07
146 | 146 | 20240206_HepCl_146 | Inactive T240 A | Analyte | 5.22 | 267962 | 6510039 | 69915.516 | 383.265 |  | 37.43 |  | bb | 2024-02-07
147 | 147 | 20240206_HepCl_147 | Inactive T240 B | Analyte | 5.22 | 266627.656 | 6609631 | 76102.867 | 350.352 |  | 34.21 |  | bb | 2024-02-07
148 | 148 | 20240206_HepCl_148 | Inactive T240 C | Analyte | 5.22 | 237742.609 | 5822674 | 72467.32 | 328.069 |  | 32.02 |  | bb | 2024-02-07
149 | 149 | 20240206_HepCl_149 | Blank | Blank | 5.21 | 101.697 | 1052 | 12.896 | 788.593 |  | 77.17 |  | bb | 2024-02-07
150 | 150 | 20240206_HepCl_150 | Blank | Blank | 5.23 | 27.967 | 472 | 23.086 | 121.143 |  | 11.74 |  | bb | 2024-02-07
151 | 151 | 20240206_HepCl_151 | Blank | Blank | 5.16 | 1.968 | 118 | 62.664 | 3.141 |  | 0.17 |  | bb | 2024-02-07
Compound 6: M8FOSA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20240206_HepCl_001 | Blank | Blank | 5.17 | 42.896 | 169 |  | 42.896 | 100 | 0.19 | -99.8 | bb | 2024-02-06
2 | 2 | 20240206_HepCl_002 | Blank | Blank | 5.28 | 17.555 | 100 |  | 17.555 | 100 | 0.08 | -99.9 | bb | 2024-02-06
3 | 3 | 20240206_HepCl_003 | Matrix Blank | Blank | 5.34 | 16.108 | 131 |  | 16.108 | 100 | 0.07 | -99.9 | bb | 2024-02-06
4 | 4 | 20240206_HepCl_004 | Crash Blank | Blank | 5.16 | 25.148 | 81 |  | 25.148 | 100 | 0.11 | -99.9 | bb | 2024-02-06
5 | 5 | 20240206_HepCl_005 | Blank | Blank | 5.16 | 5.246 | 76 |  | 5.246 | 100 | 0.02 | -100 | bb | 2024-02-06
6 | 6 | 20240206_HepCl_006 | Blank | Blank | 5.34 | 4.082 | 87 |  | 4.082 | 100 | 0.02 | -100 | bb | 2024-02-06
7 | 7 | 20240206_HepCl_007 | CC01 | Standard | 5.17 | 28015.162 | 764846 |  | 28015.162 | 100 | 121.54 | 21.5 | bb | 2024-02-06
8 | 8 | 20240206_HepCl_008 | CC02 | Standard | 5.17 | 29313.949 | 810743 |  | 29313.949 | 100 | 127.17 | 27.2 | bb | 2024-02-06
9 | 9 | 20240206_HepCl_009 | CC03 | Standard | 5.17 | 29950.836 | 830589 |  | 29950.836 | 100 | 129.94 | 29.9 | bb | 2024-02-06
10 | 10 | 20240206_HepCl_010 | CC04 | Standard | 5.17 | 28339.805 | 775903 |  | 28339.805 | 100 | 122.95 | 22.9 | bb | 2024-02-06
11 | 11 | 20240206_HepCl_011 | CC05 | Standard | 5.17 | 25832.488 | 724632 |  | 25832.488 | 100 | 112.07 | 12.1 | bb | 2024-02-06
12 | 12 | 20240206_HepCl_012 | CC06 | Standard | 5.17 | 26238.695 | 724166 |  | 26238.695 | 100 | 113.83 | 13.8 | bb | 2024-02-06
13 | 13 | 20240206_HepCl_013 | Blank | Blank | 5.16 | 17.19 | 145 |  | 17.19 | 100 | 0.07 | -99.9 | bb | 2024-02-06
14 | 14 | 20240206_HepCl_014 | CC07 | Standard | 5.17 | 27822.641 | 750505 |  | 27822.641 | 100 | 120.7 | 20.7 | bb | 2024-02-06
15 | 15 | 20240206_HepCl_015 | CC08 | Standard | 5.17 | 27143.525 | 718182 |  | 27143.525 | 100 | 117.76 | 17.8 | bb | 2024-02-06
16 | 16 | 20240206_HepCl_016 | CC09 | Standard | 5.17 | 17637.496 | 472292 |  | 17637.496 | 100 | 76.52 | -23.5 | bb | 2024-02-06
17 | 17 | 20240206_HepCl_017 | CC10 | Standard | 5.17 | 26663.059 | 733176 |  | 26663.059 | 100 | 115.67 | 15.7 | bb | 2024-02-06
18 | 18 | 20240206_HepCl_018 | CC11 | Standard | 5.17 | 27305.922 | 747337 |  | 27305.922 | 100 | 118.46 | 18.5 | bb | 2024-02-06
19 | 19 | 20240206_HepCl_019 | CC12 | Standard | 5.17 | 27131.459 | 745667 |  | 27131.459 | 100 | 117.7 | 17.7 | bb | 2024-02-06
20 | 20 | 20240206_HepCl_020 | Blank | Blank | 5.17 | 7.525 | 265 |  | 7.525 | 100 | 0.03 | -100 | bb | 2024-02-06
21 | 21 | 20240206_HepCl_021 | CC13 | Standard | 5.17 | 27395.945 | 753648 |  | 27395.945 | 100 | 118.85 | 18.9 | bb | 2024-02-06
22 | 22 | 20240206_HepCl_022 | CC14 | Standard | 5.17 | 26067.559 | 716640 |  | 26067.559 | 100 | 113.09 | 13.1 | bb | 2024-02-06
23 | 23 | 20240206_HepCl_023 | CC15 | Standard | 5.17 | 23888.445 | 647062 |  | 23888.445 | 100 | 103.64 | 3.6 | bb | 2024-02-06
24 | 24 | 20240206_HepCl_024 | CC16 | Standard | 5.17 | 45147.977 | 1246717 |  | 45147.977 | 100 | 195.87 | 95.9 | bb | 2024-02-06
25 | 25 | 20240206_HepCl_025 | CC17 | Standard | 5.17 | 35427.227 | 936213 |  | 35427.227 | 100 | 153.69 | 53.7 | bb | 2024-02-06
26 | 26 | 20240206_HepCl_026 | Blank | Blank | 5.17 | 6.313 | 212 |  | 6.313 | 100 | 0.03 | -100 | bb | 2024-02-06
27 | 27 | 20240206_HepCl_027 | Blank | Blank | 5.08 | 36.603 | 157 |  | 36.603 | 100 | 0.16 | -99.8 | bb | 2024-02-06
28 | 28 | 20240206_HepCl_028 | CCQC A | QC | 5.17 | 37207.656 | 1028814 |  | 37207.656 | 100 | 161.42 | 61.4 | bb | 2024-02-06
29 | 29 | 20240206_HepCl_029 | CCQC B | QC | 5.17 | 34276.102 | 919331 |  | 34276.102 | 100 | 148.7 | 48.7 | bb | 2024-02-06
30 | 30 | 20240206_HepCl_030 | CCQC C | QC | 5.17 | 27290.953 | 756072 |  | 27290.953 | 100 | 118.4 | 18.4 | bb | 2024-02-06
31 | 31 | 20240206_HepCl_031 | CCQC D | QC | 5.17 | 38428.234 | 1065680 |  | 38428.234 | 100 | 166.71 | 66.7 | bb | 2024-02-06
32 | 32 | 20240206_HepCl_032 | Blank | Blank | 5.17 | 10.67 | 222 |  | 10.67 | 100 | 0.05 | -100 | bb | 2024-02-06
33 | 33 | 20240206_HepCl_033 | Blank | Blank | 5.05 | 14.221 | 106 |  | 14.221 | 100 | 0.06 | -99.9 | bb | 2024-02-06
34 | 34 | 20240206_HepCl_034 | CC04 | Standard | 5.16 | 24535.77 | 665024 |  | 24535.77 | 100 | 106.44 | 6.4 | bb | 2024-02-06
35 | 35 | 20240206_HepCl_035 | CC05 | Standard | 5.16 | 23397 | 644247 |  | 23397 | 100 | 101.5 | 1.5 | bb | 2024-02-06
36 | 36 | 20240206_HepCl_036 | CC06 | Standard | 5.16 | 23039.303 | 625832 |  | 23039.303 | 100 | 99.95 | 0 | bb | 2024-02-06
37 | 37 | 20240206_HepCl_037 | CC13 | Standard | 5.16 | 26302.416 | 722371 |  | 26302.416 | 100 | 114.11 | 14.1 | bb | 2024-02-06
38 | 38 | 20240206_HepCl_038 | Blank | Blank | 5.17 | 14.408 | 148 |  | 14.408 | 100 | 0.06 | -99.9 | bb | 2024-02-06
39 | 39 | 20240206_HepCl_039 | CC04 | Standard | 5.17 | 26556.42 | 744566 |  | 26556.42 | 100 | 115.21 | 15.2 | bb | 2024-02-06
40 | 40 | 20240206_HepCl_040 | CC05 | Standard | 5.17 | 23445.711 | 645545 |  | 23445.711 | 100 | 101.71 | 1.7 | bb | 2024-02-06
41 | 41 | 20240206_HepCl_041 | CC06 | Standard | 5.17 | 24264.488 | 682297 |  | 24264.488 | 100 | 105.27 | 5.3 | bb | 2024-02-06
42 | 42 | 20240206_HepCl_042 | CC13 | Standard | 5.16 | 25713.629 | 715980 |  | 25713.629 | 100 | 111.55 | 11.6 | bb | 2024-02-06
43 | 43 | 20240206_HepCl_043 | Blank | Blank | 5.17 | 11.28 | 138 |  | 11.28 | 100 | 0.05 | -100 | bb | 2024-02-06
44 | 44 | 20240206_HepCl_044 | CC04 | Standard | 5.17 | 25185.447 | 693300 |  | 25185.447 | 100 | 109.26 | 9.3 | bb | 2024-02-06
45 | 45 | 20240206_HepCl_045 | CC05 | Standard | 5.16 | 22468.701 | 616974 |  | 22468.701 | 100 | 97.48 | -2.5 | bb | 2024-02-07
46 | 46 | 20240206_HepCl_046 | CC06 | Standard | 5.17 | 22817.016 | 625104 |  | 22817.016 | 100 | 98.99 | -1 | bb | 2024-02-07
47 | 47 | 20240206_HepCl_047 | CC13 | Standard | 5.17 | 24279.836 | 662569 |  | 24279.836 | 100 | 105.33 | 5.3 | bb | 2024-02-07
48 | 48 | 20240206_HepCl_048 | Blank | Blank | 5.17 | 12.853 | 206 |  | 12.853 | 100 | 0.06 | -99.9 | bb | 2024-02-07
49 | 49 | 20240206_HepCl_049 | Blank | Blank | 5.07 | 2.151 | 58 |  | 2.151 | 100 | 0.01 | -100 | bb | 2024-02-07
50 | 50 | 20240206_HepCl_050 | CC01 | Standard | 5.17 | 24507.627 | 671672 |  | 24507.627 | 100 | 106.32 | 6.3 | bb | 2024-02-07
51 | 51 | 20240206_HepCl_051 | CC02 | Standard | 5.17 | 23632.779 | 648234 |  | 23632.779 | 100 | 102.53 | 2.5 | bb | 2024-02-07
52 | 52 | 20240206_HepCl_052 | CC03 | Standard | 5.17 | 23528.35 | 650772 |  | 23528.35 | 100 | 102.07 | 2.1 | bb | 2024-02-07
53 | 53 | 20240206_HepCl_053 | CC04 | Standard | 5.17 | 23608.605 | 645217 |  | 23608.605 | 100 | 102.42 | 2.4 | bb | 2024-02-07
54 | 54 | 20240206_HepCl_054 | CC05 | Standard | 5.17 | 22096.361 | 611518 |  | 22096.361 | 100 | 95.86 | -4.1 | bb | 2024-02-07
55 | 55 | 20240206_HepCl_055 | CC06 | Standard | 5.17 | 21123.363 | 582288 |  | 21123.363 | 100 | 91.64 | -8.4 | bb | 2024-02-07
56 | 56 | 20240206_HepCl_056 | Blank | Blank | 5.17 | 22.688 | 234 |  | 22.688 | 100 | 0.1 | -99.9 | bb | 2024-02-07
57 | 57 | 20240206_HepCl_057 | CC07 | Standard | 5.17 | 22068.023 | 603914 |  | 22068.023 | 100 | 95.74 | -4.3 | bb | 2024-02-07
58 | 58 | 20240206_HepCl_058 | CC08 | Standard | 5.17 | 22590.502 | 634597 |  | 22590.502 | 100 | 98 | -2 | bb | 2024-02-07
59 | 59 | 20240206_HepCl_059 | CC09 | Standard | 5.17 | 14474.795 | 396937 |  | 14474.795 | 100 | 62.8 | -37.2 | bb | 2024-02-07
60 | 60 | 20240206_HepCl_060 | CC10 | Standard | 5.17 | 21487.777 | 594106 |  | 21487.777 | 100 | 93.22 | -6.8 | bb | 2024-02-07
61 | 61 | 20240206_HepCl_061 | CC11 | Standard | 5.18 | 22796.053 | 633107 |  | 22796.053 | 100 | 98.9 | -1.1 | bb | 2024-02-07
62 | 62 | 20240206_HepCl_062 | CC12 | Standard | 5.18 | 20643.158 | 562092 |  | 20643.158 | 100 | 89.56 | -10.4 | bb | 2024-02-07
63 | 63 | 20240206_HepCl_063 | Blank | Blank | 5.16 | 0.246 | 22 |  | 0.246 | 100 | 0 | -100 | bb | 2024-02-07
64 | 64 | 20240206_HepCl_064 | CC13 | Standard | 5.18 | 21099.6 | 583234 |  | 21099.6 | 100 | 91.54 | -8.5 | bb | 2024-02-07
65 | 65 | 20240206_HepCl_065 | CC14 | Standard | 5.18 | 19912.988 | 553677 |  | 19912.988 | 100 | 86.39 | -13.6 | bb | 2024-02-07
66 | 66 | 20240206_HepCl_066 | CC15 | Standard | 5.18 | 18836.156 | 503821 |  | 18836.156 | 100 | 81.72 | -18.3 | bb | 2024-02-07
67 | 67 | 20240206_HepCl_067 | CC16 | Standard | 5.18 | 37734.23 | 1023848 |  | 37734.23 | 100 | 163.7 | 63.7 | bb | 2024-02-07
68 | 68 | 20240206_HepCl_068 | CC17 | Standard | 5.18 | 31416.877 | 849521 |  | 31416.877 | 100 | 136.3 | 36.3 | bb | 2024-02-07
69 | 69 | 20240206_HepCl_069 | Blank | Blank | 5.19 | 41.587 | 250 |  | 41.587 | 100 | 0.18 | -99.8 | bb | 2024-02-07
70 | 70 | 20240206_HepCl_070 | Blank | Blank | 5.18 | 18.371 | 188 |  | 18.371 | 100 | 0.08 | -99.9 | bb | 2024-02-07
71 | 71 | 20240206_HepCl_071 | QC A | QC | 5.18 | 32357.793 | 885092 |  | 32357.793 | 100 | 140.38 | 40.4 | bb | 2024-02-07
72 | 72 | 20240206_HepCl_072 | QC B | QC | 5.18 | 30769.189 | 828347 |  | 30769.189 | 100 | 133.49 | 33.5 | bb | 2024-02-07
73 | 73 | 20240206_HepCl_073 | QC C | QC | 5.18 | 18361.693 | 481360 |  | 18361.693 | 100 | 79.66 | -20.3 | bb | 2024-02-07
74 | 74 | 20240206_HepCl_074 | QC D | QC | 5.18 | 33615.02 | 927726 |  | 33615.02 | 100 | 145.83 | 45.8 | bb | 2024-02-07
75 | 75 | 20240206_HepCl_075 | Blank | Blank | 5.18 | 11.861 | 221 |  | 11.861 | 100 | 0.05 | -99.9 | bb | 2024-02-07
76 | 76 | 20240206_HepCl_076 | Blank | Blank | 5.35 | 12.105 | 107 |  | 12.105 | 100 | 0.05 | -99.9 | bb | 2024-02-07
77 | 77 | 20240206_HepCl_077 | CC04 | Standard | 5.18 | 20045.84 | 558513 |  | 20045.84 | 100 | 86.96 | -13 | bb | 2024-02-07
78 | 78 | 20240206_HepCl_078 | CC05 | Standard | 5.18 | 17964.746 | 493692 |  | 17964.746 | 100 | 77.94 | -22.1 | bb | 2024-02-07
79 | 79 | 20240206_HepCl_079 | CC06 | Standard | 5.19 | 18868.227 | 529578 |  | 18868.227 | 100 | 81.86 | -18.1 | bb | 2024-02-07
80 | 80 | 20240206_HepCl_080 | CC13 | Standard | 5.19 | 17471.68 | 464999 |  | 17471.68 | 100 | 75.8 | -24.2 | bb | 2024-02-07
81 | 81 | 20240206_HepCl_081 | Blank | Blank | 5.18 | 23.058 | 176 |  | 23.058 | 100 | 0.1 | -99.9 | bb | 2024-02-07
82 | 82 | 20240206_HepCl_082 | CC04 | Standard | 5.19 | 20375.537 | 572269 |  | 20375.537 | 100 | 88.4 | -11.6 | bb | 2024-02-07
83 | 83 | 20240206_HepCl_083 | CC05 | Standard | 5.19 | 16582.211 | 447530 |  | 16582.211 | 100 | 71.94 | -28.1 | bb | 2024-02-07
84 | 84 | 20240206_HepCl_084 | CC06 | Standard | 5.19 | 16807.939 | 454619 |  | 16807.939 | 100 | 72.92 | -27.1 | bb | 2024-02-07
85 | 85 | 20240206_HepCl_085 | CC13 | Standard | 5.19 | 19143.832 | 518020 |  | 19143.832 | 100 | 83.05 | -16.9 | bb | 2024-02-07
86 | 86 | 20240206_HepCl_086 | Blank | Blank | 5.18 | 6.612 | 162 |  | 6.612 | 100 | 0.03 | -100 | bb | 2024-02-07
87 | 87 | 20240206_HepCl_087 | CC04 | Standard | 5.19 | 19531.541 | 545490 |  | 19531.541 | 100 | 84.73 | -15.3 | bb | 2024-02-07
88 | 88 | 20240206_HepCl_088 | CC05 | Standard | 5.19 | 17208.271 | 474056 |  | 17208.271 | 100 | 74.65 | -25.3 | bb | 2024-02-07
89 | 89 | 20240206_HepCl_089 | CC06 | Standard | 5.19 | 16733.023 | 463597 |  | 16733.023 | 100 | 72.59 | -27.4 | bb | 2024-02-07
90 | 90 | 20240206_HepCl_090 | CC13 | Standard | 5.19 | 17418.449 | 459441 |  | 17418.449 | 100 | 75.57 | -24.4 | bb | 2024-02-07
91 | 91 | 20240206_HepCl_091 | Blank | Blank | 5.19 | 33.975 | 209 |  | 33.975 | 100 | 0.15 | -99.9 | bb | 2024-02-07
92 | 92 | 20240206_HepCl_092 | Blank | Blank | 5.2 | 27.48 | 172 |  | 27.48 | 100 | 0.12 | -99.9 | bb | 2024-02-07
93 | 93 | 20240206_HepCl_093 | Active T0 A | Analyte | 5.19 | 76879.266 | 2077377 |  | 76879.266 | 100 | 333.53 | 233.5 | bb | 2024-02-07
94 | 94 | 20240206_HepCl_094 | Active T0 B | Analyte | 5.19 | 74541.352 | 2025422 |  | 74541.352 | 100 | 323.38 | 223.4 | bb | 2024-02-07
95 | 95 | 20240206_HepCl_095 | Active T0 C | Analyte | 5.19 | 73446.242 | 1989905 |  | 73446.242 | 100 | 318.63 | 218.6 | bb | 2024-02-07
96 | 96 | 20240206_HepCl_096 | Active T15 A | Analyte | 5.2 | 71795.5 | 1924504 |  | 71795.5 | 100 | 311.47 | 211.5 | bb | 2024-02-07
97 | 97 | 20240206_HepCl_097 | Active T15 B | Analyte | 5.19 | 67235.516 | 1770558 |  | 67235.516 | 100 | 291.69 | 191.7 | bb | 2024-02-07
98 | 98 | 20240206_HepCl_098 | Active T15 C | Analyte | 5.19 | 59377.094 | 1611156 |  | 59377.094 | 100 | 257.6 | 157.6 | bb | 2024-02-07
99 | 99 | 20240206_HepCl_099 | Blank | Blank | 5.2 | 11.866 | 286 |  | 11.866 | 100 | 0.05 | -99.9 | bb | 2024-02-07
100 | 100 | 20240206_HepCl_100 | Active T30 A | Analyte | 5.2 | 59786.223 | 1646308 |  | 59786.223 | 100 | 259.37 | 159.4 | bb | 2024-02-07
101 | 101 | 20240206_HepCl_101 | Active T30 B | Analyte | 5.2 | 60600.289 | 1649459 |  | 60600.289 | 100 | 262.9 | 162.9 | bb | 2024-02-07
102 | 102 | 20240206_HepCl_102 | Active T30 C | Analyte | 5.2 | 64703.051 | 1782430 |  | 64703.051 | 100 | 280.7 | 180.7 | bb | 2024-02-07
103 | 103 | 20240206_HepCl_103 | Active T60 A | Analyte | 5.2 | 71666.102 | 1965729 |  | 71666.102 | 100 | 310.91 | 210.9 | bb | 2024-02-07
104 | 104 | 20240206_HepCl_104 | Active T60 B | Analyte | 5.2 | 68193.617 | 1893662 |  | 68193.617 | 100 | 295.84 | 195.8 | bb | 2024-02-07
105 | 105 | 20240206_HepCl_105 | Active T60 C | Analyte | 5.2 | 66932.391 | 1831529 |  | 66932.391 | 100 | 290.37 | 190.4 | bb | 2024-02-07
106 | 106 | 20240206_HepCl_106 | Blank | Blank | 5.19 | 14.916 | 270 |  | 14.916 | 100 | 0.06 | -99.9 | bb | 2024-02-07
107 | 107 | 20240206_HepCl_107 | CC12 | Standard | 5.2 | 14109.446 | 384397 |  | 14109.446 | 100 | 61.21 | -38.8 | bb | 2024-02-07
108 | 108 | 20240206_HepCl_108 | Blank | Blank | 5.19 | 24.083 | 244 |  | 24.083 | 100 | 0.1 | -99.9 | bb | 2024-02-07
109 | 109 | 20240206_HepCl_109 | Active T120 A | Analyte | 5.2 | 67789.539 | 1909254 |  | 67789.539 | 100 | 294.09 | 194.1 | bb | 2024-02-07
110 | 110 | 20240206_HepCl_110 | Active T120 B | Analyte | 5.2 | 70369.023 | 1907301 |  | 70369.023 | 100 | 305.28 | 205.3 | bb | 2024-02-07
111 | 111 | 20240206_HepCl_111 | Active T120 C | Analyte | 5.2 | 72949.461 | 2025578 |  | 72949.461 | 100 | 316.48 | 216.5 | bb | 2024-02-07
112 | 112 | 20240206_HepCl_112 | Active T240 A | Analyte | 5.2 | 69046.656 | 1919022 |  | 69046.656 | 100 | 299.55 | 199.5 | bb | 2024-02-07
113 | 113 | 20240206_HepCl_113 | Active T240 B | Analyte | 5.2 | 69858.797 | 1869751 |  | 69858.797 | 100 | 303.07 | 203.1 | bb | 2024-02-07
114 | 114 | 20240206_HepCl_114 | Active T240 C | Analyte | 5.2 | 72495.445 | 2027736 |  | 72495.445 | 100 | 314.51 | 214.5 | bb | 2024-02-07
115 | 115 | 20240206_HepCl_115 | Blank | Blank | 5.2 | 8.215 | 206 |  | 8.215 | 100 | 0.04 | -100 | bb | 2024-02-07
116 | 116 | 20240206_HepCl_116 | CC12 | Standard | 5.2 | 12328.883 | 348261 |  | 12328.883 | 100 | 53.49 | -46.5 | bb | 2024-02-07
117 | 117 | 20240206_HepCl_117 | Blank | Blank | 5.19 | 6.233 | 137 |  | 6.233 | 100 | 0.03 | -100 | bb | 2024-02-07
118 | 118 | 20240206_HepCl_118 | Media T0 A | Analyte | 5.2 | 62798.77 | 1682330 |  | 62798.77 | 100 | 272.44 | 172.4 | bb | 2024-02-07
119 | 119 | 20240206_HepCl_119 | Media T0 B | Analyte | 5.2 | 60950.219 | 1672878 |  | 60950.219 | 100 | 264.42 | 164.4 | bb | 2024-02-07
120 | 120 | 20240206_HepCl_120 | Media T0 C | Analyte | 5.2 | 61018.313 | 1648730 |  | 61018.313 | 100 | 264.72 | 164.7 | bb | 2024-02-07
121 | 121 | 20240206_HepCl_121 | Media T15 A | Analyte | 5.2 | 53439.043 | 1492113 |  | 53439.043 | 100 | 231.83 | 131.8 | bb | 2024-02-07
122 | 122 | 20240206_HepCl_122 | Media T15 B | Analyte | 5.2 | 55735.617 | 1534875 |  | 55735.617 | 100 | 241.8 | 141.8 | bb | 2024-02-07
123 | 123 | 20240206_HepCl_123 | Media T15 C | Analyte | 5.21 | 58680.23 | 1601727 |  | 58680.23 | 100 | 254.57 | 154.6 | bb | 2024-02-07
124 | 124 | 20240206_HepCl_124 | Blank | Blank | 5.21 | 39.24 | 274 |  | 39.24 | 100 | 0.17 | -99.8 | bb | 2024-02-07
125 | 125 | 20240206_HepCl_125 | Media T30 A | Analyte | 5.21 | 48207.898 | 1287356 |  | 48207.898 | 100 | 209.14 | 109.1 | bb | 2024-02-07
126 | 126 | 20240206_HepCl_126 | Media T30 B | Analyte | 5.21 | 54628.371 | 1511146 |  | 54628.371 | 100 | 236.99 | 137 | bb | 2024-02-07
127 | 127 | 20240206_HepCl_127 | Media T30 C | Analyte | 5.21 | 57590.84 | 1610994 |  | 57590.84 | 100 | 249.85 | 149.8 | bb | 2024-02-07
128 | 128 | 20240206_HepCl_128 | Media T60 A | Analyte | 5.21 | 62483.848 | 1706733 |  | 62483.848 | 100 | 271.07 | 171.1 | bb | 2024-02-07
129 | 129 | 20240206_HepCl_129 | Media T60 B | Analyte | 5.21 | 54148.367 | 1483369 |  | 54148.367 | 100 | 234.91 | 134.9 | bb | 2024-02-07
130 | 130 | 20240206_HepCl_130 | Media T60 C | Analyte | 5.21 | 55963.391 | 1539156 |  | 55963.391 | 100 | 242.79 | 142.8 | bb | 2024-02-07
131 | 131 | 20240206_HepCl_131 | Blank | Blank | 5.2 | 30.685 | 278 |  | 30.685 | 100 | 0.13 | -99.9 | bb | 2024-02-07
132 | 132 | 20240206_HepCl_132 | CC12 | Standard | 5.21 | 11270.036 | 304960 |  | 11270.036 | 100 | 48.89 | -51.1 | bb | 2024-02-07
133 | 133 | 20240206_HepCl_133 | Blank | Blank | 5.21 | 4.249 | 113 |  | 4.249 | 100 | 0.02 | -100 | bb | 2024-02-07
134 | 134 | 20240206_HepCl_134 | Media T120 A | Analyte | 5.21 | 69880.008 | 1912559 |  | 69880.008 | 100 | 303.16 | 203.2 | bb | 2024-02-07
135 | 135 | 20240206_HepCl_135 | Media T120 B | Analyte | 5.21 | 67209.852 | 1777711 |  | 67209.852 | 100 | 291.58 | 191.6 | bb | 2024-02-07
136 | 136 | 20240206_HepCl_136 | Media T120 C | Analyte | 5.21 | 69773.344 | 1931253 |  | 69773.344 | 100 | 302.7 | 202.7 | bb | 2024-02-07
137 | 137 | 20240206_HepCl_137 | Media T240 A | Analyte | 5.21 | 66457.195 | 1793467 |  | 66457.195 | 100 | 288.31 | 188.3 | bb | 2024-02-07
138 | 138 | 20240206_HepCl_138 | Media T240 B | Analyte | 5.21 | 61993.867 | 1675997 |  | 61993.867 | 100 | 268.95 | 168.9 | bb | 2024-02-07
139 | 139 | 20240206_HepCl_139 | Media T240 C | Analyte | 5.21 | 67426.344 | 1841450 |  | 67426.344 | 100 | 292.52 | 192.5 | bb | 2024-02-07
140 | 140 | 20240206_HepCl_140 | Blank | Blank | 5.23 | 31.17 | 335 |  | 31.17 | 100 | 0.14 | -99.9 | bb | 2024-02-07
141 | 141 | 20240206_HepCl_141 | CC12 | Standard | 5.21 | 10386.876 | 280167 |  | 10386.876 | 100 | 45.06 | -54.9 | bb | 2024-02-07
142 | 142 | 20240206_HepCl_142 | Blank | Blank | 5.21 | 4.231 | 124 |  | 4.231 | 100 | 0.02 | -100 | bb | 2024-02-07
143 | 143 | 20240206_HepCl_143 | Inactive T0 A | Analyte | 5.21 | 76440.156 | 2082636 |  | 76440.156 | 100 | 331.62 | 231.6 | bb | 2024-02-07
144 | 144 | 20240206_HepCl_144 | Inactive T0 B | Analyte | 5.22 | 78526.922 | 2211083 |  | 78526.922 | 100 | 340.67 | 240.7 | bb | 2024-02-07
145 | 145 | 20240206_HepCl_145 | Inactive T0 C | Analyte | 5.22 | 74402.867 | 2043815 |  | 74402.867 | 100 | 322.78 | 222.8 | bb | 2024-02-07
146 | 146 | 20240206_HepCl_146 | Inactive T240 A | Analyte | 5.22 | 69915.516 | 1874804 |  | 69915.516 | 100 | 303.31 | 203.3 | bb | 2024-02-07
147 | 147 | 20240206_HepCl_147 | Inactive T240 B | Analyte | 5.22 | 76102.867 | 2106583 |  | 76102.867 | 100 | 330.16 | 230.2 | bb | 2024-02-07
148 | 148 | 20240206_HepCl_148 | Inactive T240 C | Analyte | 5.22 | 72467.32 | 2037564 |  | 72467.32 | 100 | 314.38 | 214.4 | bb | 2024-02-07
149 | 149 | 20240206_HepCl_149 | Blank | Blank | 5.22 | 12.896 | 235 |  | 12.896 | 100 | 0.06 | -99.9 | bb | 2024-02-07
150 | 150 | 20240206_HepCl_150 | Blank | Blank | 5.05 | 23.086 | 196 |  | 23.086 | 100 | 0.1 | -99.9 | bb | 2024-02-07
151 | 151 | 20240206_HepCl_151 | Blank | Blank | 5.21 | 62.664 | 263 |  | 62.664 | 100 | 0.27 | -99.7 | bb | 2024-02-07

## Raw Data-RefCmpd
Quantify Compound Subbary Report
Printed Fri Feb 09 11:09:04 2024
Compound 1: phenacetin
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20240206_HepClReRun_Ref_001 | Blank | Blank | [9 empty] | 2024-02-07
2 | 2 | 20240206_HepClReRun_Ref_002 | Blank | Blank |  |  |  | 11.192 | [5 empty] | 2024-02-07
3 | 3 | 20240206_HepClReRun_Ref_003 | Matrix Blank | Blank |  |  |  | 15.171 | [5 empty] | 2024-02-07
4 | 4 | 20240206_HepClReRun_Ref_004 | Crash Blank | Blank |  |  |  | 1.89 | [5 empty] | 2024-02-07
5 | 5 | 20240206_HepClReRun_Ref_005 | Blank | Blank | [9 empty] | 2024-02-07
6 | 6 | 20240206_HepClReRun_Ref_006 | Blank | Blank | [9 empty] | 2024-02-07
7 | 7 | 20240206_HepClReRun_Ref_007 | CC01 | Analyte | 3.8 | 1171.234 | 41010 | 99147.227 | 1.181 | 0.175 |  |  | bbI | 2024-02-07
8 | 8 | 20240206_HepClReRun_Ref_008 | CC02 | Analyte | 3.8 | 2217.497 | 81953 | 110029.922 | 2.015 | 0.281 |  |  | bbI | 2024-02-07
9 | 9 | 20240206_HepClReRun_Ref_009 | CC03 | Analyte | 3.8 | 3997.633 | 141888 | 97847.359 | 4.086 | 0.444 | 0.08 | -82.7 | bb | 2024-02-07
10 | 10 | 20240206_HepClReRun_Ref_010 | CC04 | Analyte | 3.8 | 3947.641 | 141588 | 56906.484 | 6.937 | 0.713 | 0.45 | -36.6 | bb | 2024-02-07
11 | 11 | 20240206_HepClReRun_Ref_011 | CC05 | Standard | 3.8 | 7222.345 | 258879 | 62519.641 | 11.552 | 1.138 | 1.06 | -6.9 | bb | 2024-02-07
12 | 12 | 20240206_HepClReRun_Ref_012 | CC06 | Standard | 3.8 | 12805.168 | 459773 | 88497.523 | 14.47 | 1.825 | 1.44 | -20.9 | bb | 2024-02-07
13 | 13 | 20240206_HepClReRun_Ref_013 | Blank | Blank |  |  |  | 14.119 | [5 empty] | 2024-02-07
14 | 14 | 20240206_HepClReRun_Ref_014 | CC07 | Standard | 3.8 | 16686.445 | 592203 | 76204.688 | 21.897 | 2.919 | 2.42 | -17.1 | bb | 2024-02-07
15 | 15 | 20240206_HepClReRun_Ref_015 | CC08 | Standard | 3.8 | 29356.393 | 1028975 | 76172.734 | 38.539 | 4.669 | 4.61 | -1.3 | bb | 2024-02-08
16 | 16 | 20240206_HepClReRun_Ref_016 | CC09 | Standard | 3.8 | 55292.457 | 1962972 | 79696.414 | 69.379 | 7.45 | 8.67 | 16.3 | bb | 2024-02-08
17 | 17 | 20240206_HepClReRun_Ref_017 | CC10 | Standard | 3.8 | 83447.602 | 2943994 | 81798.641 | 102.016 | 11.875 | 12.96 | 9.1 | bb | 2024-02-08
18 | 18 | 20240206_HepClReRun_Ref_018 | CC11 | Standard | 3.8 | 116151.75 | 4092842 | 72615.586 | 159.954 | 19.125 | 20.58 | 7.6 | bb | 2024-02-08
19 | 19 | 20240206_HepClReRun_Ref_019 | CC12 | Standard | 3.8 | 189893.594 | 6711653 | 76026 | 249.775 | 30.5 | 32.4 | 6.2 | bb | 2024-02-08
20 | 20 | 20240206_HepClReRun_Ref_020 | Blank | Blank |  |  |  | 11.739 | [5 empty] | 2024-02-08
21 | 21 | 20240206_HepClReRun_Ref_021 | CC13 | Standard | 3.8 | 188722.281 | 6669372 | 46317.648 | 407.452 | 48.75 | 53.14 | 9 | bb | 2024-02-08
22 | 22 | 20240206_HepClReRun_Ref_022 | CC14 | Standard | 3.8 | 622536.063 | 21863364 | 99196.797 | 627.577 | 78.125 | 82.1 | 5.1 | bb | 2024-02-08
23 | 23 | 20240206_HepClReRun_Ref_023 | CC15 | Standard | 3.8 | 797437.438 | 28023148 | 75629.461 | 1054.401 | 125 | 138.25 | 10.6 | bb | 2024-02-08
24 | 24 | 20240206_HepClReRun_Ref_024 | CC16 | Standard | 3.8 | 1173109.75 | 40990896 | 82667.789 | 1419.065 | 200 | 186.23 | -6.9 | bb | 2024-02-08
25 | 25 | 20240206_HepClReRun_Ref_025 | CC17 | Analyte | 3.8 | 1816943 | 63451664 | 103689.242 | 1752.297 | 325.625 | 230.07 | -29.3 | bb | 2024-02-08
26 | 26 | 20240206_HepClReRun_Ref_026 | Blank | Blank |  |  |  | 7.415 | [5 empty] | 2024-02-08
27 | 27 | 20240206_HepClReRun_Ref_027 | Blank | Blank |  |  |  | 18.043 | [5 empty] | 2024-02-08
28 | 28 | 20240206_HepClReRun_Ref_028 | CCQC A | QC | 3.8 | 6443.611 | 229278 | 87326.75 | 7.379 | 0.625 | 0.51 | -18.4 | bb | 2024-02-08
29 | 29 | 20240206_HepClReRun_Ref_029 | CCQC B | QC | 3.8 | 21407.742 | 760359 | 81890.242 | 26.142 | 2.506 | 2.98 | 18.8 | bb | 2024-02-08
30 | 30 | 20240206_HepClReRun_Ref_030 | CCQC C | QC | 3.8 | 60443.637 | 2140083 | 103744.516 | 58.262 | 6.269 | 7.2 | 14.9 | bb | 2024-02-08
31 | 31 | 20240206_HepClReRun_Ref_031 | CCQC D | QC | 3.8 | 205287.438 | 7265895 | 72114.469 | 284.669 | 25 | 36.99 | 48 | bb | 2024-02-08
32 | 32 | 20240206_HepClReRun_Ref_032 | Blank | Blank |  |  |  | 1.047 | [5 empty] | 2024-02-08
33 | 33 | 20240206_HepClReRun_Ref_033 | Blank | Blank |  |  |  | 12.055 | [5 empty] | 2024-02-08
34 | 34 | 20240206_HepClReRun_Ref_034 | CC04 | Analyte | 3.8 | 5296.986 | 190711 | 50823.348 | 10.422 | 0.713 | 0.91 | 27.8 | bb | 2024-02-08
35 | 35 | 20240206_HepClReRun_Ref_035 | CC05 | Standard | 3.8 | 9176.8 | 320384 | 81373.328 | 11.277 | 1.138 | 1.02 | -10.1 | bb | 2024-02-08
36 | 36 | 20240206_HepClReRun_Ref_036 | CC06 | Standard | 3.8 | 15669.537 | 564795 | 96763.203 | 16.194 | 1.825 | 1.67 | -8.5 | bb | 2024-02-08
37 | 37 | 20240206_HepClReRun_Ref_037 | CC13 | Analyte | 3.8 | 203285.75 | 7171843 | 39609.715 | 513.222 | 48.75 | 67.06 | 37.6 | bb | 2024-02-08
38 | 38 | 20240206_HepClReRun_Ref_038 | Blank | Blank | [9 empty] | 2024-02-08
39 | 39 | 20240206_HepClReRun_Ref_039 | CC04 | Analyte | 3.8 | 4708.708 | 168591 | 45127.563 | 10.434 | 0.713 | 0.91 | 28 | bb | 2024-02-08
40 | 40 | 20240206_HepClReRun_Ref_040 | CC05 | Standard | 3.8 | 8076.035 | 289767 | 66556.109 | 12.134 | 1.138 | 1.14 | -0.2 | bb | 2024-02-08
41 | 41 | 20240206_HepClReRun_Ref_041 | CC06 | Standard | 3.8 | 14739.754 | 527682 | 86041.914 | 17.131 | 1.825 | 1.79 | -1.7 | bb | 2024-02-08
42 | 42 | 20240206_HepClReRun_Ref_042 | CC13 | Standard | 3.8 | 195398.391 | 7026393 | 43471.086 | 449.49 | 48.75 | 58.67 | 20.4 | bb | 2024-02-08
43 | 43 | 20240206_HepClReRun_Ref_043 | Blank | Blank |  |  |  | 9.778 | [5 empty] | 2024-02-08
44 | 44 | 20240206_HepClReRun_Ref_044 | CC04 | Analyte | 3.8 | 4851.796 | 178003 | 50215.895 | 9.662 | 0.713 | 0.81 | 13.8 | bb | 2024-02-08
45 | 45 | 20240206_HepClReRun_Ref_045 | CC05 | Standard | 3.8 | 8571.604 | 301937 | 75406.516 | 11.367 | 1.138 | 1.03 | -9 | bb | 2024-02-08
46 | 46 | 20240206_HepClReRun_Ref_046 | CC06 | Standard | 3.8 | 16642.393 | 599708 | 89999.438 | 18.492 | 1.825 | 1.97 | 8.1 | bb | 2024-02-08
47 | 47 | 20240206_HepClReRun_Ref_047 | CC13 | Analyte | 3.8 | 209094.203 | 7393666 | 38000.164 | 550.246 | 48.75 | 71.93 | 47.5 | bb | 2024-02-08
48 | 48 | 20240206_HepClReRun_Ref_048 | Blank | Blank |  |  |  | 16.982 | [5 empty] | 2024-02-08
49 | 49 | 20240206_HepClReRun_Ref_049 | Blank | Blank | [9 empty] | 2024-02-08
50 | 50 | 20240206_HepClReRun_Ref_050 | CC01 | Analyte | 3.8 | 1288.76 | 44314 | 91640.219 | 1.406 | 0.175 |  |  | bbI | 2024-02-08
51 | 51 | 20240206_HepClReRun_Ref_051 | CC02 | Analyte | 3.8 | 2885.166 | 107134 | 118718.242 | 2.43 | 0.281 |  |  | bbI | 2024-02-08
52 | 52 | 20240206_HepClReRun_Ref_052 | CC03 | Analyte | 3.8 | 4851.163 | 177036 | 104661.281 | 4.635 | 0.444 | 0.15 | -66.4 | bb | 2024-02-08
53 | 53 | 20240206_HepClReRun_Ref_053 | CC04 | Analyte | 3.8 | 5083.494 | 184147 | 51603.223 | 9.851 | 0.713 | 0.84 | 17.3 | bb | 2024-02-08
54 | 54 | 20240206_HepClReRun_Ref_054 | CC05 | Standard | 3.8 | 8625.105 | 303739 | 76128.5 | 11.33 | 1.138 | 1.03 | -9.5 | bb | 2024-02-08
55 | 55 | 20240206_HepClReRun_Ref_055 | CC06 | Standard | 3.8 | 17092.779 | 607234 | 86843.813 | 19.682 | 1.825 | 2.13 | 16.6 | bb | 2024-02-08
56 | 56 | 20240206_HepClReRun_Ref_056 | Blank | Blank | [9 empty] | 2024-02-08
57 | 57 | 20240206_HepClReRun_Ref_057 | CC07 | Standard | 3.8 | 19873.146 | 703580 | 83129.859 | 23.906 | 2.919 | 2.68 | -8 | bb | 2024-02-08
58 | 58 | 20240206_HepClReRun_Ref_058 | CC08 | Standard | 3.8 | 34638.613 | 1236244 | 95656.969 | 36.211 | 4.669 | 4.3 | -7.8 | bb | 2024-02-08
59 | 59 | 20240206_HepClReRun_Ref_059 | CC09 | Standard | 3.8 | 64771.605 | 2311056 | 92235.273 | 70.224 | 7.45 | 8.78 | 17.8 | bb | 2024-02-08
60 | 60 | 20240206_HepClReRun_Ref_060 | CC10 | Standard | 3.8 | 104594.359 | 3677755 | 97453.102 | 107.328 | 11.875 | 13.66 | 15 | bb | 2024-02-08
61 | 61 | 20240206_HepClReRun_Ref_061 | CC11 | Standard | 3.8 | 148204.063 | 5226313 | 86978.148 | 170.392 | 19.125 | 21.96 | 14.8 | bb | 2024-02-08
62 | 62 | 20240206_HepClReRun_Ref_062 | CC12 | Standard | 3.8 | 224008.953 | 7929918 | 85374.516 | 262.384 | 30.5 | 34.06 | 11.7 | bb | 2024-02-08
63 | 63 | 20240206_HepClReRun_Ref_063 | Blank | Blank |  |  |  | 1.61 | [5 empty] | 2024-02-08
64 | 64 | 20240206_HepClReRun_Ref_064 | CC13 | Standard | 3.8 | 226042 | 8009873 | 50045.719 | 451.671 | 48.75 | 58.96 | 20.9 | bb | 2024-02-08
65 | 65 | 20240206_HepClReRun_Ref_065 | CC14 | Standard | 3.8 | 683015 | 23939198 | 116990.578 | 583.821 | 78.125 | 76.35 | -2.3 | bb | 2024-02-08
66 | 66 | 20240206_HepClReRun_Ref_066 | CC15 | Standard | 3.8 | 948660.063 | 33312734 | 93066.141 | 1019.34 | 125 | 133.64 | 6.9 | bb | 2024-02-08
67 | 67 | 20240206_HepClReRun_Ref_067 | CC16 | Standard | 3.8 | 1267215.75 | 44322560 | 96650.32 | 1311.135 | 200 | 172.03 | -14 | bb | 2024-02-08
68 | 68 | 20240206_HepClReRun_Ref_068 | CC17 | Analyte | 3.8 | 1956425.625 | 67798808 | 93195.711 | 2099.266 | 325.625 | 275.72 | -15.3 | bb | 2024-02-08
69 | 69 | 20240206_HepClReRun_Ref_069 | Blank | Blank | [9 empty] | 2024-02-08
70 | 70 | 20240206_HepClReRun_Ref_070 | Blank | Blank | [9 empty] | 2024-02-08
71 | 71 | 20240206_HepClReRun_Ref_071 | QC A | QC | 3.8 | 6273.182 | 229047 | 95330.953 | 6.58 | 0.625 | 0.41 | -35.2 | bb | 2024-02-08
72 | 72 | 20240206_HepClReRun_Ref_072 | QC B | QC | 3.8 | 22465.502 | 797937 | 79754.047 | 28.168 | 2.506 | 3.25 | 29.5 | bb | 2024-02-08
73 | 73 | 20240206_HepClReRun_Ref_073 | QC C | QC | 3.8 | 61658.652 | 2182974 | 116253.453 | 53.038 | 6.269 | 6.52 | 4 | bb | 2024-02-08
74 | 74 | 20240206_HepClReRun_Ref_074 | QC D | QC | 3.8 | 206486.453 | 7287286 | 62696.5 | 329.343 | 25 | 42.87 | 71.5 | bb | 2024-02-08
75 | 75 | 20240206_HepClReRun_Ref_075 | Blank | Blank |  |  |  | 9.576 | [5 empty] | 2024-02-08
76 | 76 | 20240206_HepClReRun_Ref_076 | Blank | Blank | [9 empty] | 2024-02-08
77 | 77 | 20240206_HepClReRun_Ref_077 | CC04 | Analyte | 3.8 | 5318.31 | 191308 | 55185.578 | 9.637 | 0.713 | 0.81 | 13.3 | bb | 2024-02-08
78 | 78 | 20240206_HepClReRun_Ref_078 | CC05 | Standard | 3.8 | 8629.438 | 304052 | 80799.367 | 10.68 | 1.138 | 0.94 | -17 | bb | 2024-02-08
79 | 79 | 20240206_HepClReRun_Ref_079 | CC06 | Standard | 3.8 | 17672.369 | 624694 | 93015.078 | 18.999 | 1.825 | 2.04 | 11.7 | bb | 2024-02-08
80 | 80 | 20240206_HepClReRun_Ref_080 | CC13 | Analyte | 3.8 | 215148.844 | 7610073 | 46678.375 | 460.918 | 48.75 | 60.18 | 23.4 | bb | 2024-02-08
81 | 81 | 20240206_HepClReRun_Ref_081 | Blank | Blank |  |  |  | 8.46 | [5 empty] | 2024-02-08
82 | 82 | 20240206_HepClReRun_Ref_082 | CC04 | Analyte | 3.8 | 5480.817 | 197308 | 57444.809 | 9.541 | 0.713 | 0.79 | 11.5 | bb | 2024-02-08
83 | 83 | 20240206_HepClReRun_Ref_083 | CC05 | Standard | 3.8 | 9571.905 | 336232 | 73697.18 | 12.988 | 1.138 | 1.25 | 9.7 | bb | 2024-02-08
84 | 84 | 20240206_HepClReRun_Ref_084 | CC06 | Standard | 3.8 | 16975.592 | 607750 | 102521.625 | 16.558 | 1.825 | 1.72 | -5.9 | bb | 2024-02-08
85 | 85 | 20240206_HepClReRun_Ref_085 | CC13 | Analyte | 3.8 | 208243.859 | 7354952 | 39008.566 | 533.841 | 48.75 | 69.77 | 43.1 | bb | 2024-02-08
86 | 86 | 20240206_HepClReRun_Ref_086 | Blank | Blank |  |  |  | 1.97 | [5 empty] | 2024-02-08
87 | 87 | 20240206_HepClReRun_Ref_087 | CC04 | Analyte | 3.8 | 4596.38 | 170385 | 46731.051 | 9.836 | 0.713 | 0.83 | 17 | bb | 2024-02-08
88 | 88 | 20240206_HepClReRun_Ref_088 | CC05 | Standard | 3.8 | 8446.699 | 304756 | 71098.961 | 11.88 | 1.138 | 1.1 | -3.1 | bb | 2024-02-08
89 | 89 | 20240206_HepClReRun_Ref_089 | CC06 | Standard | 3.8 | 16159.386 | 582782 | 93787.023 | 17.23 | 1.825 | 1.81 | -1 | bb | 2024-02-08
90 | 90 | 20240206_HepClReRun_Ref_090 | CC13 | Analyte | 3.8 | 202572.359 | 7073163 | 29047.742 | 697.377 | 48.75 | 91.29 | 87.3 | bb | 2024-02-08
91 | 91 | 20240206_HepClReRun_Ref_091 | Blank | Blank |  |  |  | 9.14 | [5 empty] | 2024-02-08
92 | 92 | 20240206_HepClReRun_Ref_092 | Blank | Blank |  |  |  | 0.86 | [5 empty] | 2024-02-08
93 | 93 | 20240206_HepClReRun_Ref_093 | Active T0 A | Analyte | 3.79 | 708850.375 | 24625148 | 159468.484 | 444.508 |  | 58.02 |  | bb | 2024-02-08
94 | 94 | 20240206_HepClReRun_Ref_094 | Active T0 B | Analyte | 3.79 | 620444.563 | 21533276 | 137772.844 | 450.339 |  | 58.79 |  | bb | 2024-02-08
95 | 95 | 20240206_HepClReRun_Ref_095 | Active T0 C | Analyte | 3.79 | 613315.25 | 21407144 | 122639.313 | 500.097 |  | 65.33 |  | bb | 2024-02-08
96 | 96 | 20240206_HepClReRun_Ref_096 | Active T15 A | Analyte | 3.79 | 779389.25 | 27266958 | 191937.047 | 406.065 |  | 52.96 |  | bb | 2024-02-08
97 | 97 | 20240206_HepClReRun_Ref_097 | Active T15 B | Analyte | 3.79 | 766846.688 | 26843304 | 270446.25 | 283.549 |  | 36.84 |  | bb | 2024-02-08
98 | 98 | 20240206_HepClReRun_Ref_098 | Active T15 C | Analyte | 3.79 | 671016.5 | 23370706 | 228889.453 | 293.162 |  | 38.11 |  | bb | 2024-02-08
99 | 99 | 20240206_HepClReRun_Ref_099 | Blank | Blank |  |  |  | 10.602 | [5 empty] | 2024-02-08
100 | 100 | 20240206_HepClReRun_Ref_100 | Active T30 A | Analyte | 3.79 | 489140.531 | 17043044 | 151167.922 | 323.574 |  | 42.11 |  | bb | 2024-02-08
101 | 101 | 20240206_HepClReRun_Ref_101 | Active T30 B | Analyte | 3.79 | 472730.625 | 16506397 | 137365.531 | 344.141 |  | 44.81 |  | bb | 2024-02-08
102 | 102 | 20240206_HepClReRun_Ref_102 | Active T30 C | Analyte | 3.79 | 468518.094 | 16473934 | 188103.422 | 249.075 |  | 32.31 |  | bb | 2024-02-08
103 | 103 | 20240206_HepClReRun_Ref_103 | Active T60 A | Analyte | 3.79 | 359900.875 | 12575928 | 140990.125 | 255.267 |  | 33.12 |  | bb | 2024-02-08
104 | 104 | 20240206_HepClReRun_Ref_104 | Active T60 B | Analyte | 3.79 | 435406.719 | 15333224 | 193763.313 | 224.711 |  | 29.1 |  | bb | 2024-02-08
105 | 105 | 20240206_HepClReRun_Ref_105 | Active T60 C | Analyte | 3.79 | 385295.563 | 13508390 | 134827.359 | 285.77 |  | 37.13 |  | bb | 2024-02-08
106 | 106 | 20240206_HepClReRun_Ref_106 | Blank | Blank | [9 empty] | 2024-02-08
107 | 107 | 20240206_HepClReRun_Ref_107 | CC12 | Standard | 3.8 | 212640.453 | 6274637 | 105078.148 | 202.364 | 30.5 | 26.16 | -14.2 | bb | 2024-02-08
108 | 108 | 20240206_HepClReRun_Ref_108 | Blank | Blank |  |  |  | 3.136 | [5 empty] | 2024-02-08
109 | 109 | 20240206_HepClReRun_Ref_109 | Active T120 A | Analyte | 3.8 | 218175.859 | 7586293 | 66314.43 | 329.002 |  | 42.82 |  | bb | 2024-02-08
110 | 110 | 20240206_HepClReRun_Ref_110 | Active T120 B | Analyte | 3.79 | 168940.172 | 5890616 | 70664.453 | 239.074 |  | 30.99 |  | bb | 2024-02-08
111 | 111 | 20240206_HepClReRun_Ref_111 | Active T120 C | Analyte | 3.8 | 218543.875 | 7734972 | 90734.82 | 240.86 |  | 31.23 |  | bb | 2024-02-08
112 | 112 | 20240206_HepClReRun_Ref_112 | Active T240 A | Analyte | 3.8 | 217568.984 | 7617007 | 150100.688 | 144.949 |  | 18.61 |  | bb | 2024-02-08
113 | 113 | 20240206_HepClReRun_Ref_113 | Active T240 B | Analyte | 3.8 | 180226.625 | 6264853 | 121538.523 | 148.288 |  | 19.05 |  | bb | 2024-02-08
114 | 114 | 20240206_HepClReRun_Ref_114 | Active T240 C | Analyte | 3.79 | 167913.188 | 5850443 | 141360.891 | 118.783 |  | 15.17 |  | bb | 2024-02-08
115 | 115 | 20240206_HepClReRun_Ref_115 | Blank | Blank |  |  |  | 7.786 | [5 empty] | 2024-02-08
116 | 116 | 20240206_HepClReRun_Ref_116 | CC12 | Standard | 3.79 | 172536.125 | 5998684 | 92012.313 | 187.514 | 30.5 | 24.21 | -20.6 | bb | 2024-02-08
117 | 117 | 20240206_HepClReRun_Ref_117 | Blank | Blank | [9 empty] | 2024-02-08
118 | 118 | 20240206_HepClReRun_Ref_118 | Media T0 A | Analyte | 3.8 | 514244.094 | 17834656 | 59452.156 | 864.971 |  | 113.33 |  | bb | 2024-02-08
119 | 119 | 20240206_HepClReRun_Ref_119 | Media T0 B | Analyte | 3.79 | 496787.813 | 17352444 | 73160.086 | 679.042 |  | 88.87 |  | bb | 2024-02-08
120 | 120 | 20240206_HepClReRun_Ref_120 | Media T0 C | Analyte | 3.8 | 458327.563 | 16039667 | 81241.563 | 564.154 |  | 73.76 |  | bb | 2024-02-08
121 | 121 | 20240206_HepClReRun_Ref_121 | Media T15 A | Analyte | 3.79 | 698198.75 | 24233036 | 203907.672 | 342.409 |  | 44.59 |  | bb | 2024-02-08
122 | 122 | 20240206_HepClReRun_Ref_122 | Media T15 B | Analyte | 3.8 | 635170.063 | 22181076 | 205969.297 | 308.381 |  | 40.11 |  | bb | 2024-02-08
123 | 123 | 20240206_HepClReRun_Ref_123 | Media T15 C | Analyte | 3.79 | 601084.375 | 20824372 | 123378.438 | 487.188 |  | 63.63 |  | bb | 2024-02-08
124 | 124 | 20240206_HepClReRun_Ref_124 | Blank | Blank |  |  |  | 2.987 | [5 empty] | 2024-02-08
125 | 125 | 20240206_HepClReRun_Ref_125 | Media T30 A | Analyte | 3.79 | 678412.188 | 23676540 | 105990.336 | 640.07 |  | 83.75 |  | bb | 2024-02-08
126 | 126 | 20240206_HepClReRun_Ref_126 | Media T30 B | Analyte | 3.79 | 638405.5 | 22125304 | 92672.016 | 688.887 |  | 90.17 |  | bb | 2024-02-08
127 | 127 | 20240206_HepClReRun_Ref_127 | Media T30 C | Analyte | 3.8 | 635660.25 | 22192772 | 122775.945 | 517.74 |  | 67.65 |  | bb | 2024-02-08
128 | 128 | 20240206_HepClReRun_Ref_128 | Media T60 A | Analyte | 3.79 | 705459.438 | 24660066 | 109065.461 | 646.822 |  | 84.63 |  | bb | 2024-02-08
129 | 129 | 20240206_HepClReRun_Ref_129 | Media T60 B | Analyte | 3.8 | 608528.625 | 21181664 | 111220.977 | 547.135 |  | 71.52 |  | bb | 2024-02-08
130 | 130 | 20240206_HepClReRun_Ref_130 | Media T60 C | Analyte | 3.79 | 686196.813 | 23924096 | 88635.961 | 774.174 |  | 101.39 |  | bb | 2024-02-08
131 | 131 | 20240206_HepClReRun_Ref_131 | Blank | Blank | [9 empty] | 2024-02-08
132 | 132 | 20240206_HepClReRun_Ref_132 | CC12 | Standard | 3.79 | 190385.391 | 6703584 | 99651.063 | 191.052 | 30.5 | 24.67 | -19.1 | bb | 2024-02-08
133 | 133 | 20240206_HepClReRun_Ref_133 | Blank | Blank |  |  |  | 7.218 | [5 empty] | 2024-02-08
134 | 134 | 20240206_HepClReRun_Ref_134 | Media T120 A | Analyte | 3.79 | 560979.875 | 19509776 | 99175.742 | 565.642 |  | 73.95 |  | bb | 2024-02-08
135 | 135 | 20240206_HepClReRun_Ref_135 | Media T120 B | Analyte | 3.79 | 632460.938 | 22164536 | 153366.016 | 412.387 |  | 53.79 |  | bb | 2024-02-08
136 | 136 | 20240206_HepClReRun_Ref_136 | Media T120 C | Analyte | 3.79 | 600658.625 | 20927208 | 156841.313 | 382.972 |  | 49.92 |  | bb | 2024-02-08
137 | 137 | 20240206_HepClReRun_Ref_137 | Media T240 A | Analyte | 3.79 | 502612.063 | 17573236 | 126010.648 | 398.865 |  | 52.01 |  | bb | 2024-02-08
138 | 138 | 20240206_HepClReRun_Ref_138 | Media T240 B | Analyte | 3.79 | 533611.875 | 18618588 | 133898.75 | 398.519 |  | 51.97 |  | bb | 2024-02-08
139 | 139 | 20240206_HepClReRun_Ref_139 | Media T240 C | Analyte | 3.79 | 533907.75 | 18771786 | 103058.133 | 518.065 |  | 67.7 |  | bb | 2024-02-08
140 | 140 | 20240206_HepClReRun_Ref_140 | Blank | Blank |  |  |  | 11.756 | [5 empty] | 2024-02-08
141 | 141 | 20240206_HepClReRun_Ref_141 | CC12 | Standard | 3.8 | 179850.547 | 6251511 | 87360.313 | 205.872 | 30.5 | 26.62 | -12.7 | bb | 2024-02-08
142 | 142 | 20240206_HepClReRun_Ref_142 | Blank | Blank |  |  |  | 2.465 | [5 empty] | 2024-02-08
143 | 143 | 20240206_HepClReRun_Ref_143 | Inactive T0 A | Analyte | 3.79 | 384702.844 | 13394410 | 67915.203 | 566.446 |  | 74.06 |  | bb | 2024-02-08
144 | 144 | 20240206_HepClReRun_Ref_144 | Inactive T0 B | Analyte | 3.79 | 398442.688 | 13978169 | 82405.945 | 483.512 |  | 63.15 |  | bb | 2024-02-08
145 | 145 | 20240206_HepClReRun_Ref_145 | Inactive T0 C | Analyte | 3.79 | 386388.594 | 13546750 | 57695.191 | 669.707 |  | 87.64 |  | bb | 2024-02-09
146 | 146 | 20240206_HepClReRun_Ref_146 | Inactive T240 A | Analyte | 3.79 | 565748.938 | 19768738 | 110553.828 | 511.741 |  | 66.86 |  | bb | 2024-02-09
147 | 147 | 20240206_HepClReRun_Ref_147 | Inactive T240 B | Analyte | 3.79 | 567725.813 | 19909150 | 79907.016 | 710.483 |  | 93.01 |  | bb | 2024-02-09
148 | 148 | 20240206_HepClReRun_Ref_148 | Inactive T240 C | Analyte | 3.79 | 596142.5 | 20803386 | 103302.727 | 577.083 |  | 75.46 |  | bb | 2024-02-09
149 | 149 | 20240206_HepClReRun_Ref_149 | Blank | Blank |  |  |  | 6.103 | [5 empty] | 2024-02-09
150 | 150 | 20240206_HepClReRun_Ref_150 | Blank | Blank | [9 empty] | 2024-02-09
151 | 151 | 20240206_HepClReRun_Ref_151 | Blank | Blank |  |  |  | 11.433 | [5 empty] | 2024-02-09
Compound 2: propranolol
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20240206_HepClReRun_Ref_001 | Blank | Blank | 4.34 | 29.845 | 1236 | [5 empty] | bb | 2024-02-07
2 | 2 | 20240206_HepClReRun_Ref_002 | Blank | Blank |  |  |  | 11.192 | [5 empty] | 2024-02-07
3 | 3 | 20240206_HepClReRun_Ref_003 | Matrix Blank | Blank |  |  |  | 15.171 | [5 empty] | 2024-02-07
4 | 4 | 20240206_HepClReRun_Ref_004 | Crash Blank | Blank |  |  |  | 1.89 | [5 empty] | 2024-02-07
5 | 5 | 20240206_HepClReRun_Ref_005 | Blank | Blank | [9 empty] | 2024-02-07
6 | 6 | 20240206_HepClReRun_Ref_006 | Blank | Blank | [9 empty] | 2024-02-07
7 | 7 | 20240206_HepClReRun_Ref_007 | CC01 | Analyte | 4.33 | 5689.459 | 100067 | 99147.227 | 5.738 | 0.175 |  |  | bbI | 2024-02-07
8 | 8 | 20240206_HepClReRun_Ref_008 | CC02 | Analyte | 4.33 | 8380.075 | 161491 | 110029.922 | 7.616 | 0.281 |  |  | bbI | 2024-02-07
9 | 9 | 20240206_HepClReRun_Ref_009 | CC03 | Analyte | 4.33 | 19523.49 | 399061 | 97847.359 | 19.953 | 0.444 | 0.09 | -80.3 | bb | 2024-02-07
10 | 10 | 20240206_HepClReRun_Ref_010 | CC04 | Analyte | 4.33 | 27509.355 | 557778 | 56906.484 | 48.341 | 0.713 | 0.95 | 33.3 | bb | 2024-02-07
11 | 11 | 20240206_HepClReRun_Ref_011 | CC05 | Standard | 4.33 | 38174.992 | 757409 | 62519.641 | 61.061 | 1.138 | 1.34 | 17.5 | bb | 2024-02-07
12 | 12 | 20240206_HepClReRun_Ref_012 | CC06 | Standard | 4.33 | 69427.516 | 1317779 | 88497.523 | 78.451 | 1.825 | 1.86 | 2.2 | bb | 2024-02-07
13 | 13 | 20240206_HepClReRun_Ref_013 | Blank | Blank | 4.33 | 10.779 | 831 | 14.119 | 76.344 |  | 1.8 |  | bb | 2024-02-07
14 | 14 | 20240206_HepClReRun_Ref_014 | CC07 | Standard | 4.33 | 85135.219 | 1578636 | 76204.688 | 111.719 | 2.919 | 2.87 | -1.5 | bb | 2024-02-07
15 | 15 | 20240206_HepClReRun_Ref_015 | CC08 | Standard | 4.33 | 121911.633 | 2212706 | 76172.734 | 160.046 | 4.669 | 4.34 | -7 | bb | 2024-02-08
16 | 16 | 20240206_HepClReRun_Ref_016 | CC09 | Standard | 4.33 | 231229.125 | 4163824 | 79696.414 | 290.137 | 7.45 | 8.29 | 11.3 | bb | 2024-02-08
17 | 17 | 20240206_HepClReRun_Ref_017 | CC10 | Standard | 4.33 | 334437.344 | 6538882 | 81798.641 | 408.854 | 11.875 | 11.9 | 0.2 | bb | 2024-02-08
18 | 18 | 20240206_HepClReRun_Ref_018 | CC11 | Standard | 4.33 | 544046.563 | 10097876 | 72615.586 | 749.215 | 19.125 | 22.24 | 16.3 | bb | 2024-02-08
19 | 19 | 20240206_HepClReRun_Ref_019 | CC12 | Standard | 4.33 | 860768.125 | 16740276 | 76026 | 1132.202 | 30.5 | 33.87 | 11.1 | bb | 2024-02-08
20 | 20 | 20240206_HepClReRun_Ref_020 | Blank | Blank | 4.34 | 1514.958 | 22611 | 11.739 | 12905.341 |  | 391.51 |  | bb | 2024-02-08
21 | 21 | 20240206_HepClReRun_Ref_021 | CC13 | Standard | 4.33 | 870822.563 | 19199102 | 46317.648 | 1880.11 | 48.75 | 56.59 | 16.1 | bb | 2024-02-08
22 | 22 | 20240206_HepClReRun_Ref_022 | CC14 | Standard | 4.33 | 2365647 | 42121644 | 99196.797 | 2384.802 | 78.125 | 71.93 | -7.9 | bb | 2024-02-08
23 | 23 | 20240206_HepClReRun_Ref_023 | CC15 | Standard | 4.33 | 3551716.25 | 66769992 | 75629.461 | 4696.207 | 125 | 142.14 | 13.7 | bb | 2024-02-08
24 | 24 | 20240206_HepClReRun_Ref_024 | CC16 | Standard | 4.33 | 5274319 | 89664752 | 82667.789 | 6380.138 | 200 | 193.29 | -3.4 | bb | 2024-02-08
25 | 25 | 20240206_HepClReRun_Ref_025 | CC17 | Analyte | 4.33 | 8462794 | 160269728 | 103689.242 | 8161.69 | 325.625 | 247.41 | -24 | bb | 2024-02-08
26 | 26 | 20240206_HepClReRun_Ref_026 | Blank | Blank | 4.33 | 17999.15 | 280841 | 7.415 | 242739.717 |  | 7373.29 |  | bb | 2024-02-08
27 | 27 | 20240206_HepClReRun_Ref_027 | Blank | Blank | 4.33 | 3748.776 | 50957 | 18.043 | 20776.9 |  | 630.63 |  | bb | 2024-02-08
28 | 28 | 20240206_HepClReRun_Ref_028 | CCQC A | QC | 4.33 | 31258.584 | 606356 | 87326.75 | 35.795 | 0.625 | 0.57 | -9 | bb | 2024-02-08
29 | 29 | 20240206_HepClReRun_Ref_029 | CCQC B | QC | 4.33 | 89345.578 | 1917429 | 81890.242 | 109.104 | 2.506 | 2.8 | 11.5 | bb | 2024-02-08
30 | 30 | 20240206_HepClReRun_Ref_030 | CCQC C | QC | 4.33 | 238971.328 | 4634862 | 103744.516 | 230.346 | 6.269 | 6.48 | 3.3 | bb | 2024-02-08
31 | 31 | 20240206_HepClReRun_Ref_031 | CCQC D | QC | 4.34 | 913307 | 16551990 | 72114.469 | 1266.468 | 25 | 37.95 | 51.8 | bb | 2024-02-08
32 | 32 | 20240206_HepClReRun_Ref_032 | Blank | Blank | 4.33 | 1958.397 | 36961 | 1.047 | 187048.424 |  | 5681.53 |  | bb | 2024-02-08
33 | 33 | 20240206_HepClReRun_Ref_033 | Blank | Blank | 4.34 | 286.214 | 5789 | 12.055 | 2374.235 |  | 71.6 |  | bb | 2024-02-08
34 | 34 | 20240206_HepClReRun_Ref_034 | CC04 | Analyte | 4.34 | 27683.043 | 577809 | 50823.348 | 54.469 | 0.713 | 1.14 | 59.4 | bb | 2024-02-08
35 | 35 | 20240206_HepClReRun_Ref_035 | CC05 | Standard | 4.33 | 40706.043 | 803446 | 81373.328 | 50.024 | 1.138 | 1 | -12 | bb | 2024-02-08
36 | 36 | 20240206_HepClReRun_Ref_036 | CC06 | Standard | 4.34 | 68328.227 | 1308721 | 96763.203 | 70.614 | 1.825 | 1.63 | -10.9 | bb | 2024-02-08
37 | 37 | 20240206_HepClReRun_Ref_037 | CC13 | Analyte | 4.34 | 922854.563 | 17027436 | 39609.715 | 2329.869 | 48.75 | 70.26 | 44.1 | bb | 2024-02-08
38 | 38 | 20240206_HepClReRun_Ref_038 | Blank | Blank | 4.33 | 1089.875 | 24160 | [5 empty] | bb | 2024-02-08
39 | 39 | 20240206_HepClReRun_Ref_039 | CC04 | Analyte | 4.34 | 24520.639 | 468125 | 45127.563 | 54.336 | 0.713 | 1.13 | 58.8 | bb | 2024-02-08
40 | 40 | 20240206_HepClReRun_Ref_040 | CC05 | Standard | 4.34 | 36367.973 | 702032 | 66556.109 | 54.643 | 1.138 | 1.14 | 0.3 | bb | 2024-02-08
41 | 41 | 20240206_HepClReRun_Ref_041 | CC06 | Standard | 4.33 | 70868.727 | 1373456 | 86041.914 | 82.365 | 1.825 | 1.98 | 8.7 | bb | 2024-02-08
42 | 42 | 20240206_HepClReRun_Ref_042 | CC13 | Standard | 4.34 | 849091 | 16840478 | 43471.086 | 1953.232 | 48.75 | 58.82 | 20.6 | bb | 2024-02-08
43 | 43 | 20240206_HepClReRun_Ref_043 | Blank | Blank | 4.33 | 1211.203 | 25076 | 9.778 | 12387.022 |  | 375.77 |  | bb | 2024-02-08
44 | 44 | 20240206_HepClReRun_Ref_044 | CC04 | Analyte | 4.34 | 25043.141 | 432078 | 50215.895 | 49.871 | 0.713 | 1 | 39.8 | bb | 2024-02-08
45 | 45 | 20240206_HepClReRun_Ref_045 | CC05 | Standard | 4.34 | 36971.234 | 738878 | 75406.516 | 49.029 | 1.138 | 0.97 | -14.7 | bb | 2024-02-08
46 | 46 | 20240206_HepClReRun_Ref_046 | CC06 | Standard | 4.33 | 60776.066 | 1089761 | 89999.438 | 67.529 | 1.825 | 1.53 | -16 | bb | 2024-02-08
47 | 47 | 20240206_HepClReRun_Ref_047 | CC13 | Analyte | 4.34 | 1039245.25 | 19600076 | 38000.164 | 2734.844 | 48.75 | 82.56 | 69.4 | bb | 2024-02-08
48 | 48 | 20240206_HepClReRun_Ref_048 | Blank | Blank | 4.33 | 1172.217 | 19600 | 16.982 | 6902.703 |  | 209.17 |  | bb | 2024-02-08
49 | 49 | 20240206_HepClReRun_Ref_049 | Blank | Blank | 4.33 | 231.257 | 5935 | [5 empty] | bb | 2024-02-08
50 | 50 | 20240206_HepClReRun_Ref_050 | CC01 | Analyte | 4.33 | 6429.335 | 117375 | 91640.219 | 7.016 | 0.175 |  |  | bbI | 2024-02-08
51 | 51 | 20240206_HepClReRun_Ref_051 | CC02 | Analyte | 4.34 | 9527.736 | 173582 | 118718.242 | 8.026 | 0.281 |  |  | bbI | 2024-02-08
52 | 52 | 20240206_HepClReRun_Ref_052 | CC03 | Analyte | 4.34 | 22009.053 | 436498 | 104661.281 | 21.029 | 0.444 | 0.12 | -73 | bb | 2024-02-08
53 | 53 | 20240206_HepClReRun_Ref_053 | CC04 | Analyte | 4.34 | 27264.674 | 495187 | 51603.223 | 52.835 | 0.713 | 1.09 | 52.4 | bb | 2024-02-08
54 | 54 | 20240206_HepClReRun_Ref_054 | CC05 | Standard | 4.34 | 40511.148 | 799950 | 76128.5 | 53.214 | 1.138 | 1.1 | -3.5 | bb | 2024-02-08
55 | 55 | 20240206_HepClReRun_Ref_055 | CC06 | Standard | 4.34 | 67995.023 | 1427447 | 86843.813 | 78.296 | 1.825 | 1.86 | 1.9 | bb | 2024-02-08
56 | 56 | 20240206_HepClReRun_Ref_056 | Blank | Blank | [9 empty] | 2024-02-08
57 | 57 | 20240206_HepClReRun_Ref_057 | CC07 | Standard | 4.34 | 102638.148 | 1975266 | 83129.859 | 123.467 | 2.919 | 3.23 | 10.7 | bb | 2024-02-08
58 | 58 | 20240206_HepClReRun_Ref_058 | CC08 | Standard | 4.34 | 162788.688 | 2998259 | 95656.969 | 170.18 | 4.669 | 4.65 | -0.4 | bb | 2024-02-08
59 | 59 | 20240206_HepClReRun_Ref_059 | CC09 | Standard | 4.34 | 244243.359 | 4911281 | 92235.273 | 264.805 | 7.45 | 7.53 | 1 | bb | 2024-02-08
60 | 60 | 20240206_HepClReRun_Ref_060 | CC10 | Standard | 4.34 | 372020.063 | 7294902 | 97453.102 | 381.743 | 11.875 | 11.08 | -6.7 | bb | 2024-02-08
61 | 61 | 20240206_HepClReRun_Ref_061 | CC11 | Standard | 4.34 | 663656.313 | 12885014 | 86978.148 | 763.015 | 19.125 | 22.66 | 18.5 | bb | 2024-02-08
62 | 62 | 20240206_HepClReRun_Ref_062 | CC12 | Standard | 4.34 | 887874.5 | 16817142 | 85374.516 | 1039.976 | 30.5 | 31.07 | 1.9 | bb | 2024-02-08
63 | 63 | 20240206_HepClReRun_Ref_063 | Blank | Blank | 4.33 | 1801.626 | 29332 | 1.61 | 111902.236 |  | 3398.78 |  | bb | 2024-02-08
64 | 64 | 20240206_HepClReRun_Ref_064 | CC13 | Standard | 4.34 | 946185.625 | 19434664 | 50045.719 | 1890.642 | 48.75 | 56.91 | 16.7 | bb | 2024-02-08
65 | 65 | 20240206_HepClReRun_Ref_065 | CC14 | Standard | 4.33 | 2664637.25 | 50079728 | 116990.578 | 2277.651 | 78.125 | 68.67 | -12.1 | bb | 2024-02-08
66 | 66 | 20240206_HepClReRun_Ref_066 | CC15 | Standard | 4.33 | 3385405.25 | 55499456 | 93066.141 | 3637.634 | 125 | 109.98 | -12 | bb | 2024-02-08
67 | 67 | 20240206_HepClReRun_Ref_067 | CC16 | Standard | 4.33 | 5772579 | 110625960 | 96650.32 | 5972.643 | 200 | 180.91 | -9.5 | bb | 2024-02-08
68 | 68 | 20240206_HepClReRun_Ref_068 | CC17 | Analyte | 4.34 | 9179669 | 166081136 | 93195.711 | 9849.884 | 325.625 | 298.7 | -8.3 | bb | 2024-02-08
69 | 69 | 20240206_HepClReRun_Ref_069 | Blank | Blank | 4.33 | 21903.471 | 355734 | [5 empty] | bb | 2024-02-08
70 | 70 | 20240206_HepClReRun_Ref_070 | Blank | Blank | 4.33 | 4349.25 | 71317 | [5 empty] | bb | 2024-02-08
71 | 71 | 20240206_HepClReRun_Ref_071 | QC A | QC | 4.34 | 32431.244 | 613701 | 95330.953 | 34.02 | 0.625 | 0.51 | -17.7 | bb | 2024-02-08
72 | 72 | 20240206_HepClReRun_Ref_072 | QC B | QC | 4.34 | 89657.414 | 1547955 | 79754.047 | 112.417 | 2.506 | 2.9 | 15.6 | bb | 2024-02-08
73 | 73 | 20240206_HepClReRun_Ref_073 | QC C | QC | 4.34 | 241690 | 4742805 | 116253.453 | 207.899 | 6.269 | 5.8 | -7.5 | bb | 2024-02-08
74 | 74 | 20240206_HepClReRun_Ref_074 | QC D | QC | 4.34 | 1043704.25 | 20803620 | 62696.5 | 1664.693 | 25 | 50.05 | 100.2 | bb | 2024-02-08
75 | 75 | 20240206_HepClReRun_Ref_075 | Blank | Blank | 4.33 | 1827.312 | 31153 | 9.576 | 19082.206 |  | 579.15 |  | bb | 2024-02-08
76 | 76 | 20240206_HepClReRun_Ref_076 | Blank | Blank | 4.34 | 536.954 | 11782 | [5 empty] | bb | 2024-02-08
77 | 77 | 20240206_HepClReRun_Ref_077 | CC04 | Analyte | 4.34 | 31107.773 | 647261 | 55185.578 | 56.369 | 0.713 | 1.19 | 67.5 | bb | 2024-02-08
78 | 78 | 20240206_HepClReRun_Ref_078 | CC05 | Standard | 4.34 | 39960.922 | 796852 | 80799.367 | 49.457 | 1.138 | 0.98 | -13.5 | bb | 2024-02-08
79 | 79 | 20240206_HepClReRun_Ref_079 | CC06 | Standard | 4.34 | 62070.469 | 1214556 | 93015.078 | 66.732 | 1.825 | 1.51 | -17.4 | bb | 2024-02-08
80 | 80 | 20240206_HepClReRun_Ref_080 | CC13 | Analyte | 4.33 | 1000756.313 | 19085440 | 46678.375 | 2143.94 | 48.75 | 64.61 | 32.5 | bb | 2024-02-08
81 | 81 | 20240206_HepClReRun_Ref_081 | Blank | Blank | 4.33 | 1757.605 | 35551 | 8.46 | 20775.473 |  | 630.59 |  | bb | 2024-02-08
82 | 82 | 20240206_HepClReRun_Ref_082 | CC04 | Analyte | 4.34 | 28283.32 | 513277 | 57444.809 | 49.236 | 0.713 | 0.98 | 37.1 | bb | 2024-02-08
83 | 83 | 20240206_HepClReRun_Ref_083 | CC05 | Standard | 4.34 | 36939.109 | 709023 | 73697.18 | 50.123 | 1.138 | 1 | -11.8 | bb | 2024-02-08
84 | 84 | 20240206_HepClReRun_Ref_084 | CC06 | Standard | 4.33 | 75094.086 | 1490398 | 102521.625 | 73.247 | 1.825 | 1.71 | -6.5 | bb | 2024-02-08
85 | 85 | 20240206_HepClReRun_Ref_085 | CC13 | Analyte | 4.34 | 966794.625 | 18279832 | 39008.566 | 2478.416 | 48.75 | 74.77 | 53.4 | bb | 2024-02-08
86 | 86 | 20240206_HepClReRun_Ref_086 | Blank | Blank | 4.34 | 2099.177 | 30580 | 1.97 | 106557.208 |  | 3236.42 |  | bb | 2024-02-08
87 | 87 | 20240206_HepClReRun_Ref_087 | CC04 | Analyte | 4.34 | 22541.699 | 376288 | 46731.051 | 48.237 | 0.713 | 0.95 | 32.8 | bb | 2024-02-08
88 | 88 | 20240206_HepClReRun_Ref_088 | CC05 | Standard | 4.34 | 34056.609 | 664162 | 71098.961 | 47.9 | 1.138 | 0.94 | -17.7 | bb | 2024-02-08
89 | 89 | 20240206_HepClReRun_Ref_089 | CC06 | Standard | 4.34 | 74154.109 | 1437870 | 93787.023 | 79.066 | 1.825 | 1.88 | 3.2 | bb | 2024-02-08
90 | 90 | 20240206_HepClReRun_Ref_090 | CC13 | Analyte | 4.34 | 1063102.875 | 19946494 | 29047.742 | 3659.847 | 48.75 | 110.66 | 127 | bb | 2024-02-08
91 | 91 | 20240206_HepClReRun_Ref_091 | Blank | Blank | 4.32 | 1580.214 | 27385 | 9.14 | 17288.993 |  | 524.68 |  | bb | 2024-02-08
92 | 92 | 20240206_HepClReRun_Ref_092 | Blank | Blank | 4.35 | 354.868 | 7123 | 0.86 | 41263.721 |  | 1252.97 |  | bb | 2024-02-08
93 | 93 | 20240206_HepClReRun_Ref_093 | Active T0 A | Analyte | 4.33 | 2294531.25 | 42666740 | 159468.484 | 1438.862 |  | 43.19 |  | bb | 2024-02-08
94 | 94 | 20240206_HepClReRun_Ref_094 | Active T0 B | Analyte | 4.33 | 2354053.25 | 45344624 | 137772.844 | 1708.648 |  | 51.39 |  | bb | 2024-02-08
95 | 95 | 20240206_HepClReRun_Ref_095 | Active T0 C | Analyte | 4.33 | 2053148 | 40953388 | 122639.313 | 1674.135 |  | 50.34 |  | bb | 2024-02-08
96 | 96 | 20240206_HepClReRun_Ref_096 | Active T15 A | Analyte | 4.33 | 2721176.25 | 52004224 | 191937.047 | 1417.744 |  | 42.55 |  | bb | 2024-02-08
97 | 97 | 20240206_HepClReRun_Ref_097 | Active T15 B | Analyte | 4.33 | 3183320 | 60800472 | 270446.25 | 1177.062 |  | 35.24 |  | bb | 2024-02-08
98 | 98 | 20240206_HepClReRun_Ref_098 | Active T15 C | Analyte | 4.32 | 2907235 | 55977992 | 228889.453 | 1270.148 |  | 38.07 |  | bb | 2024-02-08
99 | 99 | 20240206_HepClReRun_Ref_099 | Blank | Blank | 4.32 | 5665.304 | 88952 | 10.602 | 53436.182 |  | 1622.74 |  | bb | 2024-02-08
100 | 100 | 20240206_HepClReRun_Ref_100 | Active T30 A | Analyte | 4.33 | 2042914 | 38590028 | 151167.922 | 1351.42 |  | 40.53 |  | bb | 2024-02-08
101 | 101 | 20240206_HepClReRun_Ref_101 | Active T30 B | Analyte | 4.33 | 1920954.625 | 33420922 | 137365.531 | 1398.426 |  | 41.96 |  | bb | 2024-02-08
102 | 102 | 20240206_HepClReRun_Ref_102 | Active T30 C | Analyte | 4.32 | 2511156.5 | 49497012 | 188103.422 | 1334.987 |  | 40.03 |  | bb | 2024-02-08
103 | 103 | 20240206_HepClReRun_Ref_103 | Active T60 A | Analyte | 4.32 | 1781806 | 33449774 | 140990.125 | 1263.781 |  | 37.87 |  | bb | 2024-02-08
104 | 104 | 20240206_HepClReRun_Ref_104 | Active T60 B | Analyte | 4.32 | 1678213.625 | 32093362 | 193763.313 | 866.115 |  | 25.79 |  | bb | 2024-02-08
105 | 105 | 20240206_HepClReRun_Ref_105 | Active T60 C | Analyte | 4.32 | 1736233.375 | 36583268 | 134827.359 | 1287.746 |  | 38.6 |  | bb | 2024-02-08
106 | 106 | 20240206_HepClReRun_Ref_106 | Blank | Blank | 4.32 | 3944.261 | 65031 | [5 empty] | bb | 2024-02-08
107 | 107 | 20240206_HepClReRun_Ref_107 | CC12 | Standard | 4.33 | 1076495.125 | 20344942 | 105078.148 | 1024.471 | 30.5 | 30.6 | 0.3 | bb | 2024-02-08
108 | 108 | 20240206_HepClReRun_Ref_108 | Blank | Blank | 4.32 | 2588.102 | 39532 | 3.136 | 82528.763 |  | 2506.49 |  | bb | 2024-02-08
109 | 109 | 20240206_HepClReRun_Ref_109 | Active T120 A | Analyte | 4.33 | 532864.438 | 10242889 | 66314.43 | 803.542 |  | 23.89 |  | bb | 2024-02-08
110 | 110 | 20240206_HepClReRun_Ref_110 | Active T120 B | Analyte | 4.33 | 794152 | 14965115 | 70664.453 | 1123.835 |  | 33.62 |  | bb | 2024-02-08
111 | 111 | 20240206_HepClReRun_Ref_111 | Active T120 C | Analyte | 4.33 | 858142.438 | 17149398 | 90734.82 | 945.77 |  | 28.21 |  | bb | 2024-02-08
112 | 112 | 20240206_HepClReRun_Ref_112 | Active T240 A | Analyte | 4.33 | 524098.875 | 9971786 | 150100.688 | 349.165 |  | 10.09 |  | bb | 2024-02-08
113 | 113 | 20240206_HepClReRun_Ref_113 | Active T240 B | Analyte | 4.33 | 420662.5 | 8655304 | 121538.523 | 346.115 |  | 10 |  | bb | 2024-02-08
114 | 114 | 20240206_HepClReRun_Ref_114 | Active T240 C | Analyte | 4.33 | 499231.844 | 9808518 | 141360.891 | 353.161 |  | 10.21 |  | bb | 2024-02-08
115 | 115 | 20240206_HepClReRun_Ref_115 | Blank | Blank | 4.32 | 1192.125 | 20112 | 7.786 | 15311.135 |  | 464.59 |  | bb | 2024-02-08
116 | 116 | 20240206_HepClReRun_Ref_116 | CC12 | Standard | 4.33 | 986586.875 | 18464340 | 92012.313 | 1072.234 | 30.5 | 32.05 | 5.1 | bb | 2024-02-08
117 | 117 | 20240206_HepClReRun_Ref_117 | Blank | Blank | 4.32 | 1896.822 | 26264 | [5 empty] | bb | 2024-02-08
118 | 118 | 20240206_HepClReRun_Ref_118 | Media T0 A | Analyte | 4.33 | 2019758.875 | 37338576 | 59452.156 | 3397.284 |  | 102.68 |  | bb | 2024-02-08
119 | 119 | 20240206_HepClReRun_Ref_119 | Media T0 B | Analyte | 4.33 | 1882942.125 | 36278996 | 73160.086 | 2573.729 |  | 77.66 |  | bb | 2024-02-08
120 | 120 | 20240206_HepClReRun_Ref_120 | Media T0 C | Analyte | 4.33 | 1664458.75 | 30950594 | 81241.563 | 2048.777 |  | 61.72 |  | bb | 2024-02-08
121 | 121 | 20240206_HepClReRun_Ref_121 | Media T15 A | Analyte | 4.32 | 3340735.25 | 58716056 | 203907.672 | 1638.357 |  | 49.25 |  | bb | 2024-02-08
122 | 122 | 20240206_HepClReRun_Ref_122 | Media T15 B | Analyte | 4.32 | 2807132.75 | 53510668 | 205969.297 | 1362.889 |  | 40.88 |  | bb | 2024-02-08
123 | 123 | 20240206_HepClReRun_Ref_123 | Media T15 C | Analyte | 4.33 | 2809799 | 51944668 | 123378.438 | 2277.383 |  | 68.66 |  | bb | 2024-02-08
124 | 124 | 20240206_HepClReRun_Ref_124 | Blank | Blank | 4.32 | 5829.389 | 99283 | 2.987 | 195158.654 |  | 5927.9 |  | bb | 2024-02-08
125 | 125 | 20240206_HepClReRun_Ref_125 | Media T30 A | Analyte | 4.32 | 3186197.25 | 61430992 | 105990.336 | 3006.121 |  | 90.8 |  | bb | 2024-02-08
126 | 126 | 20240206_HepClReRun_Ref_126 | Media T30 B | Analyte | 4.33 | 2758034.25 | 51965436 | 92672.016 | 2976.124 |  | 89.89 |  | bb | 2024-02-08
127 | 127 | 20240206_HepClReRun_Ref_127 | Media T30 C | Analyte | 4.33 | 3248939 | 62116900 | 122775.945 | 2646.234 |  | 79.87 |  | bb | 2024-02-08
128 | 128 | 20240206_HepClReRun_Ref_128 | Media T60 A | Analyte | 4.32 | 2785065.25 | 52642948 | 109065.461 | 2553.572 |  | 77.05 |  | bb | 2024-02-08
129 | 129 | 20240206_HepClReRun_Ref_129 | Media T60 B | Analyte | 4.33 | 3777946.5 | 70176496 | 111220.977 | 3396.793 |  | 102.67 |  | bb | 2024-02-08
130 | 130 | 20240206_HepClReRun_Ref_130 | Media T60 C | Analyte | 4.33 | 2946401 | 55806236 | 88635.961 | 3324.16 |  | 100.46 |  | bb | 2024-02-08
131 | 131 | 20240206_HepClReRun_Ref_131 | Blank | Blank | 4.32 | 7125.924 | 108026 | [5 empty] | bb | 2024-02-08
132 | 132 | 20240206_HepClReRun_Ref_132 | CC12 | Standard | 4.33 | 994039.063 | 18405526 | 99651.063 | 997.52 | 30.5 | 29.78 | -2.3 | bb | 2024-02-08
133 | 133 | 20240206_HepClReRun_Ref_133 | Blank | Blank | 4.32 | 2937.011 | 42697 | 7.218 | 40690.094 |  | 1235.54 |  | bb | 2024-02-08
134 | 134 | 20240206_HepClReRun_Ref_134 | Media T120 A | Analyte | 4.33 | 2475880.25 | 46387352 | 99175.742 | 2496.458 |  | 75.32 |  | bb | 2024-02-08
135 | 135 | 20240206_HepClReRun_Ref_135 | Media T120 B | Analyte | 4.33 | 2926991.5 | 53916236 | 153366.016 | 1908.501 |  | 57.46 |  | bb | 2024-02-08
136 | 136 | 20240206_HepClReRun_Ref_136 | Media T120 C | Analyte | 4.33 | 2552598.75 | 46398920 | 156841.313 | 1627.504 |  | 48.92 |  | bb | 2024-02-08
137 | 137 | 20240206_HepClReRun_Ref_137 | Media T240 A | Analyte | 4.32 | 2148663.25 | 40641980 | 126010.648 | 1705.144 |  | 51.28 |  | bb | 2024-02-08
138 | 138 | 20240206_HepClReRun_Ref_138 | Media T240 B | Analyte | 4.32 | 2479209 | 45458668 | 133898.75 | 1851.555 |  | 55.73 |  | bb | 2024-02-08
139 | 139 | 20240206_HepClReRun_Ref_139 | Media T240 C | Analyte | 4.32 | 2299225.5 | 45059124 | 103058.133 | 2230.999 |  | 67.25 |  | bb | 2024-02-08
140 | 140 | 20240206_HepClReRun_Ref_140 | Blank | Blank | 4.32 | 4341.378 | 63516 | 11.756 | 36929.04 |  | 1121.29 |  | bb | 2024-02-08
141 | 141 | 20240206_HepClReRun_Ref_141 | CC12 | Standard | 4.33 | 975624.938 | 19203378 | 87360.313 | 1116.783 | 30.5 | 33.41 | 9.5 | bb | 2024-02-08
142 | 142 | 20240206_HepClReRun_Ref_142 | Blank | Blank | 4.33 | 2837.69 | 36534 | 2.465 | 115119.27 |  | 3496.51 |  | bb | 2024-02-08
143 | 143 | 20240206_HepClReRun_Ref_143 | Inactive T0 A | Analyte | 4.33 | 1680210.625 | 35287136 | 67915.203 | 2473.983 |  | 74.63 |  | bb | 2024-02-08
144 | 144 | 20240206_HepClReRun_Ref_144 | Inactive T0 B | Analyte | 4.33 | 1794048.625 | 34997684 | 82405.945 | 2177.086 |  | 65.62 |  | bb | 2024-02-08
145 | 145 | 20240206_HepClReRun_Ref_145 | Inactive T0 C | Analyte | 4.32 | 1585744.25 | 29815876 | 57695.191 | 2748.486 |  | 82.97 |  | bb | 2024-02-09
146 | 146 | 20240206_HepClReRun_Ref_146 | Inactive T240 A | Analyte | 4.33 | 2462339.25 | 45285364 | 110553.828 | 2227.276 |  | 67.14 |  | bb | 2024-02-09
147 | 147 | 20240206_HepClReRun_Ref_147 | Inactive T240 B | Analyte | 4.33 | 2377318.25 | 44228896 | 79907.016 | 2975.106 |  | 89.86 |  | bb | 2024-02-09
148 | 148 | 20240206_HepClReRun_Ref_148 | Inactive T240 C | Analyte | 4.32 | 2581464.75 | 45727672 | 103302.727 | 2498.932 |  | 75.39 |  | bb | 2024-02-09
149 | 149 | 20240206_HepClReRun_Ref_149 | Blank | Blank | 4.32 | 6456.986 | 100137 | 6.103 | 105800.197 |  | 3213.42 |  | bb | 2024-02-09
150 | 150 | 20240206_HepClReRun_Ref_150 | Blank | Blank | 4.32 | 1637.52 | 28260 | [5 empty] | bb | 2024-02-09
151 | 151 | 20240206_HepClReRun_Ref_151 | Blank | Blank | 4.34 | 737.453 | 11797 | 11.433 | 6450.214 |  | 195.42 |  | bb | 2024-02-09
Compound 3: M8PFOA
| # | Name | Sample Text | Type | RT | Area | Height | IS Area | Response | Std. Conc | nM | %Dev | Primary Flags | Acq.Date
1 | 1 | 20240206_HepClReRun_Ref_001 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-07
2 | 2 | 20240206_HepClReRun_Ref_002 | Blank | Blank | 3.84 | 11.192 | 512 |  | 11.192 | 100 | 0.01 | -100 | bb | 2024-02-07
3 | 3 | 20240206_HepClReRun_Ref_003 | Matrix Blank | Blank | 3.82 | 15.171 | 367 |  | 15.171 | 100 | 0.02 | -100 | bb | 2024-02-07
4 | 4 | 20240206_HepClReRun_Ref_004 | Crash Blank | Blank | 3.78 | 1.89 | 77 |  | 1.89 | 100 | 0 | -100 | bb | 2024-02-07
5 | 5 | 20240206_HepClReRun_Ref_005 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-07
6 | 6 | 20240206_HepClReRun_Ref_006 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-07
7 | 7 | 20240206_HepClReRun_Ref_007 | CC01 | Analyte | 3.79 | 99147.227 | 3690400 |  | 99147.227 | 100 | 118.98 | 19 | bb | 2024-02-07
8 | 8 | 20240206_HepClReRun_Ref_008 | CC02 | Analyte | 3.79 | 110029.922 | 4115721 |  | 110029.922 | 100 | 132.04 | 32 | bb | 2024-02-07
9 | 9 | 20240206_HepClReRun_Ref_009 | CC03 | Analyte | 3.79 | 97847.359 | 3658686 |  | 97847.359 | 100 | 117.42 | 17.4 | bb | 2024-02-07
10 | 10 | 20240206_HepClReRun_Ref_010 | CC04 | Analyte | 3.79 | 56906.484 | 2091590 |  | 56906.484 | 100 | 68.29 | -31.7 | bb | 2024-02-07
11 | 11 | 20240206_HepClReRun_Ref_011 | CC05 | Standard | 3.79 | 62519.641 | 2550542 |  | 62519.641 | 100 | 75.03 | -25 | bb | 2024-02-07
12 | 12 | 20240206_HepClReRun_Ref_012 | CC06 | Standard | 3.79 | 88497.523 | 3262129 |  | 88497.523 | 100 | 106.2 | 6.2 | bb | 2024-02-07
13 | 13 | 20240206_HepClReRun_Ref_013 | Blank | Blank | 3.79 | 14.119 | 559 |  | 14.119 | 100 | 0.02 | -100 | bb | 2024-02-07
14 | 14 | 20240206_HepClReRun_Ref_014 | CC07 | Standard | 3.79 | 76204.688 | 2781862 |  | 76204.688 | 100 | 91.45 | -8.6 | bb | 2024-02-07
15 | 15 | 20240206_HepClReRun_Ref_015 | CC08 | Standard | 3.79 | 76172.734 | 2958551 |  | 76172.734 | 100 | 91.41 | -8.6 | bb | 2024-02-08
16 | 16 | 20240206_HepClReRun_Ref_016 | CC09 | Standard | 3.79 | 79696.414 | 3734757 |  | 79696.414 | 100 | 95.64 | -4.4 | bb | 2024-02-08
17 | 17 | 20240206_HepClReRun_Ref_017 | CC10 | Standard | 3.79 | 81798.641 | 3844104 |  | 81798.641 | 100 | 98.16 | -1.8 | bb | 2024-02-08
18 | 18 | 20240206_HepClReRun_Ref_018 | CC11 | Standard | 3.79 | 72615.586 | 2704232 |  | 72615.586 | 100 | 87.14 | -12.9 | bb | 2024-02-08
19 | 19 | 20240206_HepClReRun_Ref_019 | CC12 | Standard | 3.79 | 76026 | 2870378 |  | 76026 | 100 | 91.23 | -8.8 | bb | 2024-02-08
20 | 20 | 20240206_HepClReRun_Ref_020 | Blank | Blank | 3.78 | 11.739 | 323 |  | 11.739 | 100 | 0.01 | -100 | bb | 2024-02-08
21 | 21 | 20240206_HepClReRun_Ref_021 | CC13 | Standard | 3.79 | 46317.648 | 1555216 |  | 46317.648 | 100 | 55.58 | -44.4 | bb | 2024-02-08
22 | 22 | 20240206_HepClReRun_Ref_022 | CC14 | Standard | 3.79 | 99196.797 | 4421671 |  | 99196.797 | 100 | 119.04 | 19 | bb | 2024-02-08
23 | 23 | 20240206_HepClReRun_Ref_023 | CC15 | Standard | 3.79 | 75629.461 | 2825389 |  | 75629.461 | 100 | 90.76 | -9.2 | bb | 2024-02-08
24 | 24 | 20240206_HepClReRun_Ref_024 | CC16 | Standard | 3.79 | 82667.789 | 3326529 |  | 82667.789 | 100 | 99.21 | -0.8 | bb | 2024-02-08
25 | 25 | 20240206_HepClReRun_Ref_025 | CC17 | Analyte | 3.79 | 103689.242 | 3858924 |  | 103689.242 | 100 | 124.43 | 24.4 | bb | 2024-02-08
26 | 26 | 20240206_HepClReRun_Ref_026 | Blank | Blank | 3.78 | 7.415 | 304 |  | 7.415 | 100 | 0.01 | -100 | bb | 2024-02-08
27 | 27 | 20240206_HepClReRun_Ref_027 | Blank | Blank | 3.8 | 18.043 | 493 |  | 18.043 | 100 | 0.02 | -100 | bb | 2024-02-08
28 | 28 | 20240206_HepClReRun_Ref_028 | CCQC A | QC | 3.79 | 87326.75 | 3300575 |  | 87326.75 | 100 | 104.8 | 4.8 | bb | 2024-02-08
29 | 29 | 20240206_HepClReRun_Ref_029 | CCQC B | QC | 3.79 | 81890.242 | 3061537 |  | 81890.242 | 100 | 98.27 | -1.7 | bb | 2024-02-08
30 | 30 | 20240206_HepClReRun_Ref_030 | CCQC C | QC | 3.79 | 103744.516 | 3924788 |  | 103744.516 | 100 | 124.5 | 24.5 | bb | 2024-02-08
31 | 31 | 20240206_HepClReRun_Ref_031 | CCQC D | QC | 3.79 | 72114.469 | 2688628 |  | 72114.469 | 100 | 86.54 | -13.5 | bb | 2024-02-08
32 | 32 | 20240206_HepClReRun_Ref_032 | Blank | Blank | 3.79 | 1.047 | 123 |  | 1.047 | 100 | 0 | -100 | bb | 2024-02-08
33 | 33 | 20240206_HepClReRun_Ref_033 | Blank | Blank | 3.84 | 12.055 | 251 |  | 12.055 | 100 | 0.01 | -100 | bb | 2024-02-08
34 | 34 | 20240206_HepClReRun_Ref_034 | CC04 | Analyte | 3.79 | 50823.348 | 1909275 |  | 50823.348 | 100 | 60.99 | -39 | bb | 2024-02-08
35 | 35 | 20240206_HepClReRun_Ref_035 | CC05 | Standard | 3.79 | 81373.328 | 3072050 |  | 81373.328 | 100 | 97.65 | -2.3 | bb | 2024-02-08
36 | 36 | 20240206_HepClReRun_Ref_036 | CC06 | Standard | 3.79 | 96763.203 | 3627207 |  | 96763.203 | 100 | 116.12 | 16.1 | bb | 2024-02-08
37 | 37 | 20240206_HepClReRun_Ref_037 | CC13 | Analyte | 3.79 | 39609.715 | 1478978 |  | 39609.715 | 100 | 47.53 | -52.5 | bb | 2024-02-08
38 | 38 | 20240206_HepClReRun_Ref_038 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-08
39 | 39 | 20240206_HepClReRun_Ref_039 | CC04 | Analyte | 3.79 | 45127.563 | 1700804 |  | 45127.563 | 100 | 54.16 | -45.8 | bb | 2024-02-08
40 | 40 | 20240206_HepClReRun_Ref_040 | CC05 | Standard | 3.78 | 66556.109 | 2338083 |  | 66556.109 | 100 | 79.87 | -20.1 | bb | 2024-02-08
41 | 41 | 20240206_HepClReRun_Ref_041 | CC06 | Standard | 3.79 | 86041.914 | 2949560 |  | 86041.914 | 100 | 103.25 | 3.3 | bb | 2024-02-08
42 | 42 | 20240206_HepClReRun_Ref_042 | CC13 | Standard | 3.79 | 43471.086 | 1413638 |  | 43471.086 | 100 | 52.17 | -47.8 | bb | 2024-02-08
43 | 43 | 20240206_HepClReRun_Ref_043 | Blank | Blank | 3.79 | 9.778 | 359 |  | 9.778 | 100 | 0.01 | -100 | bb | 2024-02-08
44 | 44 | 20240206_HepClReRun_Ref_044 | CC04 | Analyte | 3.79 | 50215.895 | 1900730 |  | 50215.895 | 100 | 60.26 | -39.7 | bb | 2024-02-08
45 | 45 | 20240206_HepClReRun_Ref_045 | CC05 | Standard | 3.79 | 75406.516 | 2832145 |  | 75406.516 | 100 | 90.49 | -9.5 | bb | 2024-02-08
46 | 46 | 20240206_HepClReRun_Ref_046 | CC06 | Standard | 3.79 | 89999.438 | 3457415 |  | 89999.438 | 100 | 108 | 8 | bb | 2024-02-08
47 | 47 | 20240206_HepClReRun_Ref_047 | CC13 | Analyte | 3.79 | 38000.164 | 1436974 |  | 38000.164 | 100 | 45.6 | -54.4 | bb | 2024-02-08
48 | 48 | 20240206_HepClReRun_Ref_048 | Blank | Blank | 3.81 | 16.982 | 459 |  | 16.982 | 100 | 0.02 | -100 | bb | 2024-02-08
49 | 49 | 20240206_HepClReRun_Ref_049 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-08
50 | 50 | 20240206_HepClReRun_Ref_050 | CC01 | Analyte | 3.78 | 91640.219 | 3479469 |  | 91640.219 | 100 | 109.97 | 10 | bb | 2024-02-08
51 | 51 | 20240206_HepClReRun_Ref_051 | CC02 | Analyte | 3.78 | 118718.242 | 4508226 |  | 118718.242 | 100 | 142.47 | 42.5 | bb | 2024-02-08
52 | 52 | 20240206_HepClReRun_Ref_052 | CC03 | Analyte | 3.79 | 104661.281 | 3932631 |  | 104661.281 | 100 | 125.6 | 25.6 | bb | 2024-02-08
53 | 53 | 20240206_HepClReRun_Ref_053 | CC04 | Analyte | 3.79 | 51603.223 | 1962689 |  | 51603.223 | 100 | 61.93 | -38.1 | bb | 2024-02-08
54 | 54 | 20240206_HepClReRun_Ref_054 | CC05 | Standard | 3.79 | 76128.5 | 2872571 |  | 76128.5 | 100 | 91.36 | -8.6 | bb | 2024-02-08
55 | 55 | 20240206_HepClReRun_Ref_055 | CC06 | Standard | 3.78 | 86843.813 | 3293688 |  | 86843.813 | 100 | 104.22 | 4.2 | bb | 2024-02-08
56 | 56 | 20240206_HepClReRun_Ref_056 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-08
57 | 57 | 20240206_HepClReRun_Ref_057 | CC07 | Standard | 3.78 | 83129.859 | 3025645 |  | 83129.859 | 100 | 99.76 | -0.2 | bb | 2024-02-08
58 | 58 | 20240206_HepClReRun_Ref_058 | CC08 | Standard | 3.79 | 95656.969 | 3566732 |  | 95656.969 | 100 | 114.79 | 14.8 | bb | 2024-02-08
59 | 59 | 20240206_HepClReRun_Ref_059 | CC09 | Standard | 3.79 | 92235.273 | 4208975 |  | 92235.273 | 100 | 110.69 | 10.7 | bb | 2024-02-08
60 | 60 | 20240206_HepClReRun_Ref_060 | CC10 | Standard | 3.79 | 97453.102 | 4459191 |  | 97453.102 | 100 | 116.95 | 16.9 | bb | 2024-02-08
61 | 61 | 20240206_HepClReRun_Ref_061 | CC11 | Standard | 3.79 | 86978.148 | 3261943 |  | 86978.148 | 100 | 104.38 | 4.4 | bb | 2024-02-08
62 | 62 | 20240206_HepClReRun_Ref_062 | CC12 | Standard | 3.78 | 85374.516 | 3281335 |  | 85374.516 | 100 | 102.45 | 2.5 | bb | 2024-02-08
63 | 63 | 20240206_HepClReRun_Ref_063 | Blank | Blank | 3.83 | 1.61 | 79 |  | 1.61 | 100 | 0 | -100 | bb | 2024-02-08
64 | 64 | 20240206_HepClReRun_Ref_064 | CC13 | Standard | 3.79 | 50045.719 | 1651115 |  | 50045.719 | 100 | 60.06 | -39.9 | bb | 2024-02-08
65 | 65 | 20240206_HepClReRun_Ref_065 | CC14 | Standard | 3.78 | 116990.578 | 5138411 |  | 116990.578 | 100 | 140.39 | 40.4 | bb | 2024-02-08
66 | 66 | 20240206_HepClReRun_Ref_066 | CC15 | Standard | 3.78 | 93066.141 | 3368788 |  | 93066.141 | 100 | 111.68 | 11.7 | bb | 2024-02-08
67 | 67 | 20240206_HepClReRun_Ref_067 | CC16 | Standard | 3.79 | 96650.32 | 3640127 |  | 96650.32 | 100 | 115.99 | 16 | bb | 2024-02-08
68 | 68 | 20240206_HepClReRun_Ref_068 | CC17 | Analyte | 3.78 | 93195.711 | 3550052 |  | 93195.711 | 100 | 111.84 | 11.8 | bb | 2024-02-08
69 | 69 | 20240206_HepClReRun_Ref_069 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-08
70 | 70 | 20240206_HepClReRun_Ref_070 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-08
71 | 71 | 20240206_HepClReRun_Ref_071 | QC A | QC | 3.79 | 95330.953 | 3615751 |  | 95330.953 | 100 | 114.4 | 14.4 | bb | 2024-02-08
72 | 72 | 20240206_HepClReRun_Ref_072 | QC B | QC | 3.79 | 79754.047 | 3041261 |  | 79754.047 | 100 | 95.71 | -4.3 | bb | 2024-02-08
73 | 73 | 20240206_HepClReRun_Ref_073 | QC C | QC | 3.79 | 116253.453 | 4423421 |  | 116253.453 | 100 | 139.51 | 39.5 | bb | 2024-02-08
74 | 74 | 20240206_HepClReRun_Ref_074 | QC D | QC | 3.79 | 62696.5 | 2378496 |  | 62696.5 | 100 | 75.24 | -24.8 | bb | 2024-02-08
75 | 75 | 20240206_HepClReRun_Ref_075 | Blank | Blank | 3.78 | 9.576 | 462 |  | 9.576 | 100 | 0.01 | -100 | bb | 2024-02-08
76 | 76 | 20240206_HepClReRun_Ref_076 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-08
77 | 77 | 20240206_HepClReRun_Ref_077 | CC04 | Analyte | 3.79 | 55185.578 | 2096160 |  | 55185.578 | 100 | 66.23 | -33.8 | bb | 2024-02-08
78 | 78 | 20240206_HepClReRun_Ref_078 | CC05 | Standard | 3.79 | 80799.367 | 3053570 |  | 80799.367 | 100 | 96.96 | -3 | bb | 2024-02-08
79 | 79 | 20240206_HepClReRun_Ref_079 | CC06 | Standard | 3.78 | 93015.078 | 3374432 |  | 93015.078 | 100 | 111.62 | 11.6 | bb | 2024-02-08
80 | 80 | 20240206_HepClReRun_Ref_080 | CC13 | Analyte | 3.79 | 46678.375 | 1515470 |  | 46678.375 | 100 | 56.02 | -44 | bb | 2024-02-08
81 | 81 | 20240206_HepClReRun_Ref_081 | Blank | Blank | 3.84 | 8.46 | 296 |  | 8.46 | 100 | 0.01 | -100 | bb | 2024-02-08
82 | 82 | 20240206_HepClReRun_Ref_082 | CC04 | Analyte | 3.78 | 57444.809 | 2196922 |  | 57444.809 | 100 | 68.94 | -31.1 | bb | 2024-02-08
83 | 83 | 20240206_HepClReRun_Ref_083 | CC05 | Standard | 3.78 | 73697.18 | 2806167 |  | 73697.18 | 100 | 88.44 | -11.6 | bb | 2024-02-08
84 | 84 | 20240206_HepClReRun_Ref_084 | CC06 | Standard | 3.78 | 102521.625 | 3753753 |  | 102521.625 | 100 | 123.03 | 23 | bb | 2024-02-08
85 | 85 | 20240206_HepClReRun_Ref_085 | CC13 | Analyte | 3.78 | 39008.566 | 1492758 |  | 39008.566 | 100 | 46.81 | -53.2 | bb | 2024-02-08
86 | 86 | 20240206_HepClReRun_Ref_086 | Blank | Blank | 3.78 | 1.97 | 59 |  | 1.97 | 100 | 0 | -100 | bb | 2024-02-08
87 | 87 | 20240206_HepClReRun_Ref_087 | CC04 | Analyte | 3.79 | 46731.051 | 1784513 |  | 46731.051 | 100 | 56.08 | -43.9 | bb | 2024-02-08
88 | 88 | 20240206_HepClReRun_Ref_088 | CC05 | Standard | 3.78 | 71098.961 | 2720564 |  | 71098.961 | 100 | 85.32 | -14.7 | bb | 2024-02-08
89 | 89 | 20240206_HepClReRun_Ref_089 | CC06 | Standard | 3.79 | 93787.023 | 3541642 |  | 93787.023 | 100 | 112.55 | 12.5 | bb | 2024-02-08
90 | 90 | 20240206_HepClReRun_Ref_090 | CC13 | Analyte | 3.78 | 29047.742 | 1175002 |  | 29047.742 | 100 | 34.86 | -65.1 | bb | 2024-02-08
91 | 91 | 20240206_HepClReRun_Ref_091 | Blank | Blank | 3.8 | 9.14 | 264 |  | 9.14 | 100 | 0.01 | -100 | bb | 2024-02-08
92 | 92 | 20240206_HepClReRun_Ref_092 | Blank | Blank | 3.78 | 0.86 | 57 |  | 0.86 | 100 | 0 | -100 | bb | 2024-02-08
93 | 93 | 20240206_HepClReRun_Ref_093 | Active T0 A | Analyte | 3.78 | 159468.484 | 6234953 |  | 159468.484 | 100 | 191.37 | 91.4 | bb | 2024-02-08
94 | 94 | 20240206_HepClReRun_Ref_094 | Active T0 B | Analyte | 3.78 | 137772.844 | 5556907 |  | 137772.844 | 100 | 165.33 | 65.3 | bb | 2024-02-08
95 | 95 | 20240206_HepClReRun_Ref_095 | Active T0 C | Analyte | 3.78 | 122639.313 | 4941604 |  | 122639.313 | 100 | 147.17 | 47.2 | bb | 2024-02-08
96 | 96 | 20240206_HepClReRun_Ref_096 | Active T15 A | Analyte | 3.78 | 191937.047 | 7737548 |  | 191937.047 | 100 | 230.33 | 130.3 | bb | 2024-02-08
97 | 97 | 20240206_HepClReRun_Ref_097 | Active T15 B | Analyte | 3.78 | 270446.25 | 10975762 |  | 270446.25 | 100 | 324.55 | 224.5 | bb | 2024-02-08
98 | 98 | 20240206_HepClReRun_Ref_098 | Active T15 C | Analyte | 3.78 | 228889.453 | 9083325 |  | 228889.453 | 100 | 274.68 | 174.7 | bb | 2024-02-08
99 | 99 | 20240206_HepClReRun_Ref_099 | Blank | Blank | 3.78 | 10.602 | 443 |  | 10.602 | 100 | 0.01 | -100 | bb | 2024-02-08
100 | 100 | 20240206_HepClReRun_Ref_100 | Active T30 A | Analyte | 3.78 | 151167.922 | 5595652 |  | 151167.922 | 100 | 181.41 | 81.4 | bb | 2024-02-08
101 | 101 | 20240206_HepClReRun_Ref_101 | Active T30 B | Analyte | 3.78 | 137365.531 | 4915725 |  | 137365.531 | 100 | 164.85 | 64.8 | bb | 2024-02-08
102 | 102 | 20240206_HepClReRun_Ref_102 | Active T30 C | Analyte | 3.78 | 188103.422 | 6732630 |  | 188103.422 | 100 | 225.73 | 125.7 | bb | 2024-02-08
103 | 103 | 20240206_HepClReRun_Ref_103 | Active T60 A | Analyte | 3.78 | 140990.125 | 4797147 |  | 140990.125 | 100 | 169.19 | 69.2 | bb | 2024-02-08
104 | 104 | 20240206_HepClReRun_Ref_104 | Active T60 B | Analyte | 3.78 | 193763.313 | 6576127 |  | 193763.313 | 100 | 232.53 | 132.5 | bb | 2024-02-08
105 | 105 | 20240206_HepClReRun_Ref_105 | Active T60 C | Analyte | 3.78 | 134827.359 | 4495299 |  | 134827.359 | 100 | 161.8 | 61.8 | bb | 2024-02-08
106 | 106 | 20240206_HepClReRun_Ref_106 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-08
107 | 107 | 20240206_HepClReRun_Ref_107 | CC12 | Standard | 3.78 | 105078.148 | 3696228 |  | 105078.148 | 100 | 126.1 | 26.1 | bb | 2024-02-08
108 | 108 | 20240206_HepClReRun_Ref_108 | Blank | Blank | 3.82 | 3.136 | 179 |  | 3.136 | 100 | 0 | -100 | bb | 2024-02-08
109 | 109 | 20240206_HepClReRun_Ref_109 | Active T120 A | Analyte | 3.78 | 66314.43 | 2659309 |  | 66314.43 | 100 | 79.58 | -20.4 | bb | 2024-02-08
110 | 110 | 20240206_HepClReRun_Ref_110 | Active T120 B | Analyte | 3.78 | 70664.453 | 2822867 |  | 70664.453 | 100 | 84.8 | -15.2 | bb | 2024-02-08
111 | 111 | 20240206_HepClReRun_Ref_111 | Active T120 C | Analyte | 3.79 | 90734.82 | 3446046 |  | 90734.82 | 100 | 108.89 | 8.9 | bb | 2024-02-08
112 | 112 | 20240206_HepClReRun_Ref_112 | Active T240 A | Analyte | 3.78 | 150100.688 | 5841353 |  | 150100.688 | 100 | 180.13 | 80.1 | bb | 2024-02-08
113 | 113 | 20240206_HepClReRun_Ref_113 | Active T240 B | Analyte | 3.78 | 121538.523 | 4873387 |  | 121538.523 | 100 | 145.85 | 45.9 | bb | 2024-02-08
114 | 114 | 20240206_HepClReRun_Ref_114 | Active T240 C | Analyte | 3.78 | 141360.891 | 5586628 |  | 141360.891 | 100 | 169.64 | 69.6 | bb | 2024-02-08
115 | 115 | 20240206_HepClReRun_Ref_115 | Blank | Blank | 3.82 | 7.786 | 316 |  | 7.786 | 100 | 0.01 | -100 | bb | 2024-02-08
116 | 116 | 20240206_HepClReRun_Ref_116 | CC12 | Standard | 3.78 | 92012.313 | 3411119 |  | 92012.313 | 100 | 110.42 | 10.4 | bb | 2024-02-08
117 | 117 | 20240206_HepClReRun_Ref_117 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-08
118 | 118 | 20240206_HepClReRun_Ref_118 | Media T0 A | Analyte | 3.78 | 59452.156 | 2364865 |  | 59452.156 | 100 | 71.35 | -28.7 | bb | 2024-02-08
119 | 119 | 20240206_HepClReRun_Ref_119 | Media T0 B | Analyte | 3.78 | 73160.086 | 2981668 |  | 73160.086 | 100 | 87.8 | -12.2 | bb | 2024-02-08
120 | 120 | 20240206_HepClReRun_Ref_120 | Media T0 C | Analyte | 3.78 | 81241.563 | 3158303 |  | 81241.563 | 100 | 97.49 | -2.5 | bb | 2024-02-08
121 | 121 | 20240206_HepClReRun_Ref_121 | Media T15 A | Analyte | 3.78 | 203907.672 | 8179395 |  | 203907.672 | 100 | 244.7 | 144.7 | bb | 2024-02-08
122 | 122 | 20240206_HepClReRun_Ref_122 | Media T15 B | Analyte | 3.78 | 205969.297 | 8093554 |  | 205969.297 | 100 | 247.17 | 147.2 | bb | 2024-02-08
123 | 123 | 20240206_HepClReRun_Ref_123 | Media T15 C | Analyte | 3.78 | 123378.438 | 5218319 |  | 123378.438 | 100 | 148.06 | 48.1 | bb | 2024-02-08
124 | 124 | 20240206_HepClReRun_Ref_124 | Blank | Blank | 3.82 | 2.987 | 122 |  | 2.987 | 100 | 0 | -100 | bb | 2024-02-08
125 | 125 | 20240206_HepClReRun_Ref_125 | Media T30 A | Analyte | 3.78 | 105990.336 | 4481414 |  | 105990.336 | 100 | 127.19 | 27.2 | bb | 2024-02-08
126 | 126 | 20240206_HepClReRun_Ref_126 | Media T30 B | Analyte | 3.78 | 92672.016 | 3787446 |  | 92672.016 | 100 | 111.21 | 11.2 | bb | 2024-02-08
127 | 127 | 20240206_HepClReRun_Ref_127 | Media T30 C | Analyte | 3.78 | 122775.945 | 4730254 |  | 122775.945 | 100 | 147.34 | 47.3 | bb | 2024-02-08
128 | 128 | 20240206_HepClReRun_Ref_128 | Media T60 A | Analyte | 3.78 | 109065.461 | 4490368 |  | 109065.461 | 100 | 130.88 | 30.9 | bb | 2024-02-08
129 | 129 | 20240206_HepClReRun_Ref_129 | Media T60 B | Analyte | 3.78 | 111220.977 | 4359841 |  | 111220.977 | 100 | 133.47 | 33.5 | bb | 2024-02-08
130 | 130 | 20240206_HepClReRun_Ref_130 | Media T60 C | Analyte | 3.78 | 88635.961 | 3907847 |  | 88635.961 | 100 | 106.37 | 6.4 | bb | 2024-02-08
131 | 131 | 20240206_HepClReRun_Ref_131 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-08
132 | 132 | 20240206_HepClReRun_Ref_132 | CC12 | Standard | 3.78 | 99651.063 | 3638737 |  | 99651.063 | 100 | 119.59 | 19.6 | bb | 2024-02-08
133 | 133 | 20240206_HepClReRun_Ref_133 | Blank | Blank | 3.78 | 7.218 | 368 |  | 7.218 | 100 | 0.01 | -100 | bb | 2024-02-08
134 | 134 | 20240206_HepClReRun_Ref_134 | Media T120 A | Analyte | 3.78 | 99175.742 | 4210011 |  | 99175.742 | 100 | 119.02 | 19 | bb | 2024-02-08
135 | 135 | 20240206_HepClReRun_Ref_135 | Media T120 B | Analyte | 3.78 | 153366.016 | 6537087 |  | 153366.016 | 100 | 184.05 | 84 | bb | 2024-02-08
136 | 136 | 20240206_HepClReRun_Ref_136 | Media T120 C | Analyte | 3.78 | 156841.313 | 6584501 |  | 156841.313 | 100 | 188.22 | 88.2 | bb | 2024-02-08
137 | 137 | 20240206_HepClReRun_Ref_137 | Media T240 A | Analyte | 3.78 | 126010.648 | 5389455 |  | 126010.648 | 100 | 151.22 | 51.2 | bb | 2024-02-08
138 | 138 | 20240206_HepClReRun_Ref_138 | Media T240 B | Analyte | 3.78 | 133898.75 | 5696388 |  | 133898.75 | 100 | 160.68 | 60.7 | bb | 2024-02-08
139 | 139 | 20240206_HepClReRun_Ref_139 | Media T240 C | Analyte | 3.78 | 103058.133 | 4536677 |  | 103058.133 | 100 | 123.67 | 23.7 | bb | 2024-02-08
140 | 140 | 20240206_HepClReRun_Ref_140 | Blank | Blank | 3.79 | 11.756 | 472 |  | 11.756 | 100 | 0.01 | -100 | bb | 2024-02-08
141 | 141 | 20240206_HepClReRun_Ref_141 | CC12 | Standard | 3.78 | 87360.313 | 3242247 |  | 87360.313 | 100 | 104.84 | 4.8 | bb | 2024-02-08
142 | 142 | 20240206_HepClReRun_Ref_142 | Blank | Blank | 3.83 | 2.465 | 96 |  | 2.465 | 100 | 0 | -100 | bb | 2024-02-08
143 | 143 | 20240206_HepClReRun_Ref_143 | Inactive T0 A | Analyte | 3.78 | 67915.203 | 2974119 |  | 67915.203 | 100 | 81.5 | -18.5 | bb | 2024-02-08
144 | 144 | 20240206_HepClReRun_Ref_144 | Inactive T0 B | Analyte | 3.78 | 82405.945 | 3620762 |  | 82405.945 | 100 | 98.89 | -1.1 | bb | 2024-02-08
145 | 145 | 20240206_HepClReRun_Ref_145 | Inactive T0 C | Analyte | 3.78 | 57695.191 | 2476753 |  | 57695.191 | 100 | 69.24 | -30.8 | bb | 2024-02-09
146 | 146 | 20240206_HepClReRun_Ref_146 | Inactive T240 A | Analyte | 3.78 | 110553.828 | 4797423 |  | 110553.828 | 100 | 132.67 | 32.7 | bb | 2024-02-09
147 | 147 | 20240206_HepClReRun_Ref_147 | Inactive T240 B | Analyte | 3.78 | 79907.016 | 3404979 |  | 79907.016 | 100 | 95.89 | -4.1 | bb | 2024-02-09
148 | 148 | 20240206_HepClReRun_Ref_148 | Inactive T240 C | Analyte | 3.78 | 103302.727 | 4376327 |  | 103302.727 | 100 | 123.97 | 24 | bb | 2024-02-09
149 | 149 | 20240206_HepClReRun_Ref_149 | Blank | Blank | 3.84 | 6.103 | 264 |  | 6.103 | 100 | 0.01 | -100 | bb | 2024-02-09
150 | 150 | 20240206_HepClReRun_Ref_150 | Blank | Blank | [5 empty] | 100 |  |  |  | 2024-02-09
151 | 151 | 20240206_HepClReRun_Ref_151 | Blank | Blank | 3.83 | 11.433 | 297 |  | 11.433 | 100 | 0.01 | -100 | bb | 2024-02-09